## Sheet1
Problem | Suggested solution(s)
The PDMS does not bond to the glass. | - Make sure both the PDMS and glass are clean.
- Make sure there is enough PDMS around the chamber for bonding.
- Optimize the settings of the plasma oven.
There are air bubbles on my chip. | - Check if the pump was turned on.
- Degas the chip before loading tissue onto the chip.
- Degas the medium before start of the experiment.
- Make sure the humidity in the incubation chamber is high enough.
The tissue dies at the beginning of the experiment. | - Add HEPES to the medium when loading and assembling the chip.
- Do not flush the tissue in and out too often during loading.
- Check that the flow rate is not too high.
The tissue dies during imaging. | - Make sure there is no phototoxicity from the imaging
- Make sure there is no contamination.
- The flow rate should not be too high.
The focus changes during imaging. | - Use autofocus during imaging to adjust for drift of the imaging slide.
- Let the chip equilibrate to the temperature within the microscope for at least 20 minutes.
Outlets/inlets detach. | - Push all tubing completely down to the bottom.
The glass of the chip breaks. | - Be more careful, the glass is very fragile.
- The thin glass is only required for imaging. If imaging is not performed, a normal thicker glass slide can be used.
- If the break is outside of the chip, it might not be a problem. If the glass seal to the chip is broken, liquid will leak out and air will come in.
There is a contamination on my chip. | - Add antibiotics to the culture medium.
- Sterilize all equipment and tubing.
- Work fast and clean.