Dear Editor,

Thank you very much for considering this manuscript for publication. Please find enclosed below the response to the comments and suggestions by the reviewers. We apologize for slower than expected video processing – this is due to continued coronavirus limitations and should be completed and uploaded in a couple of days.

Thank you very much for your help and consideration,

Sincerely,

A. Klibanov

Editorial and production comments:
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Thank you for the suggestion, an additional check was performed.
2. Please provide an institutional email address for each author.

Institutional address for all authors was provided during original submission; they all receive the email from the JOVE office. 

3. Please revise the text to avoid the use of any personal pronouns (e.g., "we", "you", "our" etc.).
4. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials.
For example: speedvac, Vialmix, ESPE Capmix, Wig-L-Bug, Hamamatsu Orca video camera, Gemini XS, Molecular Devices, Accusizer, Beckman Coulter, Multisizer, Coulter counter, Corning, Cadence, etc.

Done, except CPS Gain – that parameter is required to describe exact imaging conditions. 

5. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly. Please include all safety procedures and use of hoods, etc.

Done.

6. Line 91-94: Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. How is the evaporation performed? How much chloroform is used for redissolving the nonvolatile residue?

Information added.

7. Please move the notes compiled at the end of a particular section to the appropriate steps of the protocol. In the JoVE Protocol format, “Notes” should be concise and used sparingly. They should only be used to provide extraneous details, optional steps, or recommendations that are not critical to a step. Any text that provides details about how to perform a particular step should either be included in the step itself or added as a sub-step. Please consider moving some of the notes about the protocol to the discussion section (Lines: 107-134, 197-202, 225-242, 256-263, 287-309, 358-373, 417-442)

We would prefer to disagree with this suggestion; the technical details as we describe in the notes are not of the “general discussion” type, but specifically address the reasons, potential issues and pitfalls of the experimental steps as well. Reviewers also pointed at these details as the positive feature of the manuscript.
8. Line 249/350: Please use standard abbreviations for time units preceded by a numeral. Do not abbreviate day, week, month, and year. Examples: 5 h, 10 min, 100 s, 8 days, 10 weeks

Corrected as suggested.

9. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

The permission to reuse one manuscript figure was obtained earlier, prior to the time of the original submission.

10. Please include the limitations of the Protocol in the Discussion section.

Limitation paragraph added to Discussion as suggested.

11. Figure 1: Please use standard abbreviations for time units in the X- axis. Please revise “time, seconds “ to “time (s)”

X-axis adjusted to “time, s”, as suggested.

12. Figure 2/5: Please include scalebars in all the images of the panel and define them in the Figure legends. Consider adding the details of the magnification to make the image more informative.
Figure 2 has the fragment of the stage micrometer image embedded in the lower right corner of each image (distance between vertical bars of the pictured stage micrometer is 10 m). Frame width is 85 um, which is also marked in the figure legend. 
Figure 5 has the scale bar added; frame size and scale bar information added to the figure legend. 
 
13. Please sort the Table of Materials in alphabetical order.
Chemicals had been sorted in alphabetical order. We now, separately, sort equipment in alphabetical order, separately from chemicals.
Reviewers' comments:
Reviewer #1: 
Manuscript Summary:
The goal of this protocol is to prepare, purify and characterize gas-filled microbubbles, targeted contrast agents for ultrasound molecular imaging. The authors describe two targeting systems: biotinylated bubbles adherent to streptavidin and cyclic RGD peptide microbubbles that bind to αvβ3, a known tumor neovasculature biomarker. Very detailed description and explanation of the steps necessary for targeted microbubble preparation. In addition, methods to test bubble preparation and transfer of lipid materials.

Overall: This protocol will be very helpful with the microbubble and ultrasound community to have a well-documented protocol to making bubbles in the research lab. There are several bubble preparation or sizing bubble protocols previously published, but hard to replicate, if at all. The authors did a great job at detailing the steps throughout the steps. It is very detailed but sometimes hard to follow steps. I would encourage to go through the document one more time and make it a bit clearer in certain steps. If more sentences are needed to make it detailed, please add them. There are some parts not in the video footage. For a researcher getting started with the microbubble process for the first time, it would be great to have more portions shown in the video footage, if possible. Or even include footage you submit that could be added as supplementary. The notes documentation is great for the microbubble preparation throughout the document and for alternatives. This protocol could serve as a big stepping stone for a detailed microbubble fabrication document for many researchers.

We thank the reviewer for the positive assessment of our effort.  In the revision we have tried to address all of the suggestions to the best of our ability.

Major Concerns:
Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., "Do this," "Ensure that," etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as "could be," "should be," and "would be" throughout the Protocol.

Thank you for noting these issues; numerous amendments were made as required.

Any text that cannot be written in the imperative tense may be added as a "Note." There are many sections throughout the document that just make statements. Please elaborate these areas. Line 136 -139: this is a little hard to read, please rewrite for clear message

Text edited to achieve better clarity and state the intended goal.

Line 372-373: This needs to be cited and does normal saline irrigation solution cause the bubbles to expand when being used to characterize the bubbles in the Counter?

The phenomena that take place following introduction of fluorocarbon bubbles into air-saturated medium have been discussed in detail since last century, starting with Van Liew (PMID:8828702) and Kabalnov (PMID: 9695277) - here we just suggest to minimize this potential source of a significant variability, because dependence on the level of dissolved air and temperature history of the medium on the bubble size can be drastic. Indeed, as suggested by the reviewer, degassing Coulter diluent saline and saturating it with perfluorocarbon gas will provide a somewhat different size distribution pattern (the number and volume of largest bubbles will drop). A separate systematic study may be considered separately from this manuscript, which is already close to the allowed page limit. Normally, for particle counting, regular air-saturated saline, equilibrated at room temperature, is used: some say that it would serve as a better mimic for in vivo following intravenous administration, when compared with degassed medium. Ideally, all bubbles should be of uniform size, which is not the current real use scenario. That is why we present here a simple flotation tool to remove largest microbubbles, so they will not grow further in any medium.

Line 418: cover the biomarker protein surface with what?

Thank you for suggesting the clarification. We want to warn the user to avoid covering the target receptor surface with silicon grease. Text amended to point at that.

Line 437-442: This portion is hard to understand of what the reader is to do as an alternative approach.

Thank you for the suggestion; text reworded to clarify.

Line 583-584: citation needed for bubbles lasting years in a sealed vial.

There is no formal citation, sorry, this is our own personal experience; we have adjusted the text to remove “years”. May need to perform a separate formal study to investigate long-term stability of microbubbles, if we obtain funding.

Figures
Figure 2: Both images have white boxes in bottom right. Was the scale not properly added or was corrupted in uploading? Lines 525 - 528 state fluorescence, upload the fluorescence image to differentiate between the fluorescence and brightfield

These white boxes with black marks in them are fragments of the microscopy image of the stage micrometer where vertical black lines represent the “ruler” with the ten-micrometer “step”. Figure legend expanded to clarify. The frame on the left is a fluorescence-only image. The frame on the right is brightfield-only image (same frame, taken several seconds later).

Figure 5: It is very blurry, unless the upload disrupted the quality. Please upload a clear image of the bubbles on the petri dish and add a scale.

Dynamic range of the camera resulted in high background level observed, following conversion to pdf format, and resulting fuzziness, now corrected by background adjustment. The width of the image is 106 um (information has been added to the figure legend), size bar added.

Figure 7: In the notes, the author described that increasing the MI to 1.9 would remove the bubbles. Please add an additional image that shows another panel for removing bubbles with high MI to compare with the streptavidin coated bubbles panel.

Third imaging frame added (now on the right). Figure legend amended to include.
Lines 53-55: cite the two published examples

Amended as suggested.

Minor Concerns:
Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Thank you for making this suggestion – we did an additional check.

Please sort the Materials Table alphabetically by the name of the material.

First part of the table is Equipment (now in alphabetical order). Then we have reagents – also in alphabetical order.

There are certain portions of the document that has common acronyms without spelling it out. Please look throughout the document to catch these.

Done.

The words coulter counter and Multisizer is used interchangeably but this can be confusing for the reader.

Both words have been removed and replaced with “electrozone sensing”, which is an established generic term for coulter counting.

In certain steps, it mentions notes. Please make sure to remove the notes or commentary into note sections. Line 254: Regarding the coulter counter, notes on using the coulter counter, add that it will be discussed in section 5 Microbubble size distribution assessment.

Text added as suggested.

Line 265 - 285: For the Microbubble counter steps. Use the term Particle sizing system as mentioned in your materials document. Multisizer and coulter counter are used interchangeably.

Amended as suggested.

List all centrifugation speeds in terms of centrifugal g-force instead of rpm. Avoid the use of commercial language, including ™/®/© symbols or company brand names before/after an instrument or reagent. Cite these in the Table of Materials instead.
Two mentions of rpm are not related to centrifugation: one is the rotation of the magnetic stirrer, and another is the movement of the amalgamator. Centrifugation description (yet another procedure) does report g-force value.

We have not used ™/®/© in the initial version of the manuscript. We have removed company and brand names from the text into the Materials spreadsheet.
Spell out acronyms and use proper symbols, here are a few examples found:
Lines 56: Spell out PEG

This has been spelled out in the original version of the manuscript, line 83-84

Lines 59: Spell out NHS-PEG-DSPE

This has been spelled out in the original version of the manuscript, line 83-84

Line 75: spell out c(RDGfK) -> cyclo[Arg-Gly-Asp-D-Phe-Lys]

Done.

Line 81: write epsilon as ε

Done.

Reviewer #2: 
Manuscript Summary:
As discussed in the paper, microbubbles targeted to upregulated vascular proteins has been pursued for multiple disease states over many years. This manuscript is of interest particularly as regards targeting using peptide ligands and the attachment of these to microbubbles. The description of microbubble manufacture and purifications of these as well as the in vitro assessment is generally well written, although there are a few typos etc that need correcting.

We thank the reviewer for the positive assessment of our effort.  In the revision we have tried to address all of the suggestions to the best of our ability.

Major Concerns:
None

Minor Concerns:
1.4 could the authors add a brief description of creating the micellar solution of peptide-PEG-DSPE in saline (or biotin version)?

PEG-DSPE molecules turn into small spherical micelles upon addition of aqueous medium, by themselves, with sufficient time. Note N1.5 added for clarification.

1.5.2 needs to change to 260nm in last sentence. 
Thank you for noticing this misprint! It is now corrected.

2.1.1, states that DSPC and PEG Stearate were co-dissolved in propylene glycol (PG, 10mg/ml for both). Could the authors clarify that the PG used at a neat concentration and that the 10mg/ml refers to the DSPC and PEG stearate?

It was indeed each of these components at 10mg/ml, in neat PG. The text has been amended to clarify. 

Reviewer #3: 
Manuscript Summary:
The authors have described the process of formulating targeted microbubbles with enough detail to make the procedure accessible to those without extensive experience in microbubble fabrication. This will be a valuable resource for anyone interested in formulating targeted or non-targeted microbubbles. The authors have taken care to point out procedural precautions that may not have been performed if they were not called to the reader's attention. I wish all protocols were written with this much attention to detail and I look forward to viewing the video protocol.

We thank the reviewer for the kind assessment of the manuscript and our effort.

Major Concerns:
The authors used the uptake of DiI into the microbubble shell to monitor the efficacy of lipid transfer. They cited the loss of 85% of DiI fluorescence of the unshaken lipid solution compared to the infranatant recovered after centrifuging microbubble suspensions as evidence that 85% of the available lipid in the solution was incorporated into the microbubble shell. The authors' prior publication (#6, doi:10.1021/acs.langmuir.8b03551) used fluorescently tagged lipids to assess the efficacy of lipid transfer and saw about 70% transfer to microbubble shells. Interpreting the result from the fluorescently tagged lipid experiments is straightforward but much more expensive than using DiI dye. Interpreting the DiI dye experiment is a little more complicated since fluorescence of this dye is enhanced when it is incorporated into a lipid environment. Do the authors have additional data to validate the use of DiI dye for assessing the efficacy of lipid transfer? (My concern is that DiI fluorescence could be enhanced by lipid micelles in the lipid solution or by nano bubbles remaining in the infranatant - complicating interpretation of these results.) If it's simple to explain the validity of using DiI, please add text as appropriate. Otherwise, let the reader know that 70% lipid transfer was observed in Ref #6 and that the DiI method provided a similar measure.

Thank you for pointing at this difference. DiI (as all of the bis-C18-lipid cyanine dyes in the family, from DiO to DiR) have a moderate size fluorescence molecule, outfitted with two stearyl residues. So from physical chemistry standpoint, such as solubility in water, DiI is similar to diacyl lipid with similar C18 fatty residues. Reference 6 (now Reference 5) study describes transfer of a different molecule, not the basic phospholipid or labeled dye, but a mimic of ligand-PEG-DSPE, namely, fluorescein-PEG5000-DSPE, which has a somewhat higher critical micellar concentration (i.e., solubility of individual molecules in water) than DiI or phospholipids. Therefore, it is not surprising that PEG-lipids will transfer to the shell to a somewhat lower extent. The validity of fluorescence measurement performed with DiI in this study had been achieved by the selection of the fluorescence measurement medium, as outlined in Section 3.3 (lines 219-223 of the first version of the manuscript): PBS:ethanol 1:1, with 1% Triton X-100. In these conditions, with large excess of Triton X-100 and presence of ethanol, all lipid molecules are disassociated, all membrane structures or potential nanobubbles are destroyed, so the dye molecules cannot be subjected to potential concentration quenching anymore. All measurements, prior to or after the separation of bubbles from the preparation medium are made in these conditions.

Minor Concerns:
1) In the description of the parallel-plate flow apparatus, please include the name of the manufacturer (Glycotech) otherwise a reader may expect more technical details to be included so that they can reproduce this apparatus.

Thank you for suggesting this clarification. As the author instructions require, the name of the parallel plate plug manufacturer could not be present in the text. It is present in the Table of Materials.

. 
2) In Figure 6, please provide complete units for the Y-axis. Was this microbubbles per mm^2, per video frame, per 35mm plate?

Figure 6  Y-scale represents the number of particles in the field of view following four minutes of accumulation time. Figure legend is now slightly amended to clarify.