Detailed responses and changes to

“A robust method for packing high resolution C18 RP-nano-HPLC columns”

By

Yasaman Jami-Alahmadi, Vijaya Pandey, Adarsh K Mayank, James A Wohlschlegel

First of all, we thank the reviewers for their very thorough review of our paper.  We strongly believe that by taking into account their comments the paper is much stronger.

Editorial comments:

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

We have addressed this issue in the revised manuscript.

2. Please provide email addresses of all authors.

We have addressed this issue in the revised manuscript.

3. For in-text citations, use “… research1-4.” Instead of “…research.1-4 ” (line 35, 38, 45, etc.)

We have addressed this issue in the revised manuscript.

4. Please adjust the numbering of the Protocol to follow the JoVE Instructions for Authors. For example, 1 should be followed by 1.1 and then 1.1.1 and 1.1.2 if necessary.

We have addressed this issue in the revised manuscript.

5. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.”. E.g. Line 92-94: Use ”Place a hot soldering iron along the tip to initiate the polymerization of the frit solution…” instead.

We have addressed this issue in the revised manuscript.

6. Use appropriate symbols and abbreviations of SI units. E.g. use “µL / µm” instead of “uL/um” (line 80,84, 112-113, etc.), “nL” instead of “nl” (line 172). Add a single space between quantity and its unit e.g. “1.7-3.5 mm” instead of “1.7-3.5mm” (line 45), “1300 psi” instead of “1300psi” (line 128), “275 oC” instead of “275oC” (line 174). Check the same for the Table of Materials as well.

We fixed this.

7.  Line 68: Please add appropriate units for the settings of heat value, velocity and delay.

We fixed this.

8. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials. E.g. Ploymicro technology, Kimwipe, Sutter, Kasil, Eppendorf, Orbitrap Fusion Lumos etc. Also mask the manufacturer/model name on instruments used (e.g. Figure 5a).

We agree with the editor’s comments about removing commercial language from the manuscript. However, one of the reviewers has suggested sharing in the protocol, the commercial source of empty fused silica capillary with integrated ESI emitter. These capillaries have not been used in our protocol, so we are not adding them to the table of materials. In this situation, where do you recommend, we mention the commercial source?

9. Please highlight up to 3 pages of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol. Remember that non-highlighted Protocol steps will remain in the manuscript, and therefore will still be available to the reader.

We highlighted the essential steps of the protocol for the video.

10. Please include a title and a description of each figure and/or table. All figures and/or tables showing data must include measurement definitions, scale bars, and error bars (if applicable). Please include all the Figure Legends together at the end of the Representative Results in the manuscript text.

We fixed this.

11. Please include a scale bar for all images taken with a microscope to provide context to the magnification used. Define the scale in the appropriate Figure Legend. E.g. Figure 3

We inserted dimension of the column to provide context to the magnification in Figure 3.

12. Figure 7: Please adjust the Y-axis length. Include appropriate axis labels and shift the legend (it is covering the bar).

We fixed this.

13. Please include a Disclosures section, providing information regarding the authors’ competing financial interests or other conflicts of interest. If authors have no competing financial interests, then a statement indicating no competing financial interests must be included

We fixed this.

14. Please ensure that the references appear as the following: [Lastname, F.I., LastName, F.I., LastName, F.I. Article Title. Source. Volume (Issue), FirstPage – LastPage (YEAR).] For more than 6 authors, list only the first author then et al. Do not abbreviate the journal name. Do not use “& / and” in the author list in the references. Please include volume and issue numbers for all references.

We have addressed this issue in the revised manuscript.

15. Please sort the Materials Table alphabetically by the name of the material.

We have addressed this issue in the revised manuscript.

Reviewers' comments:

Reviewer #1:

Major Concerns:

One of the major concerns in this article is that the authors have entirely disregarded the recent fundamental studies done in the field of column packing science and yet claim their method as "state of the art" in the abstract. It may not be fair to call the method outlined here "state of the art" because authors are perhaps unaware of recent contributions in this field of column packing science.

We agree with the reviewer’s comment that the “state of the art” claims made in the protocol are overstated (not to mention not really our intention) and have revised the text accordingly. Importantly, the revised protocol attempts to better emphasize the true goal of the protocol which is to provide an efficient approach for preparing reverse-phase LC columns that offer a strong blend of performance and durability in proteomic applications at a fraction of the cost. This protocol has been routinely and reproducibly employed in our laboratory to publish >100 proteomics-based publications and we believe that this protocol can be easily replicated in laboratories equipped with basic supplies for proteomic applications 1-10.

This reviewer recognizes the importance of well-packed capillaries, but have you considered what others have investigated and what were their findings? Those findings are certainly relevant to this work because the whole paper revolves around developing a protocol for packing a good capillary.

We apologize for largely ignoring the fundamental studies in the field of column packing science.  We agree that the literature suggested by the reviewer is extremely relevant for putting into perspective the column packing approach described in the manuscript. We have revised the introduction to reflect the importance of previously published work and have incorporated the corresponding references.

I suggest the authors go through the following references and notice the importance of optical microscopy of suspensions before packing their capillaries. How did they choose the solvent? Is that solvent choice universal? More importantly, does that capillary represent the highest efficiency?

We agree with reviewer’s comments about the significance of optical microscopy in selecting the most appropriate solvent for slurry packing. We also appreciate that when dealing with slurry particles of sub- 2mm size range, as is often the case in modern proteomics, the adhesion forces between the particles are strong enough to necessitate a stable suspension before packing can commence. Hence it is critical to experimentally determine solvent or mixture of solvents for efficient suspension of the slurry. Although we did not perform optical microscopy and sedimentation velocity measurements to study the behavior of the slurry particles in solvent, our choice of methanol as the slurry packing solvent was informed by previously published literature11-15 and also subsequently empirically determined in our laboratory.  Methanol is a commonly used and frequently reported solvent for sample preparation and chromatography for its ready availability and low costs. As suggested by the reviewer we have noted the importance of solvent-particle interaction in the introduction and cited references where these interactions have been discussed extensively.

2) In any packing protocol, the design of the packing hardware is also critical. Please show the inside of the packing chamber and schematically show how helium pressurizes the slurry in this particular chamber design. Also, show the nature of connections from the helium tank to the packing chamber.

We agree with the reviewer and have added the schematic as represented in Figure 5.

3) The authors must include a quality control test that shows the packed capillary chromatographic efficiency. For that, the authors must include a quality control peak efficiency test using a compound of their choice after packing a capillary. How would the user know that the efficiency of their packed capillary is at the optimum? Again, please see the flow chart in the above references.

We agree that it is important to assess the quality of packed columns for proteomic applications. Based on our experience, the number of peptide identifications resulting from a standardized LC-MS/MS analysis (e.g. 750 ng of HEK293 cell tryptic digest in a 70-minute water-acetonitrile gradient) is a very sensitive measure of column quality and one that we use in conjunction with classic properties such as retention time reproducibility and peak width. Importantly, we track these metrics longitudinally to ensure reproducibility over time.  Fig. 8 shows the consistency of our approach for columns packed at different points in time.  It is also worth noting that using complex peptide digests to benchmark LC-MS/MS setups is common in the field and the columns described here provide results that are comparable to previously published reports using either commercial or in-house packed columns based on similar column length, slurry, solvent and MS parameters similar to the current study12,16,17.

4) The word nano repeatedly appears like a buzzword in the introduction. Please reduce its usage and define the dimensions of the columns instead of using nanoLC. These are just commercial names, and hence should be avoided.

We have revised the text as suggested by the reviewer.

6) In the protocol, point no. 3. "heat value of 300, velocity of 10 and delay of 180" write the units (these numbers are meaningless without their units).

We have added the proper units to the text.

7) In Step 3, under the heading of "Preparation of Stationary Phase,": Please do not call C18 silica as a resin until and unless polymers are being used. C18 silica is a fully porous silica particle, not a resin. What is the particle size?

We agree with the reviewer and have revised the text accordingly.

8) A good practice in writing is to write the figure caption as a stand-alone statement and detail everything so that the readers do not have to search the main text. There might be some repetition, but it is better to repeat in a caption. For example, in Figure 4, explain what HF-eater is and what is being done there.

We have revised the figure captions based on the reviewer’s suggestions.

Minor Concerns:

1) Add a cautionary note in "column tip is immersed in a 50% hydrofluoric acid (HF)".

We have added the cautionary note per the reviewer’s suggestion.

2) Correct reference 10.

We have corrected the reference.

Reviewer #2:

Manuscript Summary:

The abstract reads well. However, I would advise to replace the statement: "the size and nature of stationary phase particles" to a more specific and clearer statement: "the particle size, pore size and chemistry of stationary phase particles"

We have corrected this error.

Major Concerns:

Protocol

1. Preparation of capillary tip

Protocol Point 2 When talking about the bare-fused silica capillary from (Polymicro Technology) it should be clarified that the capillary is "coated with Polyimide buffer" as documented in the manufacturer's website. It is important to state this as this is the justification for the procedure in point 2 of the protocol: removing the polyimide coating from the capillary along 10 cm in the middle. The authors should also explain the purpose of wanting to have these 10 cm of bare-fused silica capillary with clear glass windows and where they intend to locate this 10 cm segment.

We have revised the text based on the reviewer’s suggestion.

Protocol Point 3 The units should be included to the value in this list: heat value of 300, velocity of 10 and delay of 180,

We have revised the manuscript to indicate the units.

Question: are the authors talking about 1-5 mm or 1-5 µm diameter tip? Take care with the units.

We have revised the manuscript to make sure the units are correct. µm is the correct unit for the diameter tip as shown in figure 3.

Related figure 3: A size bar should be shown in Fig. 3 as to get a clear idea about the difference between the tip of the capillary and the main ID of the capillary.

We agree with the reviewer and have revised Fig. 3A accordingly

2. Polymerization/etching of the tip

Protocol Point 1: These instructions are misleading: "the emitter is blocked with a porous frit comprised". It reads as the composition of this solid/gel frit is composed by these chemicals. Instead, the porous frit is prepared from these chemicals as starting materials a suspension or gel which later need to be allowed to polymerise.

We apologize for the confusion and have clarified this point in the revised manuscript.

This Note can be misleading: For example, one can use 100uL of Kasil 1, 300 uL of Kasil 1624 and 100uL of formamide acid to get the polymerized frit solution.

Do the authors mean here:" …and 100uL of formamide acid to ensure the frit solution will be able to polymerize and form the frit with the required porosity"?

We have clarified the description of this step of the protocol in the revised manuscript.

Protocol Point 2: This sentence is unclear: "Immerse the laser-pulled capillary tip in the clear solution of the frit (not the precipitate)". Do the authors mean: "in the clear mother liquor of the frit suspension (not the precipitate)

We have revised the text based on the reviewer’s suggestion.

Protocol Point 4: "To account for the possibility" may be more accurately described as "to prevent"

We have revised the text accordingly.

The authors should add 3 extra pictures of the nano-column tip: 1) before etching with HF; 2) after etching with HF and 3) after washing the HF from the capillary tip with water.

Figure 3B shows the column tip before etching. However, taking an image of the column after HF-etching and before washing it, does not align with the safety protocols followed in the lab. Since HF is a highly corrosive chemical, we strictly avoid contact with etched tip unless it has been thoroughly washed first with HF-neutralizer and then with copious amounts of water. In our experience, the emitter tip before and after etching looks very similar such that an image of post-etching tip provide little useful information.  We have revised the text to make this point.

Under notes, No 3: the authors could suggest from which suppliers could "the empty fused silica capillary columns with an integrated electrospray emitter" be purchased.

We have included the requested information.

3. Preparation of stationary phase

Protocol Point 1: C18 reversed phase is not a resin as it is covalently bonded to silica. It is a C18 stationary phase. The term resin can only be used when the C18 is bonded to an organic polymer. This is not the case here.

We thank the reviewer for his/her insight and have revised the text accordingly.

C18 stationary phase material that the authors are using to pack the nanocolumn is not a straightforward C18; they are using the "Reprosil-Pur 120 C18-AQ". The authors should learn from the manufacturers more about the chemistry of this stationary phase. I am disappointed with the lack of sensible information on the chemistry and applications of this column in dr-maisch.com website. AQ typically stands with most column manufacturers for aqueous, and these C18 AQ phases have been designed with a polar group (amide, carbamate, ether, other) embedded in the C18 strand so as to prevent the stationary phase from collapsing when using 100% aqueous mobile phase. Stationary phase collapse means that the alkyl strands in the presence of highly aqueous phase collapse away from the highly aqueous mobile phase just as oil droplets will collapse in water. The effect of phase collapsing will result over time in shorter retention times for the analytes.

The phrase: Resuspend ~ 0.05 g of is wrongly used. Reprosil-Pur 120 C18-AQ is sold as a solid powder, not as a suspension. So, it would be less confusing to the reader if the instruction is written as ": Suspend ~ 0.05 g. The authors should not use this rough quantity (0.05g) but should instead use 50 mg. It is important for the column to column reproducibility how the concentration of the suspension/slurry of the stationary phase is prepared. The authors should give more careful advice.

We agree with and value the points being made by the reviewer. We have now corrected the text wherever possible to reflect the correct nomenclature of the particles being used as the stationary phase. As suggested by the reviewer, we have also accurately stated the weight of Reprosil-Pur 120 C18-AQ particles used for preparing the stationary phase suspension.

The authors use MeOH to prepare the slurry of the C18 Aq. stationary phase. From my own experience, the solvent for packing is carefully selected to give an ideal slurry preventing the particles from sedimenting or agglomerating. Different solvents are more suitably for different stationary phases.

We agree with reviewer’s comments about the significance of meticulous selection of appropriate solvent for slurry packing. We also appreciate that when dealing with slurry particles of sub- 2mm size range, as is often the case in modern proteomics, the adhesion forces between the particles are strong enough to necessitate a stable suspension before packing can commence. Hence it is critical to experimentally determine solvent or mixture of solvents for efficient suspension of the slurry. Although we did not perform optical microscopy and sedimentation velocity measurements to study the agglomeration behavior of the slurry particles in solvent, our choice of methanol as the slurry packing solvent was informed by previously published literature11-15 and also subsequently empirically determined in our laboratory.  In the same note, we have also stated the importance of solvent-particle interaction in the introduction and have cited references where these interactions have been discussed extensively.

The authors should report with accuracy the details of their stationary phase: "C18 reversed phase resin with 1.9 um diameter and 100 Å pore". Two points:

1) 1.9 µm is the particle size which is most likely spherical but it should not be referred as 1.9 µm diameter; dr-maisch.com website claims particle size= 1.9 µm;

We have revised the text accordingly.

2) the details of the pore size are incorporated into the name Reprosil-Pur 120 C18-AQ; looking at dr-maisch.com, this stationary phase is reported as having 120 Å pore size.

We thank the reviewer for noting this issue and have corrected it in the revised manuscript.

Protocol Point 2: When describing the HPLC packing bomb, the authors should explain to the reader the marked difference between the operation of HPLC pumps and HPLC packing bombs. HPLC pumps gives very accurate constant flowrates in order to reproduce the retention time of analytes. On the other hand, HPLC packing bombs have to be operated at constant pressure to avoid disrupting the HPLC slurry during packing.

We have revised the text based on the reviewer’s suggestions.

I am surprised that the packing Bomb is operated with helium as this gas is expensive and not very sustainable. Have they authors checked if the packing pump could be also operated with other gases like nitrogen? This should be discussed.

Although we appreciate the reviewer’s concern, we and others have been routinely using this gas for the purpose of capillary column packing 4,5. In our experience, balloon grade helium gas is readily available in our lab. Although we have not tested the suitability of nitrogen gas in our system, we expect nitrogen gas is equally suitable for this purpose.  Indeed, Kovalchuk et al have used nitrogen gas in their protocol for packing columns with dimensions fairly similar to ours.2

Fig 5a which is relevant for this section, should indicate with arrows how the 3-way valve of the packing bomb is to be used for packing the nano-LC column.

We have added a schematic representation (Figure 5) of the column packing system where we have indicated the direction of helium gas flow inside the pressure bomb system.

4. Packing the column with stationary phase

Protocol Point 1

The specific plastic fitting shown as a white fitting on relevant Figures 5, 6 and 7 is intended to seal the end of nano- LC column to the female fitting at the outlet of the packing bomb to prevent leaking and loss of pressure during packing. This HPLC fitting is probably a PEEK One-Piece Fingertight fitting required for sealing 360µm OD tubing. This must be a very specific fitting and should be documented under the section "Name of Material and Equipment". Additionally, a picture showing the end of the nanocolumn being threatened through the One-Piece Fingertight fitting should be shown making sure that there is a clear view of One-Piece Fingertight fitting.

We have added an image as Figure 5 which shows the column being threaded through the one-piece fingertight fitting. We have also documented the information related to this fitting in the "Name of Material and Equipment" table.

Protocol Point 2 The term "slurry" is correctly used in this context and should be consistently used over the protocol.

Per the reviewer’s suggestion, we have corrected the usage of the term ‘slurry’ in the text where relevant.

Protocol Point 3 This statement requires an explanation "If the formation of liquid droplet on the column tip is delayed, quickly flame the tip to ensure it is open."

In my mind there are at least 2 reasons why the tip could prevent MeOH from coming out. 1) either the frit pores were much too large and particles are coming through and blocking it or 2) the stationary phase has too many fine particles which are coming through a well prepare frit with the right pore size range and clogging the tip. Modern HPLC stationary phases do not tend to have fines. Even if these problems were to be resolved by flaming the tip, a higher backpressure will be obtained through a blockage of the frit or continuous blockage of the capillary end through the stationary phase fines.

Based on our experience, 5-10% of the columns packed using our approach have blocked tips that prevent the slurry from flowing through the length of the capillary column. This may happen either for the reasons cited by the reviewer or due to sub-optimal pulling of the tip by the laser puller. In this scenario, a quick flaming of the tip often releases the blockage by pushing it out enabling the packing procedure to resume. Columns rescued by this technique perform comparably to ones packed normally.

Point 4: The authors do not discuss the possibility of getting He bubbles into the packed column, if packing had to be stopped and restarted due to blockage. The authors should discuss strategies to avoid He bubbles inside the nano LC column. Any disruption of good packing, like gaps due to air bubbles, inside a typical HPLC column results in peak splitting of the eluted peaks.

We appreciate the concerns raised by the reviewer related to the possibility of He bubbles inside the column and have addressed them in the section 5. As mentioned in section 5.2, we recommend packing columns longer than the desired length. In the same note, we also advise the users to continue packing beyond the visible length of the column above the one-piece finger-tight fitting for the reason that there is 5-10 cm of column length which is inside the bomb and not visible to eyes (see the schematic). Additionally, we recommend allowing the packing apparatus to slowly self-depressurize after the helium flow has been turned off, in order to ensure that the particles remain uniformly packed throughout the length of the column, avoiding the introduction of air bubbles.

The authors talk about tightening the ferrule; this is confusing as in the picture it is clear that the authors are using a One-Piece Fingertight fitting where the ferrule is an integral part of this fitting.

We have revised the text to address the reviewer’s concerns.

5. Finishing up the column and making the back-frit.

The authors jumped by mistake from Protocol Point 3 to 5, bypassing Point 4

We thank the reviewer for noticing this error which we have corrected in the revised manuscript.

Protocol Point 4: The way the instructions have been written here implies that the column end is cut with stationary phase flash with the end. If this end is dipped into the frit solution, then the polymerisation will trap stationary phase probably blocking many of the pores of this frit. This should be better explained.

We apologize for the confusion and have edited the text in section 5 to make the instructions clearer.

Representative Results

Here the authors should clearly state the starting composition of the mobile phase gradient. Is it probably implied that the gradient starts at 0% Buffer B? This is not stated. If so, it is important to state the equilibration time required to come from 80% buffer B to 0% buffer B to ensure the chromatograms are reproducible.

We have revised the manuscript to address the reviewer’s concerns.

I am not familiar with this LCMS term: 120,000 "resolution". Maybe the authors could make it clearer for the readers. The supplier, quality and preferred storage conditions of DMSO should be documented.

In mass spectrometry, resolution refers to the measure of ability of instrument to distinguish two peaks or peptide ions with slight difference in their mass-to-charge ratio. We have described the resolution of the MS and MS/MS scans in results section using standard mass spectrometry nomenclature.

Per the reviewer’s suggestion, we have also added the supplier and quality of DMSO to the table of materials. Storage conditions have also been added to the materials table section.

Title of Figure 8

The title inside the figure is unclear and confusing: It should say: "Example of total number of identified proteins and peptides"; It may be necessary to reduce the font of this title

We have revised the title of this figure to make it clearer.

Minor Concerns:

The titles of the selected references cited here are relevant and up to date. The chromatography examples presented in Figure 8 are as reproducible as can be expected from any HPLC analysis.

Title: the concept of nano- such as C18 RP-nano-HPLC column is important and should be incorporated into the title.

We have renamed the title of the manuscript in accordance with the reviewer’s suggestion.

Abstract: The abstract reads well. However, I would advise to replace the statement: "the size and nature of stationary phase particles" to a more specific and clearer statement: "the particle size, pore size and chemistry of stationary phase particles"

We have revised the manuscript in accordance with the the reviewer’s suggestion.

Keywords: If only 10 keywords are allowed, I would then recommend to replace "mobile phase" which does not add value here for "nanoflow chromatography" or "nanoLC". The term "frit" is vague. "nanoLC column frit" is more specific, as this type of frit can be only used for nano LC columns.

We have implemented the reviewer’s suggestion.

Correct use of terminology over the document: Better to use nanoLC column than LC column; uL or mL are incorrect; should be µL; um or mm are incorrect; should be µm; please check all the occurrences of these units in the document.

We apologize for the oversight on our part and have corrected the spellings and units throughout the text.

Introduction: Same advice for keywords. The "gradient" should be referred more correctly as "solvent gradient".

We have made necessary changes in the introduction to address the reviewer’s concerns.

References:

- Stehling, O., et al. MMS19 Assembles Iron-Sulfur Proteins Required for DNA Metabolism and Genomic Integrity. Science. 337, (6091), 195–199 (2012).Blue, L. E., Jorgenson, J. W. 1.1 μm Superficially porous particles for liquid chromatography: Part II: Column packing and chromatographic performance. Journal of Chromatography A. 1380, 71-80 (2015)

- Mayank, A. K., et al. An Oxygen-Dependent Interaction between FBXL5 and the CIA-Targeting Complex Regulates Iron Homeostasis. Molecular cell. 75, (2), 282-393 (2019).

- Nie, M., Oravcová, M., Jami-Alahmadi, Y., Wohlschlegel, J. W., Lazzerini‐Denchi, E., Boddy, M. N. FAM111A induces nuclear dysfunction in disease and viral restriction. European Molecular Biology Organization. 22, (2), e50803 (2021).

- Vélez-Ramírez, D. E., et al. APEX2 Proximity Proteomics Resolves Flagellum Subdomains and Identifies Flagellum Tip-Specific Proteins in Trypanosoma brucei. mSphere. 6, (1) e01090 (2021)

- Wang, Y., et al. NAP1-RELATED PROTEIN1 and 2 negatively regulate H2A.Z abundance in chromatin in Arabidopsis. Nature Communications. 11, (1) 2887 (2020)

- Nadipuram, S. M, Thind, A. C., Rayatpisheh, S., Wohlschlegel, J. A., Bradley, P. J. Proximity biotinylation reveals novel secreted dense granule proteins of Toxoplasma gondii bradyzoites. PLoS One. 6(15) e0232552 (2020)

- Grolla, A. A., et al.  A nicotinamide phosphoribosyltransferase-GAPDH interaction sustains the stress-induced NMN/NAD + salvage pathway in the nucleus. Journal of Biological Chemistry. 295 (11) 3635-3651 (2020)

- Wongpalee, S. P., et al.  CryoEM structures of Arabidopsis DDR complexes involved in RNA-directed DNA methylation. Nature Communications. 10(1) 3916 (2019)

- Mota, C. M., et al. Corrigendum: Neospora Caninum Activates p38 MAPK as an Evasion Mechanism against Innate Immunity. Frontiers in Microbiology. 10(548), 1456 (2019)

- Barshop, W. D., Rayatpisheh, S., Kim, H. J., Wohlschlegel, J. A. Sequential Windowed Acquisition of Reporter Masses for Quantitation-First Proteomics. Journal of Proteome Research. 18,(4), 1893-1901 (2019)

- Kovalchuk, S. I., Jensen, O. N., Rogowska-Wrzesinska, A., FlashPack. Fast and Simple Preparation of Ultrahigh-performance Capillary Columns for LC-MS. Molecular and cellular proteomics. 18, (2), 383-390 (2019).

- Annesley, T. M Methanol-Associated Matrix Effects in Electrospray Ionization Tandem Mass Spectrometry. Clinical Chemistry, 53, (10) 1827–1834 (2007)

- Moore, S. M., Wu, C. C. Proteomic Characterization of Integral Membrane Proteins Using Thermostatted Liquid Chromatography Coupled with Tandem Mass Spectrometry (Methods and Protocols), 914. Humana Press, Totowa, NJ. https://doi.org/10.1007/978-1-62703-023-6_9

- Ishihama, Y., Rappsilber, J., Andersen, J. S., Mann, M. Microcolumns with self-assembled particle frits for proteomics. Journal of Chromatography A. 979, (1-2):233-239 (2002)

- Newton, R. P., Brenton, A. G., Smith, C. J., Dudley, E., Plant proteome analysis by mass spectrometry: principles, problems, pitfalls and recent developments. Phytochemistry, 65, (11) 1 1449-1485 (2004)

- Peng, J., Elias, J. E., Thoreen, C. C., Licklider, L. J., Gygi, S. P. Evaluation of Multidimensional Chromatography Coupled with Tandem Mass Spectrometry (LC/LC−MS/MS) for Large-Scale Protein Analysis:  The Yeast Proteome. Journal of Proteome Research, 2, (1) 43–50 (2003)