# Submission ID #:  62383

# Scriptwriter Name: Madhulika Pathak

# Supervisor Name: Anastasia Gomez

# Project Page Link: https://www.jove.com/account/file-uploader?src=19026863

# Title:   Strategies for Optimization of Cryogenic Electron Tomography Data Acquisition

# Authors and Affiliations:

Felix Weis1, Wim J. H. Hagen1, Martin Schorb2, Simone Mattei1,3

1Structural and Computational Biology Unit, European Molecular Biology Laboratory, Heidelberg, Germany

2Electron Microscopy Core Facility, European Molecular Biology Laboratory, Heidelberg, Germany

3Imaging Centre, European Molecular Biology Laboratory, Heidelberg, Germany

# Corresponding Authors:

Simone Mattei		(simone.mattei@embl.de)

# Email Addresses for All Authors:

felix.weis@embl.de

wim.hagen@embl.de

martin.schorb@embl.de

simone.mattei@embl.de

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

3. Interview statements: Considering the COVID-19-imposed mask-wearing and social distancing recommendations, which interview statement filming option is the most appropriate for your group?

☒ 	Interview Statements are read by JoVE’s voiceover talent.

4. Filming location: Will the filming need to take place in multiple locations?   No

Current Protocol Length

Number of Steps:  30

Number of Shots:  37

# Introduction

- Introductory Interview Statements

REQUIRED:

  - The set-up of a large-scale data collection session is a time-consuming procedure that can considerably reduce the remaining microscope time available for tilt series acquisition.

    - 8.4.3.

  - This method enables fast setup of large-scale automated tomography data acquisition to maximize the time-efficiency of tomography experiments that require a careful selection of acquisition targets.

    - 6.4.1

# Protocol

- Selecting the Targets

  - Begin by opening [1] a dummy SerialEM instance [2].

    - WIDE: Establishing shot of talent at the computer. Videographer: Obtain a few shots of talent clicking the mouse and typing on keyboard to use as b-roll throughout the video.

    - SCREEN: 62383_2.1.2_mp4 00:00:00 – 00:00:14

  - Once the first grid square is mapped, use the SerialEM Navigator menu option Merge to see the montage in the dummy SerialEM instance [1].

    - SCREEN: 62383_2.2.1_mp4 00:00:00 – 00:00:13

  - Double click on the Navigator window to open the grid square map [1].

    - SCREEN: 62383_2.3.1_mp4 00:00:00 – 00:00:15

  - Search the map and use the dummy SerialEM Navigator option Add Points to add image acquisition points on the target of interest [1].

    - SCREEN: 62383_2.4.1_mp4 00:00:00 – 00:00:11

  - After mapping the new squares, save the Navigator file and merge the Navigator again. Continue until all grid squares are mapped [1].

    - SCREEN: 62383_2.5.1_mp4 00:00:00 – 00:00:29

- Generate Virtual Maps

  - Again, merge the navigator file with the dummy SerialEM instance [1].

    - SCREEN: 62383_3.1.1_mp4 00:00:00 – 00:00:15:

  - Run the PyEM Virtual maps script from the dummy SerialEM menu, Select Tools and select Virtual Anchor Maps. This may take some time depending on the size and amount of the grid square maps, as well as the binning of the View and Preview maps [1].

    - SCREEN: 62383_2.5.1_mp4 00:00:00 – 00:01:34

- Microscope Tuning

  - To ensure proper microscope performance, use the same magnification and beam size setup for data acquisition, in the following order [1]. Run SeriaLEM Coma-free alignment by CTF [2].

    - SCREEN: 62383_4.1.1_mp4 00:00:00 – 00:00:27

    - SCREEN: 62383_4.1.2_mp4 00:00:00 – 00:00:22

  - Insert and center an objective aperture [1]. Run SerialEM Correct Astigmatism by CTF [2] and GIF Quick Tune [3].

    - SCREEN: 62383_4.2.1_mp4 00:00:00 – 00:00:20

    - SCREEN: 62383_4.2.2_mp4 00:00:00 – 00:00:20

    - SCREEN: 62383_4.2.3_mp4 00:00:28 – 00:00:30, 00:02:46 – 00:02:50.

- Set up Navigator

  - In SerialEM, open the new navigator file. Deselect all A points, select the first View maps, select Collapse, click on A twice, and deselect Collapse [1].

    - SCREEN: 62383_5.1.1_mp4 00:00:00 – 00:00:18

  - Select the first view map position, press Shift plus T, then select the very last view map position, and press Shift plus T again [1].

    - SCREEN: 62383_5.2.x_mp4 00:00:00 – 00:00:19

  - Choose Single Frame Images in the properties of the file to open a dialog [1].

    - SCREEN: 62383_5.2.x_mp4 00:00:19 – 00:00:23

  - In the next File Properties dialog, select the desired parameters according to the imaging needs and the instrument setup [1].

    - SCREEN: 62383_5.2.x_mp4 00:00:23 – 00:00:29

  - When prompted, give a name with a number and click on Save [1].

    - SCREEN: 62383_5.2.x_mp4 00:00:29 – 00:00:39

  - Set up the tilt series controller for the first TS position. When done, click OK to set these parameters for all tilt series after this acquisition item. All Preview maps are now selected as TS with a numbered file name [1].

    - SCREEN: 62383_5.2.x_mp4 00:00:39 – 00:01:11

- Set the Focus/Track Positions

  - If needed, set the focus/track distance for each target [1].

    - SCREEN: 62383_6.x.x_mp4 00:00:00 – 00:00:05

  - Double click on the View Map to load it, select the Preview Map in the Navigator list, then select Edit Focus in the Navigator window [1].

    - SCREEN: 62383_6.x.x_mp4 00:00:00 – 00:00:05

  - In the Low Dose Control panel, deselect Rotate inter-area axis to position Trial and Focus along the stage tilt axis [1].

    - SCREEN: 62383_6.x.x_mp4 00:00:05 – 00:00:12

  - Click on the desired region in the loaded view map to set the focus-trial position for this tilt series and ensure that the Navigator item has TSP set. Repeat the procedure for all items [1].

    - SCREEN: 62383_6.x.x_mp4 00:00:12 – 00:00:30

- Set up Additional Scripts

  - Edit the focus range in the script pretomo, which runs before each tilt series [1].

    - SCREEN: 62383_7.1.1_mp4 00:00:00 – 00:00:23

  - In the SerialEM Tilt Series menu, check Run Script in TS and select the script number of the duringtomo script during each tilt [1].

    - SCREEN: 62383_7.2.1_mp4 00:00:00 – 00:00:18

- Run

  - Check the nitrogen tank level [1] and whether the Autoloader turbo off is selected [2].

    - Talent checking the nitrogen tank level.

    - SCREEN: 62383_8.1.2_mp4 00:00:00 – 00:00:10

  - Check the data storage free space. In the SerialEM file menu, deselect continuous saving for the log file. Each tilt series will get its own log file [1].

    - SCREEN: 62383_8.1.2_mp4 00:00:00 – 00:00:14

  - After clicking on Acquire at Items in the Navigation menu, run the script PreTomo [1]. Select Primary task Acquire tilt series [2], then select Run script after PostTomo [3].

    - SCREEN: 62383_8.x.x_mp4 00:00:00 – 00:00:07

    - SCREEN: 62383_8.x.x_mp4 00:00:07 – 00:00:08

    - SCREEN: 62383_8.x.x_mp4 00:00:08 – 00:00:09

  - Select Close column valves at end [1] and send email at end, then click on GO [1].

    - SCREEN: 62383_8.x.x_mp4 00:00:09 – 00:00:17

# Results

- Results: Optimization of Cryogenic Electron Tomography Data Acquisition

  - Seventy-one suitable squares were selected on the grid map at lower magnification. Medium magnification maps were acquired with settings that allow for the direct visualization and identification of the sample of interest, coronaviruses in this case. The acquisition time was 3 minutes per square, 3 hours and 45 minutes in total [1].

    - LAB MEDIA: Figure 1A

  - As the first square map was created, a dummy SerialEM instance was opened on a separate computer to visualize the square map and to add points on targets suitable for tilt series acquisition [1].

    - LAB MEDIA: Figure 1B

  - SerialEM low-dose was set up and reference view and preview images were taken and saved as maps. The latter maps could then be used immediately on the dummy SerialEM instance to generate the virtual view and virtual preview from the corresponding square map images [1].

    - LAB MEDIA: Figure 1C,1D

  - The virtual view maps were used for an initial centering of the target [1] followed by a final centering performed at the actual tilt series acquisition magnification [2] using the virtual preview map [3].

    - LAB MEDIA: Figure 1C,1E

    - LAB MEDIA: Figure 1F

    - LAB MEDIA: Figure 1D

# Conclusion

- Conclusion Interview Statements

  - Users should allow themselves plenty of time to explore the sample landscape. After having searched some maps, one develops a better feel for good quality areas to collect data from.

    - 2.4.1