Response Letter to Editorial Team and Reviewers

Editorial comments:
Changes to be made by the Author(s):
1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Re: We appreciate the editorial comment. We have checked and revised the manuscript to avoid spelling or grammar issues.

2. Please provide an email address for each author.

Re: We appreciate the editorial comment. We have added the email address for each author.

3. Please rephrase the Short Abstract/Summary to clearly describe the protocol and its applications in complete sentences between 10-50 words: “This protocol presents…”

Re: We appreciate the editorial comment. We have rephrased the Summary as following:

“This protocol presents rapid antimicrobial susceptibility testing (AST) within 2.5 h by single-cell stimulated Raman scattering imaging of D2O metabolism. This method is applicable to bacteria in urine or whole blood environment, which is transformative for rapid single-cell phenotypic AST in clinic.”

4. Please ensure that the Abstract is within 150-300-word limit.

Re: We appreciate the editorial comment. We have confirmed the Abstract is in 151 words.

5. Please revise the following lines to avoid overlap with previously published work: 30-31, 49-51, 54-56, 60-62, 75-76, 85-88, 160-163, 167-170, 199-200, 375-377, 386-388, 405-408,

Re: We appreciate the editorial comment. We have modified these sentences to avoid overlap.

6. Please include an ethics statement before the numbered protocol steps, indicating that the protocol follows the guidelines of your institution’s human research ethics committee since human urine and blood is used in the study?

Re: We appreciate the editorial comment. This protocol doesn’t include an ethics statement because the human urine and blood are deidentified samples.

7. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.”

Re: We appreciate the editorial comment. We have revised the protocol section to avoid usage of phrases such as “could be,” “should be,” and “would be”.

8. Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed? This can be done by including mechanical actions, button clicks in the software, knob turns, command lines, etc.

Re: We appreciate the editorial comment. We have modified part 5, 6, 7 to including mechanical actions and button clicks in the software.

9. 1.4 What bacterial strain is used?

Re: We appreciate the editorial comment. We added more information in the text: Escherichia coli BW25113 or Pseudomonas aeruginosa ATCC 47085.

10. 2.5: Which antibiotic is used in this experiment? Concentration used?

Re: We appreciate the editorial comment. We added more information and clarified in part 1.1 and part 2.4:

  - Prepare the antibiotics (gentamicin sulfate or amoxicillin) stock solution at a concentration of 10 mg/mL dissolved in 1x phosphate-buffered saline (PBS) or dimethyl sulfoxide (DMSO) solvent in 1.5 mL micro tubes. Gentamicin sulfate is dissolved in sterile PBS solution. Amoxicillin is dissolved in sterile DMSO solvent.

2.4 2.4	Prepare two-fold serial dilutions of the test antibiotics, gentamicin or amoxicillin, using the aliquots of bacteria solution as a dilution medium. For both of gentamicin or amoxicillin, the serial concentrations cover from 0.25 to 8 µg/mL.

11. 3.3: Citation to show that this is mimicking UTI?

Re: We appreciate the editorial comment. We added the following citations:

14	Baltekin, Ö., Boucharin, A., Tano, E., Andersson, D. I. & Elf, J. Antibiotic susceptibility testing in less than 30 min using direct single-cell imaging. Proceedings of the National Academy of Sciences. 114 (34), 9170-9175 (2017).

18	Schoepp, N. G., Schlappi, T. S., Curtis, M. S., Butkovich, S. S., Miller, S., Humphries, R. M. et al. Rapid pathogen-specific phenotypic antibiotic susceptibility testing using digital LAMP quantification in clinical samples. Science Translational Medicine. 9 (410), eaal3693, doi:10.1126/scitranslmed.aal3693, (2017).

41	Michael, I., Kim, D., Gulenko, O., Kumar, S., Kumar, S., Clara, J. et al. A fidget spinner for the point-of-care diagnosis of urinary tract infection. Nature Biomedical Engineering. 4 (6), 591-600, doi:10.1038/s41551-020-0557-2, (2020).

12. 4.2: Citations to show that is mimicking blood stream infection?

Re: We appreciate the editorial comment. We added the following citations:

42	Bhattacharyya, R. P., Bandyopadhyay, N., Ma, P., Son, S. S., Liu, J., He, L. L. et al. Simultaneous detection of genotype and phenotype enables rapid and accurate antibiotic susceptibility determination. Nature Medicine. 25 (12), 1858-1864, doi:10.1038/s41591-019-0650-9, (2019).

43	Stupar, P., Opota, O., Longo, G., Prod'hom, G., Dietler, G., Greub, G. et al. Nanomechanical sensor applied to blood culture pellets: a fast approach to determine the antibiotic susceptibility against agents of bloodstream infections. Clinical Microbiology and Infection. 23 (6), 400-405, doi:10.1016/j.cmi.2016.12.028, (2017).

13. 4.4. Please check and reword for clarity… centrifugation speed in x g, time and temperature?

Re: We appreciate the editorial comment. We added more information in part 4.4:

“…, and then harvest bacteria to 1 mL volume by centrafuged at 8000 rpm for 5 min at 4 °C.”

and in part 2.9:

“… was first centrifuged at 8000 revolutions per minute (rpm) for 5 min at 4 °C, …”

14. 5.3.1: What is C-D vibrational frequency?

Re: We appreciate the editorial comment. We added: “To image bacteria at the C-D vibrational frequency at 2167 cm-1, …”

15. 5,6, 7: How is this done?

Re: We appreciate the editorial comment. We have modified part 5, 6, 7 to including mechanical actions and button clicks in the software.

16. 8: We do not have paragraph of text in the protocol section. The Protocol should contain only action items that direct the reader to do something. This can be moved to representative result/discussion section as appropriate. If this needs to be here, please make numbered action step and include how this is done.

Re: We appreciate the editorial comment. The “Conclusion” part is moved down to the place between Representative result and Discussion.

17. Please include a single line space between each steps of the protocol and then highlight 3 pages or less of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol.

Re: We appreciate the editorial comment. The line space is set to 1.0. And the essential steps of the protocol for the video are highlighted in bold.

18. Please remove the embedded figure(s) from the manuscript. All figures should be uploaded separately to your Editorial Manager account. Each figure must be accompanied by a title and a description after the Representative Results of the manuscript text.

Re: We appreciate the editorial comment. We have removed the figures, and kept the captions for each figure.

19. Please remove the embedded Table from the manuscript. All tables should be uploaded separately to your Editorial Manager account in the form of an .xls or .xlsx file. Each table must be accompanied by a title and a description after the Representative Results of the manuscript text. Table of materials should be uploaded separately as a .xlsx file.

Re: We appreciate the editorial comment. We have removed the “TABLE OF MATERIALS” part and the “Troubleshooting table”. The two tables are uploaded separately in the form of an .xlsx file.

20. Please obtain explicit copyright permission to reuse any figures from a previous publication. Explicit permission can be expressed in the form of a letter from the editor or a link to the editorial policy that allows re-prints. Please upload this information as a .doc or .docx file to your Editorial Manager account. The Figure must be cited appropriately in the Figure Legend, i.e. “This figure has been modified from [citation].”

Re: We appreciate the editorial comment. The publisher John Wiley and Sons clarified that this is an open access article under the terms of the Creative Commons Attribution License, which permits use, distribution and reproduction in any medium, provided the original work is properly cited.

21. As we are a methods journal, please ensure that the Discussion explicitly cover the following in detail in 3-6 paragraphs with citations. Please do not make bullet points.
a) Critical steps within the protocol
b) Any modifications and troubleshooting of the technique
c) Any limitations of the technique
d) The significance with respect to existing methods
e) Any future applications of the technique

Re: We appreciate the editorial comment. The bullet points have been modified to a), b), … as instructed.

22. Please ensure that the references appear as the following: [Lastname, F.I., LastName, F.I., LastName, F.I. Article Title. Source. Volume (Issue), FirstPage – LastPage, (YEAR).] For more than 6 authors, list only the first author then et al.
Re: We appreciate the editorial comment. The references are modified to the format as instructed.

Reviewers' comments:
Reviewer #1:
Manuscript Summary:
In this manuscript the authors describe the protocol for rapid single cell antimicrobial susceptibility testing, by using D2O incorporated stimulated Raman scattering imaging. This protocol will be of interest to researchers for rapid AST, and may be considered for acceptance after the following concerns have been addressed.

Re: We appreciate the reviewer’s positive comments. 

Major issue:
The authors did not give enough and appropriate references to related and previous work.
--- SRS imaging of D2O metabolism in single cells (as well as animals) was first developed by Shi et al. (Nat Commun. 2018. 9,2995), which provides the underlying foundation of the current work, but this reference was ignored in the manuscript. For example, Line 68, "Cellular metabolic activity can be probed via monitoring … so called heavy water (Figure 1)" - lacks appropriate references. Same to Line 80.
--- Appropriate citations should also be added to the description of figures in both the text and figure captions, in addition to the reference cited in Figure 4 caption.

Re: We appreciate the reviewer’s comments. We added references in the text and Figure 1 caption.

In "REPRESENTATIVE RESULTS", many sentences are identical to those in the authors' previous publication.

Re: We appreciate the reviewer’s comments. We have modified these sentences to avoid overlap.

All through the Protocol, the authors used inconsistent tenses (present/past) and active/passive voice.

Re: We appreciate the reviewer’s comments. We corrected all the sentences into present and active voice.

In addition, many places in the Protocol the description needs to be clearer and more straightforward. For example, Line 120. It is unclear what the concentration of D2O is, 100%? And what is the final concentration of D2O in the medium for the test?

Re: We appreciate the reviewer’s comments. We modified the sentences in part 2.7 as following:

“2.7	In the meantime of incubation, prepare the serial diluted concentration of antibiotics in the 100% D2O containing medium with the same concentration gradient of antibiotics.”

In part 2.8, we added explanation as following:

“... In this step, the final concentration of D2O in the medium for the test is 70%.”

Did the authors conduct the D2O toxicity testing in bacteria?

Re: We appreciate the reviewer’s comments. Yes, we conducted the D2O toxicity testing in bacteria. We note that D2O labelling above 70% affects bacterial metabolism significantly was observed during a 25-hour incubation period (Figure S1). Our current results show that 70% D2O does not affect the performance of our method in P. aeruginosa, because we have performed MIC by gold standard broth dilution and obtained SC-MIC results in two P. aeruginosa strains (Table S1). The category agreement of our method (SC-MIC) with the conventional culture-based method is 100% for all the tested P. aeruginosa and antibiotic combinations, as shown in the following Table. We attribute this good agreement to the minimal toxicity of 70% D2O on P. aeruginosa during the 30 min incubation period in SC-MIC determination.

Figure S1. Testing D2O toxicity to P. aeruginosa cultured in LB medium with different D2O concentrations. Error bars indicate standard deviation values (number of measurements = 5).

Table. Comparison of SC-MICs and MICs in normal MHB of P. aeruginosa upon antibiotics. S: sensitive; R: resistant; I: Intermediate.

| Bacteria strains | Antibiotics | SC-MIC (µg/mL) | MIC in normal MHB (µg/mL) | MIC/SCMIC category |
| --- | --- | --- | --- | --- |
| P. aeruginosa ATCC 47085 (PAO1) | Cefotaxime | 32 | 16 | I/I |
|  | Colistin | 0.5 | 2 | S/S |
|  | Gentamicin | 2 | 4 | S/S |
|  | Amikacin | ≤4 | 2 | S/S |
|  | Tobramycin | ≤1 | 0.5 | S/S |
|  | Ciprofloxacin | ≤0.25 | 0.125 | S/S |
| P. aeruginosa ATCC 1133 | Colistin | 8 | ≥16 | R/R |

Reviewer #2:
Manuscript Summary:
A rapid antibiotic susceptibility testing (AST) method for a single bacterium is proposed in this article by adopting D2O and Stimulated Raman Spectroscopy (SRS). Living bacteria can incorporate deuterium from D2O into their organic compounds that form a C-D bond, which reflected their metabolism activity. The intensities of C-D bond Raman signals from a single bacterium with different antibiotic concentration treatments are measured by the SRS image with further data process. By comparing with the standard minimum inhibition concentration (MIC) result, the single-cell metabolism inactivation concentration (SC-MIC) of the bacterium is determined by the relative intensity threshold of 0.6. The major contribution of this manuscript is to provide a detailed SC-MIC protocol, including bacteria sample preparation, treatments, SRS image measurement, data processing, and analysis. Representative results of this method are also provided.

Re: We appreciate the reviewer’s positive comments. 

Major Concerns:
Overall, the experimental information and the description are quite clear. My only concern is that most figures used in this manuscript are duplicated from the authors' previous article (ref 38), not sure if it is common in JoVE. If the above issue is acceptable, I think the manuscript can be considered as acceptable. Some minor comments are shown below.

Re: We appreciate the reviewer’s comments. We have confirmed with the editor Dr. Vineeta Bajaj that we can reuse the figures from previous publications as long as reprint permission is obtained and the original publication is cited in the figure legend.

Minor Concerns:
1. Some experimental descriptions are not complete. For example, how to determine the relative intensity threshold of 0.6 is not described in this manuscript. The authors only add a citation of ref 38 without any explanation, which may be confused to the readers (line 196 on page 6).

Re: We appreciate the reviewer’s comments. We added a paragraph on how to determine the threshold in part 7.1:

7.1 The cut-off value at 0.60 to determine the SC-MIC is established according to the statistical analysis of the SRS C-D intensities of the metabolism-active and metabolism-inhibited conditions for bacteria upon various concentrations of drug exposure.40 The C-D intensities for the antibiotic-susceptible and antibiotic-resistant groups were fitted with normal distribution. Plot the receiver operating characteristic (ROC) curve and the cut-off threshold is evaluated at 0.60. Based on this cut-off value, the SC-MIC as an indicator of the efficacy of antibiotics can be defined to determine the metabolically inactive and metabolically active group.

2. Various spelling errors or repeated words appeared in the manuscript. (Ex. line 167 171 on page 5; "spliked" line 250 on page 7)
Re: We appreciate the reviewer’s comments. We corrected the typos in the manuscript.

Reviewer #3:
Manuscript Summary:
The paper titled "Rapid Antimicrobial Susceptibility Testing by Stimulated Raman Scattering Imaging of Deuterium Incorporation in a Single Bacterium" presents excellent results based on the investigation of bacterial metabolic capacity when exposed to antibiotics, with bacterial response measured via the C-D vibrational band on Raman spectrum using Stimulated Raman scattering imaging. The authors demonstrate that bacteria which are susceptible to antibiotics can be identified within 2.5 h which is ideal for timely intervention of appropriate prescription of antibiotic drugs.
Re: We appreciate the reviewer’s comments.

Major Concerns:
Figures of Raman spectra showing C-D and C-H bands and their intensity dynamics are needed to make the article clearer. Include more explanation on the dynamics of C-D intensities for antibiotic-susceptible bacteria; is there any time course or concentration level dependence demonstrated by the C-D band? Computing the C-D/C-H ratios between microbes exposed to D2O/antibiotics and microbes in the control class (i.e., microbes that are not treated with antibiotics) would also help with elucidating the extent of D incorporation into biomass and level of antimicrobial susceptibility.

Re: We appreciate the reviewer’s comments. We added a paragraph and a figure discussing effect of incubation time on deuterium incorporation measured by spontaneous Raman microspectroscopy as following.

The effect of incubation time on deuterium incorporation is measured by spontaneous Raman microspectroscopy at the C-D (2070 to 2250 cm-1) and C-H ﻿(2,800 to 3,100 cm-1) region (Figure 4a). The time-lapse single-cell Raman spectra of P. aeruginosa cultured in 70% D2O containing medium show increasing CD/CH intensity over incubation time from 0 to 180 min. (Figure 4b) The increasing C-D abundance in single microbial cells reveals that D2O is incorporated into deuterated biomolecules inside the cell.

Figure 4. Effect of incubation time on deuterium incorporation into bacteria. (a) Time-lapse measurement at C-D (2070 to 2250 cm-1) and C-H ﻿(2,800 to 3,100 cm-1) region by spontaneous Raman microspectroscopy (averaged from 20 spectra). (b) Histogram plot of CD/CH intensity ratio plot over D2O incubation time for bacteria in (a). Each colored point stands for a measurement from a single bacterium. Error bars represent the standard error of the mean (SEM).

Minor Concerns:
Page 3, line 71: "tob e" should be "to be"

Re: We appreciate the reviewer’s comments. We corrected this typo.

Page 4, line 90: Delete "time" repeated word

Re: We appreciate the reviewer’s comments. We corrected this typo.

The "Protocols section" is written in both passive and active voices; authors should choose one format and be consistent.

Re: We appreciate the reviewer’s comments. We corrected all the sentences into active voice.

Page 4, line 105: It is not clear if the 2mL of normal MHB medium drawn from D2O-labelled or H2O-labelled MHB medium? Does "normal MHB" imply H2O-labelled medium? Please make this clear.

Re: We appreciate the reviewer’s comments. We modified this sentence as following:

“… add 2 mL of normal MHB media, which does not contain deuterium, …”

Page 4, line 120: Please mention specific antibiotics and concentration at this stage.

Re: We appreciate the reviewer’s comments. We modified this sentence as following:

“Prepare two-fold serial dilutions of the test antibiotics, gentamicin or amoxicillin, using the aliquots of bacteria solution as a dilution medium. For both of gentamicin or amoxicillin, the serial concentrations cover from 0.25 to 8 µg/mL.”

Page 4, line 123: What were the actual antibiotic concentrations in serially diluted samples?

Re: We appreciate the reviewer’s comments. We clarified the actual antibiotic concentrations in parts 2.4 and 2.7:

“…For both of gentamicin and amoxicillin, the serial concentrations cover from 0.25 to 8 µg/mL.”

Page 4, line 125-129: Please include specific shaking and centrifugation parameters.

Re: We appreciate the reviewer’s comments. We added specific shaking and centrifugation parameters as following:

2.8 “…and then incubate at 37 °C in a shaker at 200 revolutions per minute (rpm)…”

2.9 “First centrifuge the 1 mL of antibiotic and D2O-treated bactria sample at 8000 rpm for 5 min at 4 °C, …”

Page 5, line 169: What was the step size of the SRS mapping?

Re: We appreciate the reviewer’s comments. We added the step size information as following:

5.5 “Set the step size for 150 nm, ...”

Page 7, line 230: Bacteria were grown in 70% D2O-labelled MHB. D2O labelling above 50% affects bacterial metabolism significantly (e.g., Berry et al., Proc. Natl. Acad. Sci. USA, 2015, 112, E194). So, did you observed any 70% D2O-mediated metabolic response in microbes, in addition to the influence of antibiotics? Any justification as to why such a high concentration of D2O was used? Would the results be different if say, 30-40% D2O was used instead?

Re: We appreciate the reviewer’s comments. We note that D2O labelling above 50% affects bacterial metabolism significantly was observed during a 23-hour incubation period in this paper (Berry et al., Proc. Natl. Acad. Sci. USA, 2015, 112, E194). We observed bacterial growth inhibition when D2O labelling concentration is above 70% during a 25-hour incubation period (Figure S1). Our current results show that 70% D2O does not affect the performance of our method in P. aeruginosa, because we have performed MIC by gold standard broth dilution and obtained SC-MIC results in two P. aeruginosa strains (Table S1). The category agreement of our method (SC-MIC) with the conventional culture-based method is 100% for all the tested P. aeruginosa and antibiotic combinations, as shown in the following Table. We attribute this good agreement to the minimal toxicity of 70% D2O on P. aeruginosa during the 30 min incubation period in SC-MIC determination.

Figure S1. Testing D2O toxicity to P. aeruginosa cultured in LB medium with different D2O concentrations. Error bars indicate standard deviation values (number of measurements = 5).

Table. Comparison of SC-MICs and MICs in normal MHB of P. aeruginosa upon antibiotics. S: sensitive; R: resistant; I: Intermediate.

| Bacteria strains | Antibiotics | SC-MIC (µg/ml) | MIC in normal MHB (µg/ml) | MIC/SCMIC category |
| --- | --- | --- | --- | --- |
| P. aeruginosa ATCC 47085 (PAO1) | Cefotaxime | 32 | 16 | I/I |
|  | Colistin | 0.5 | 2 | S/S |
|  | Gentamicin | 2 | 4 | S/S |
|  | Amikacin | ≤4 | 2 | S/S |
|  | Tobramycin | ≤1 | 0.5 | S/S |
|  | Ciprofloxacin | ≤0.25 | 0.125 | S/S |
| P. aeruginosa ATCC 1133 | Colistin | 8 | ≥16 | R/R |

Page 14, line 353, Table 1 row number 2: Would "centrifuge for longer time" advice not damage or lyse bacterial cells?

Re: We appreciate the reviewer’s comments. According to the protocol listed in the paper [﻿Anal Bioanal Chem (2017) 409:3043], the high-speed centrifugation ﻿at 10,000 rpm for 20 min can be applied to the bacterial cells. In our protocol, we usually centrifuge less than 10 min for the fixed bacterial cells.