## Table1
Table1_Troubleshooting for high working backpressure of the column
There are several possible causes for high backpressure, some of which are shared between fritted and pulled emitter columns, while some are specific for pulled emitter columns.
Problem* | Possible cause and solution
A pulled emitter column has a backpressure much higher than expected for the column length. | Possible cause: long pulled region with narrow ID.
Comment: capillary pulling in a Sutter laser puller implies a part of the capillary with a gradually narrowing ID. If the pulling program is not optimized, a prolonged narrow ID region can be formed which, when packed, has a backpressure much higher than the same length of the original ID capillary. As a result the final pulled-emitter column backpressure is higher than the for the same length of a fritted capillary without narrowing. To reduce the effect, optimize the pulling program to achieve a better hour-glass shape approximation of the pulled end.
A pulled emitter column end is blocked by sorbent particles clogging the pulled end (visible under the microscope). | Cause: self-assembling frit structure is broken.
Comment: to make a working column, the sorbent in the pulled end is supposed to form a self-assembling frit –a sorbent wall at some distance from the emitter opening. If the sorbent packs up to the very tip, sorbent particles can partially or completely clog the tip, leading to column blocking. For a very narrow emitter end, a single sorbent particle can block it completely. There are two infallible solutions:
- use a specially prepared pulled capillaries with frits integrated into the pulled end (see PicoFrit self-pack columns from NewObjective), which do not rely on the self-assembling frit formation;
- adjust the program to get the emitter end to be several times the size of the sorbent particle. In this case, even if the self-assembling frit structure is broken, the column will still work at higher backpressure.
The former is good but expensive, and at the moment the item might be out of production.
The latter is practically extremely hard to achieve in a Sutter P2000 laser puller for 75 µm ID capillaries. When pulled normally (the capillary is heated evenly and the walls and capillary ID narrow proportionally when pulled), 75 µm ID capillary separates with the opening around 2-3 µm. Achieving 5-7 µm is possible, but requires a very special laser puller program where the outer capillary layer has a higher temperature than the capillary inside. When pulled, capillary lengthening happens mostly due to outer layer flowing to compensate for lengthening, and the capillary wall narrows faster than the capillary ID. In the end, the emitter separation happens with larger opening diameter. However, such programming is individual for each P2000 instrument. Today, Sutter makes such programs at the factory and provides them for each new P2000 laser puller instrument installed. However, old instruments do not have such programs (though it was possible to send the instrument to the Sutter factory to make such programs), and from our experience, we could not make a required program by ourselves. So, we suggest using 100 µm ID capillaries instead. They have thinner walls from the beginning which separate with the opening around 3-5 µ m under normal pulling conditions (an example of the pulling program is in Table 2). Thus, they are much less prone to clogging due to separate particles getting into the column end – separate particles just flush out of the capillary while the major particle front forms the self-assembling frit. At the same time, to achieve exactly the same sensitivity for 100 µm ID columns in comparison with 75 µm ID columns, it might be necessary to increase the peptide loading amount ~ 1.5-2 times.
Independently of the capillary used for the column preparation, avoid the factors leading to the self-assembling frit breaking:
- do not immerse pulled emitter end into the sonication bath during packing;
- avoid column end drying: do not leave the column without flow when connected the LC system; fill the emitter end protecting tip with water before disconnecting the column from the LC system and place the column into the storage box filled with ethanol-water solution as soon as possible;
- avoid dramatic pressure drops, e.g. valve opening, tubing connection disassembling under pressure.
The backpressure is higher than expected from the column length and the effect is reproducible from column to column – happens for both fritted and pulled emitter columns | Cause: the sorbent in the working vial is ground by the magnet bar with the accumulation of small sorbent debris.
Comment: The cause can be pinpointed based on the following indications:
- gradual rise of the columns pressure in time with each new packed column batch with the use of the same sorbent vial for packing: e.g. columns produced two weeks ago were fine, last week columns had the pressure a little higher, today it is even higher;
- very long usage of the same working sorbent vial for packing;
- the working sorbent vial was forgotten in the pressure bomb with the rotation turned on for a prolonged period of time.
To avoid that problem, prepare fresh working sorbent as a first-step in packing troubleshooting. The amount of the sorbent in the working vial is cheaper than the time and resources spent on finding the problem.
*Before going to packing troubleshooting, make sure, that the overpressure is not due to an incorrectly working pump, wrong buffers or blocked or restricted connections, capillaries and emitters, but is really a column problem.