2. Mass spectrometry

1. Spot 0.75 μL aliquot of the sample supernatant onto the

stainless steel MALDI target and allow it to dry. Overlay

the dried sample spot with a 0.75 μL aliquot of a saturated

solution of sinapinic acid in 33% acetonitrile, 67% water,

and 0.2% trifluoracetic acid. Allow the spot to dry.

2. Analyze the dried sample spots using a MALDI-TOF-TOF

mass spectrometer.

1. After loading the MALDI target into the mass

spectrometer, SHOT#1_click the button for MS linear mode

acquisition in the acquisition software. SHOT#2_Enter the m/z

range to be analyzed by entering the m/z of the lower

and upper bounds (e.g., 2 kDa to 20 kDa) into their

respective fields in the acquisition method software.

2. SHOT#3_Click on the sample spot to be analyzed on

the MALDI target template in the software. Then,

SHOT#4_depress the left mouse button and drag the

mouse cursor over the sample spot to specify

the rectangular region to be sampled for laser

ablation/ionization. Release the mouse button and

the acquisition will initiate. SHOT#5_Collect 1,000 laser shots

for each sample spot.

NOTE: Data acquisition is displayed in real-time in

the software acquisition window.

3. If no ions are detected, SHOT#5_increase the laser intensity

by adjusting the Sliding Scale Bar under Laser

Intensity in the software until the protein ion signal

is detected. This is referred to as threshold.

NOTE: Prior to the sample spot analysis, externally

calibrate the instrument in MS linear mode with

protein calibrants whose m/z span the range being

analyzed, e.g., the +1 and +2 charge states of

protein calibrants: cytochrome-C, lysozyme, and

myoglobin cover a mass range of 2 kDa to 20 kDa.

An intermediate mass within the specified mass

range is used as a focus mass, e.g., 9 kDa. The

focus mass is the ion whose m/z is optimally focused

for detection by the linear mode detector.

4. When the MS linear mode acquisition is complete,

SHOT#6_click the button for MS/MS reflectron mode

acquisition in the acquisition software. SHOT#7_Enter the

precursor mass to be analyzed into the Precursor

Mass field. SHOT#8_Next, enter an isolation width (in Da) into

the Precursor Mass Window for the low and high

mass side of the precursor mass, e.g., ±100 Da.

5. SHOT#9_Click on the CID Off button. Click on the Metastable

Suppressor ON button. SHOT#10_Adjust the laser intensity to

at least 90% of its maximum value by adjusting the

sliding scale bar under the Laser Intensity in the

software.

6. SHOT#11_Click on the sample spot to be analyzed on

the MALDI target template in the software. SHOT#12_Then

depress the left mouse button and drag the

mouse cursor over the sample spot to specify the

rectangular region to be sampled for laser ablation/

ionization. Release the mouse button, and the

acquisition will initiate. SHOT#13_Collect 10,000 laser shots for

each sample spot.

NOTE: Prior to the sample spot analysis, the

instrument should be externally calibrated in MS/

MS-reflectron mode using the fragment ions from

post-source decay (PSD) of the +1 charge state of

alkylated thioredoxin35.

3. Do not process raw MS data. Process MS/MS-PSD

raw data using the following sequence of steps in the

specified order: advanced baseline correction (32, 0.5,

0.0) followed by noise removal (two standard deviations)

followed by Gaussian smoothing (31 points).

4. Manually inspect MS/MS-PSD data for the presence of

prominent fragment ions generated by PBC19 , 20.

5. Evaluate MS/MS data with respect to the absolute and

relative abundance of fragment ions and their signal-to-noise

(S/N). Use only the most abundant fragment ions

for protein identification, especially if the MS/MS-PSD

data is noisy.