TITLE:

Supplementary Protocol 3 – Construction of the Pneumatic Powered OG Injury Device

PROTOCOL:

Open globe injuries are induced via a pneumatic powered linear actuator capable of producing high force injuries of various shapes and sizes. The system was constructed using readily available parts and supplies and includes the framework to hold the device static while inducing injuries. While an air compressor was utilized to generate the air pressure to drive the device, any compressed air source can be utilized as long as 50 psi can be generated. Below are instructions and diagrams to construct the device and a Table of Materials identifying each required component.

- A list of required parts is detailed in the Table of Materials for this supplementary protocol.

- Assemble the electrical components as detailed in Supplementary Figure 4B. Crimp connectors were used throughout for ease of connecting and disconnecting as was required.

- Pressure components can be assembled as detailed in Supplementary Figure 4C. Ensure each connection is pressure tight by using Teflon tape for each threaded part.

- Construct the t-slot framing setup to house the flow and electrical components as detailed in Supplementary Figure 4A.

- Set the air pressure regulator at 50psi and ensure the relief valve in the system is set at 75psi. Power on the electronics and ensure the solenoid valves open and close with the momentary switch.

- Connect compressed air to the system and ensure pressure in the system reached 50psi. Check for any leaks in the system by closing of the compressed air source and confirmed pressure holds in the system.

- Turn on/off the solenoid valves to confirm they drive the linear actuator forward and back. The device is now ready for injury induction in ASOC setups as detailed in the protocol.

FIGURE AND TABLE LEGENDS:

Supplementary Figure 4: Overview of the open globe induction setup. (A) The injury device fully constructed. (B) Electrical diagram for controlling the solenoid valves for forwarding and retracting the linear actuator. (C) flow diagram for compressed air through the regulating valve and to the solenoid valves which direct air flow to the linear actuator and puncture object held in the drill chuck. Numbers indicate the required parts as identified in the Table of Materials for this supplementary protocol.

TABLE OF MATERIALS:

# correspond to labels in the Supplementary Figure 4

| # | Name | Company | Catalog Number | Comments/Description |
| --- | --- | --- | --- | --- |
|  | General Device Components |  |  |  |
| 9 | 304 Stainless Steel Cross Connector | McMaster-Carr | 4464K311 | 1/8” NPT Female |
| 7 | 304 Stainless Steel Straight Connector | McMaster-Carr | 4464K711 | 1/8” NPT Male Quantity = 3 |
| 8 | 304 Stainless Steel Tee Connector | McMaster-Carr | 4464K47 | 1/8” NPT Female |
| 1 | Air Cylinder Double Acting, 5/16” Bore, 1-½” Stroke | McMaster-Carr | 6498K13 |  |
| 14 | Air Compressor | McMaster-Carr | 9965K62 | Any source of compressed above 50psi will suffice |
| 6 | Digital High Accuracy Pressure Gauge | McMaster-Carr | 1287N1 | 1/8” NPT |
| 2 | Drill Chuck | Grainger | CSS-06 R-5/16 GE-DS | 5/16” - 24 |
| 3 | Electric Solenoid Control Valve | McMaster-Carr | 61245K84 | 1/8” NPT, 12V DC Quantity = 2 |
| 11 | Hard Nylon Plastic Tubing | McMaster-Carr | 5548K84 | ¼” Tubing OD, 11/64” ID; 25ft Length |
| 5 | Precision Compressed Air Regulator (1 – 120psi) | McMaster-Carr | 2227T2 | 1/8” NPT |
| 4 | Pressure Relief Valve | McMaster-Carr | 5784T11 | Set at 75psi, 1/8” NPT |
| 10 | Push-to Connect Rub Fitting | McMaster-Carr | 5779K108 | ¼" Tubing OD, 1/8” NPT Male Quantity = 7 |
| 13 | Thread Sealant Tape | McMaster-Carr | 1726N12 | 3/8” wide, for Stainless Steel Threads |
|  | Electronic Components |  |  |  |
| 15 | DC Power Supply | Digi-Key | WSU120-1000 | 12V 1A DC |
| 18 | Female Crimp Connector | McMaster-Carr | 7243K11 |  |
| 17 | Male Crimp Connector | McMaster-Carr | 7243K12 |  |
| 16 | Push-Button Switch | API Electric | B07932GYNH |  |
| 20 | On-Off Switch | SparkFun | 11138 |  |
| 19 | Wire | McMaster-Carr | 8054T15 |  |
|  | Structural Framing |  |  |  |
| 27 | 1” T-Slotted Framing Bracket | McMaster-Carr | 6498K992 | Quantity = 2 |
| 23 | 4 Slot Rail , 1” L x W x H | McMaster-Carr | 47065T101 | Quantity = 2 |
| 22 | Corner Bracket, 2” long, 1” rail | McMaster-Carr | 47065T239 | Quantity = 5 |
| 28 | Foot Bracket for 5/16” Bore Size Air Cylinder | McMaster-Carr | 6498K581 | Quantity = 2 |
| 20 | Optical Breadboard Plate | Newport | SA2-11 | 12”x12”, ¼” - 20 |
| 26 | Tie Cable Holder | McMaster-Carr | 47065T268 | Quantity = 2 |
| 21 | T-Slotted Framing, 1” Rail, 12” Length | McMaster-Carr | 47065T188 | Quantity = 1 |
| 29 | T-Slotted Framing End Cap | McMaster-Carr | 3136N2 | Quantity = 1 |
| 24 | T-Slotted Framing, End- Fed Single Nut | McMaster-Carr | 47065T139 | Quantity = 3 |
| 25 | T Slotted Framing Snap Cap Cable Holder | McMaster-Carr | 47065T389 | Quantity = 2 |