Editorial comments:

1. Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues. Please define all abbreviations at first use and use American English.

The document was proofread to correct spelling and grammar. All abbreviations were defined upon first use, e.g., outer diameter (OD) in step 3.5. and digital multiplex (DMX) in step 2.7. The text has been converted to American English.

2. Please reduce your summary’s word count to <50.

The summary is now 47 words. Summary - “This publication describes the design of laboratory photobioreactors (PBRs) with customizable light regimes.  The growth of cyanobacteria or microalgae, using bicarbonate as their carbon source, is monitored continuously by measuring volumetric oxygen production. These PBRs facilitate rapid, replicated, laboratory growth comparisons with little user intervention during experiments.”

3. JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials and Reagents. For example: Rigamo; DMX; ESA etc.

Thank-you for clarification. I was unaware that software should also be included in the Table of Materials. This has been amended. DMX stands for Digital MultipleX. It is not a commercial product but rather a type of lighting control, thus this remains throughout the text. The text has been updated so that commercial names are mentioned only once when software is first introduced as suggested. Generic terms like ‘data acquisition software’ and ‘LED control software’ are used throughout the remainder of the text.

Line 244 - “5.7. Install the data acquisition software, Rigamo, on a Windows operating system and plug in licensing key dongle. Add sensor calibration files to the software’s calibration directory.”

Line 252 – “6.2. Download Store Upgrade Tools (SUT) and the LED control software, Easy Stand Alone (ESA). Register the DMX lighting controller online.”

4. Obviously, if you have optimized your protocol with certain instruments/reagents/software, you will need to mention them, but include just the bare minimum information (e.g., name of product) in the manuscript and come up with a generic term to refer to the product (include this in the comments column in the Table of Materials) after the first mention so that you don’t keep repeating commercial terms throughout the paper. You can mention products or software in the discussion if you wish to compare them to what you have used.

Refer to the response to comment #3. The Table of Materials has been updated to include the generic term in the comments section. Further, all URL links for software have also been shifted to the Table of Materials.

5. Please revise the text, especially in the protocol, to avoid the use of any personal pronouns (e.g., "we", "you", "our" etc.).

Text has been revised and personal pronouns have been removed.

6. Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly. Please include all safety procedures and use of hoods, etc.

Steps within the protocol that were not in the imperative sense have been modified, e.g., step 2.4. was modified to remove “The fit should be snug”. Notes have been revised to be more concise. Three notes have been removed altogether, with the information they conveyed instead shifted to the adjacent steps. In particular, the protocol now only recommends colour coding of materials once after step 3.7.

7. Please note that your protocol will be used to generate the script for the video and must contain everything that you would like shown in the video. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol.

Additional details were added to the protocol to better describe how each step is performed. For example, in step 1.1. and 1.2. it is now specified that fastening/securing is achieved with bolts. In step 2.5, it is now detailed that “medium length rivets” are used to attach the piano hinge. In step 7.8. the text has been updated from “Insert rubber stoppers into each bottle opening and seal with threaded aperture screw caps” to “Seal each bottle opening with a rubber stopper and threaded aperture screw cap.” Further, references for determining the ash-free dry weight of microalgal cultures was added to the note after step 7.6.

8. Please sort the Materials Table alphabetically by the name of the material.

The Table of Materials is now sorted alphabetically.

Reviewers' comments:

Reviewer #1:
The manuscript Design, construction, and operation of laboratory photobioreactors with online growth measurements and customisable light regimes provides a detailed protocol for building a DIY PBR and a thorough discussion of specifications.

Clarity: The introduction needs a bit of rephrasing, but the protocol for building the PBR and discussion is clear.

Usefulness: the usefulness of the described DIY PBR is obvious despite some details mentioned in comments 12 and 13.

My comments and suggestions:

1. Please improve the English. The introduction and abstract are a bit difficult to follow.

We agree with reviewer #1s comment and thank them for this advice. The abstract and introduction have been extensively reworked with an emphasis on reader comprehension.

2. Avoid using slash e.g. in and/or, cyanobacteria/microalgae, species/strains...

We have heeded reviewer #1’s suggestion and defined microalgae as including both cyanobacteria and microalgae. See line 52 – “Microalgae and cyanobacteria, herein collectively referred to as microalgae for simplicity, are championed for their potential in sustainable biotechnology.” This eliminated the need for slashes in this context. Further, all and/or and species/strains have now been removed.

3. It might be better to define microalgae as algae including cyanobacteria in the introduction. That would make it easier to refer to all photosynthesizing microbes (phototrophs, microalgae, microbes, cyanobacteria) at once as well as avoid using a bit unusual terms like microbial photosynthesis.

See the response to comment #2.

4. Line 22 and 34: "...using bicarbonate." - using bicarbonate for what? Probably as a carbon source? Please rephrase.

These sentences have been updated to specify that bicarbonate is used as a carbon source. Line 21 now reads “The growth of cyanobacteria or microalgae, using bicarbonate as their carbon source, is monitored continuously by measuring volumetric oxygen production.” Line 46 now reads “The method is suited for microalgae grown with bicarbonate as the carbon source.” Line 104 now reads “For microalgae that use bicarbonate as the carbon source, there are two additional growth metrics that can be measured continuously.”

5. Line 41: "...and can be adjusted on fine time scales (1 sec)." - not clear what you mean, probably photoperiod? It would be better if this part of the sentence was put after the following sentence dealing with light regimes.

The abstract has been updated for clarity. The segment “…and can be adjusted on fine time scales (1 sec).” has been updated to “Light intensity can be adjusted each second or held constant for longer durations.” This refers to the duration of ‘steps’ within the LED control software. Each step can be set to last between 1 second and 43 minutes and from 0-2400 µmolphotons.m-2.s-1. This is further detailed in protocol step 6.7 and the subsequent note.

6. Line 53: "...and for using sunlight to drive carbon dioxide conversion." - conversion to what? Please specify.

The introduction has been altered to clarify that CO2 can be converted into products. Line 53-55 - “They are attractive candidates due to their rapid growth, ability to be cultivated on non-arable land, and for their use of sunlight to drive carbon dioxide conversion to biomass.  Microalgal biomass can be converted into products like…”

7. Lines 74-82 talk about biomass assessment by measuring DW and the need for large sample volumes. BUT, biomass can be (and very often is) nicely estimated by measuring OD against calibration to DW or pigment content. This can be done within the cultivation vessel if sensors are present (no need to extract samples) or measured by an external spectrophotometer needing usually just 250-1000 uL of sample volume. This makes your paragraph on large volumes a bit obsolete.

This is valuable feedback which has prompted the inclusion of additional information about common growth metrics. It has been noted in line 72 that “a suitable method to accurately quantify growth must be selected.” The text goes on to discuss cell counts, chlorophyll A content, dry weight (DW), ash-free DW (AFDW), and optical density (OD) between lines 74 and 86. OD is indeed a powerful non-invasive continuous method for measuring culture growth. We have previously explored it as an option in our laboratory, however due to the clumping nature of the organisms we work with it was unsuitable. This disadvantage is highlighted in line 83-84. In lines 104-117 we highlight the advantage of pH and O2 measurements as alternative growth metrics due to their insensitivity to changes in microalgal morphology. We emphasize in lines 84-86 that researchers should choose methods that are both practical and relevant to their experimental aims.

8. Line 84: "When the target organisms use bicarbonate as their source of CO2, no external supply of CO2 is needed." - Please rephrase e.g. external CO2 supply can be replaced by addition of bicarbonate into the cultivation media...

This paragraph containing this sentence has been reordered and rephrased for clarity. Instead, line 112-114 now reads “For pH and oxygen measurements, carbon must be supplied in the form of bicarbonate, not by CO2 sparging.”

9. It would be nice to mention the list of needed material before you start describing how to build the PBR.

The materials list is provided as a supplementary file. This format was requested by JoVE.

10. Figure 1B - pls change the background of letter B so that the letter is visible.

Figure 1B has been updated.

11. It would be nice to have a picture of the fully equipped PBR, open, including bottles filled with algae culture, preferentially with lights on. But this is just a non-essential detail.

The accompanying video should display the PBR in its entirety during operation.

12. Just a note: The PBR design is good. Nevertheless, I would miss temperature controller and the possibility to control light for each vessel independently.

Temperature control would indeed be a nice addition to the design. We have added a comment in the discussion at line 633-634 – “If temperature control of medium is desired, a recirculating chiller could be added.” The design could be altered to facilitate individual lighting control. However, we have refrained from suggesting this as the PBR is intended to facilitate equal lighting intensities to all three bottles for the sake of triplicate replication.

13. It would be good to include information on the possibility to autoclave (or sterilize otherwise) the materials that will get into contact with algae culture. Some tubes or plastic parts might not be autoclavable (rubber stoppers, tubing, sampling port valve...).

Thank-you for pointing out this oversight, we have amended section 7 of the protocol, Starting an experiment’. Now the first step, 7.1, reads “If required, sterilize media, bottles, stir bars, rubber stoppers, sampling ports, threaded aperture screw caps, and tubing.” This is followed by a subsequent note, “NOTE: All components used in this design are autoclavable except the valves and luer fittings—there are autoclavable alternatives from other manufacturers.”

14. The authors might want to include some comparison of their PBR design with some other existing commercial or DIY PBRs and highlight the advantages of their product preferably in the introduction and also in the conclusion.

Thank-you for this suggestion. We are aware that there are copious numbers of commercial PBRs and DIY prototypes available. What we believe is lacking, is a PBR that bridges the gap between simple benchtop flasks and very expensive, highly technical PBRs. We hope that this design is accessible to individuals or groups without a fabrication or engineering background. The JoVE format aids this by providing a step-by-step construction guide–something that is missing in published DIY PBR protocols. Given the sheer magnitude and variety of PBRs (each suited to a different experimental application) we believe a comparison of PBRs is best suited to a more extensive review article. We have added a paragraph at the end of the introduction which explains the gap this PBR hopes to fill. Further, we recommend that readers consider the suitability of a PBR design to their own culture conditions, financial position, and research questions.

Lines 119-130- “PBRs in use today range in complexity. Laboratories may use simple benchtop flasks, custom prototypes, or commercially available products. For research groups seeking to upgrade from flasks, the cost of commercial PBRs or technical skill and part fabrication required to build many prototypes, may be a barrier. The goal of this manuscript is to describe the step-by-step design, construction, and operation of laboratory PBRs that bridge this gap. These PBRs have a customizable light regime and monitor growth continuously by recording volumetric oxygen production. This design houses three culture vessels for triplicate replication, can be built with moderate skill, and is made with easily accessible materials. This PBR is a valuable addition to a laboratory looking to expand their capacity for microalgal research without investing in very technical or expensive products. When choosing to acquire or build a PBR, researchers must consider the suitability of a design to their own culture conditions, financial position, and research questions.”

In the conclusion, we have focused primarily on the Protocol, as direct by the Instructions for Authors document. Further, the discussion is limited to 6 paragraphs.

Reviewer #2:

Comments for JoVE62910

Design, construction, and operation of laboratory photobioreactor with online growth measurements and customisable light regimes

Keywords: photobioreactor, oxygen production, microalgae, cyanobacteria, biotechnology, photosynthesis.

The JoVE MS showed a lot of interesting information. However, the work just lacks some specific quantitative data ranges. Some comments are as below:

1. All the pages need consecutively numbering.

All pages have been numbered.

2. Line 51 - 53, "Microbial phototrophs………carbon dioxide conversion" What is sustainable biotechnology? How fast is the development?

We welcome the reviewer’s suggestion to elaborate on the definition of sustainable biotechnology and comment on the speed of development. However, we feel the introduction adequately conveys how technology which uses microalgal biomass is sustainable. Especially in lines 53-54 – “They are attractive candidates due to their rapid growth, ability to be cultivated on non-arable land, and for their use of sunlight to drive the conversion of carbon dioxide to biomass.” We believe that commenting on the speed of development is beyond the scope of this paper and is perhaps instead best suited to a review article.

3. Line 54 - 57, " Microbial phototrophs……….water remediation", more references are needed to clarify their application in making these products. Moreover, how specifically is the application in water and wastewater treatment?

Additional references have been added to the product examples. We have clarified the application briefly in water remediation and wastewater treatment in lines 57-58 – “In addition, they can be used to treat wastewater or remediate waterbodies by consuming excess nutrients”.

4. Line 62 - 63, " In cyanobaterial/microalgal…………, and method to quantify growth", needs adding references to clarify how the growth of cyanobacteria/microalgae in laboratory studies.

We have added a paragraph in the introduction which outlines the most common ways that microalgal growth is measured. Lines 74-86 – “Growth is commonly measured by cell count, optical density (OD), chlorophyll A content, dry weight (DW) density, and ash-free dry weight (AFDW) density.  Cell counts, chlorophyll A content, and gravimetric methods are manual processes which produce discrete data points. OD can be measured continuously and non-invasively with a spectrophotometer, providing it is well calibrated against another method like AFDW density. However, OD measurements and Chlorophyll A content can be unreliable as results vary under different culture conditions, e.g., between species and throughout the growth cycle. For Chlorophyll A, the extraction method used can also impact pigment yield. Chlorophyll A is typically used to assess the microalgal composition of microbial communities which also include non-photosynthetic organisms. When choosing a method to assess growth, the morphology of the suspension is also important. When organisms clump and are not-well mixed, OD and cell counts aren’t possible. Any one method is not suited to all experimental applications—researchers must decide which methods are both practical and relevant to their experimental aims.”

The references for this information are available within the manuscript.

5. The sentence "AFDW quantifies the organic or "useful" portion of the biomass by excluding contaminating inorganic salts" (Line 77-78) What is AFDW? How is this quantification?

Ash-free dry weight (AFDW) is defined in the introduction. Paragraph 4, lines 88-102 of the introduction has been expanded to include a thorough overview of the AFDW calculation.

“AFDW enables growth comparisons among a range of culture conditions, notably, between species and culture media. To calculate AFDW, first a sample of microalgal culture is concentrated, either by filtration or centrifugation, and dried. At this stage the DW can be determined. Usually, the DW sample contains at least 8-10% ash—inorganic material like salts and particulate matter. DW tracks growth trends but can be skewed if the contribution of inorganics varies. To determine AFDW density, dry biomass is combusted at high temperature, this vaporizes the organic or “useful” portion while leaving ash (inorganics) behind. To calculate the AFDW the weight of the ash fraction is subtracted from that of the DW fraction. Typically, microalgal suspensions range from 0.1-3 gAFDW per liter. Small volumes of dilute suspensions yield little dry biomass, < 10 mg.  After combustion ash may only weigh 1 mg. Therefore, depending on the culture density, this method requires volumes of 5-100 mL, and analytical scales accurate to 0.1 mg. Laboratory PBRs are typically small, a couple of liters at most, hence every liquid sample depletes culture volume. Further, the AFDW method is manual and takes 2-3 days. For replicated and repetitive experiments, an automated and continuous method is preferable.”

The references for this information are available within the manuscript.

6. Line 374 - 377, Please explain the basis for choosing this daylight intensity and initial biomass concentration (IBC) for comparison?

For the basis of providing representative results these IBCs and light intensities were chosen arbitrarily for demonstrative purposes only. This information has been added at line 403.

7. The statement, "In our experimental setup (using local temperature and pressure data) we calculated that an 80 mL headspace required 340 mL of exsolved O2 to establish an O2 partial pressure of 99% of the ambient pressure", additional basis for calculating the amount of O2 needed to establish the partial pressure in this study?

The reason this calculation is important, is that over the course of an experiment, the gas composition of the bottle headspaces will change. Initially the headspace gas composition will be identical to that of the atmosphere (~ 80% N2 and ~20% O2). “With each flip of the measurement cell, O2 displaces atmospheric gas” – Lines 416-417. This means that with time, the headspace partial pressure of O2 will increase, also increasing the DO of the medium. Here, the 340 mL of O2 required to reach 99% O2 partial pressure, is compared to the actual values produced in the study. This demonstrates to the reader, that the headspace gas composition can shift to almost pure oxygen. We have added an additional sentence to line 433 to clarify this – “Therefore, by the end of the experiment, the headspace of all bottles would have contained mostly O2.” The disadvantage of high DO to microalgal cultures and a potential solution is discussed in lines 578-597.

Reviewer #3:

Manuscript Summary:
The paper "Design, construction, and operation of laboratory photobioreactors with online growth measurements and customisable light regimes" nicely presents the build-up of laboratory photo-bioreactor that can be useful to work triplicates in parallel and that may monitor with high resolution the growth of cyanobacterial/microbial biomass. The construction is well detailed, some interesting results and discussion are also provided. The paper is well organized, interesting to be read and the discussion valuable for the following research and it will be very valuable for the filming.

Major Concerns:
My main doubt is that the scientific literature is not well considered, for example there are many examples regarding manometric (or other methods, like microcalorimetric, etc.) that have been previously applied (some systems are also commercial ones) and are not included in this manuscript. Even though the proposed system is more flexible, etc. this comparison is lacking. Most of the presented literature is updated but only demonstrates the need of the system, not the comparison with other systems.

Thank-you for this suggestion. In the introduction we have added a paragraph which considers common methods of growth measurement.

Lines 74-86 – “Growth is commonly measured by cell count, optical density (OD), chlorophyll A content, dry weight (DW) density, and ash-free dry weight (AFDW) density. Cell counts, chlorophyll A content, and gravimetric methods are manual processes which produce discrete data points. OD can be measured continuously and non-invasively with a spectrophotometer, providing it is well calibrated against another method like AFDW density. However, OD measurements and Chlorophyll A content can be unreliable as results vary under different culture conditions, e.g., between species and throughout the growth cycle. For Chlorophyll A, the extraction method used can also impact pigment yield. Chlorophyll A is typically used to assess the microalgal composition of microbial communities which also include non-photosynthetic organisms. When choosing a method to assess growth, the morphology of the suspension is also important. When organisms clump and are not-well mixed, OD and cell counts aren’t possible. Any one method is not suited to all experimental applications—researchers must decide which methods are both practical and relevant to their experimental aims.”

It is our opinion that microcalorimetric methods are not commonplace in this field and that their inclusion would be beyond the scope of this manuscript. Methods to measure photosynthetic oxygen production, including manometry, are included in lines 110-111 - “Oxygen production is measured by many different techniques: probes gauge DO concentration, manometric devices assess headspace pressure, gas chromatography measures headspace composition, and volumetric sensors record gas outflow."

We are aware that there are copious numbers of commercial PBRs and DIY prototypes available. What we believe is lacking, is a PBR that bridges the gap between simple benchtop flasks and very expensive, highly technical PBRs. We hope that this design is accessible to individuals or groups without a fabrication or engineering background. The JoVE format aids this by providing a step-by-step construction guide–something that is missing in published DIY PBR protocols. Given the sheer magnitude and variety of PBRs (each suited to a different experimental application) we believe a comparison of PBRs is best suited to a more extensive review article. We have added a paragraph at the end of the introduction which explains the gap this PBR hopes to fill. Further, we recommend that readers consider the suitability of a PBR design to their own culture conditions, financial position, and research questions.

Lines 119-130 - “PBRs in use today range in complexity. Laboratories may use simple benchtop flasks, custom prototypes, or commercially available products. For research groups seeking to upgrade from flasks, the cost of commercial PBRs or technical skill and part fabrication required to build many prototypes, may be a barrier. The goal of this manuscript is to describe the step-by-step design, construction, and operation of laboratory PBRs that bridge this gap. These PBRs have a customizable light regime and monitor growth continuously by recording volumetric oxygen production. This design houses three culture vessels for triplicate replication, can be built with moderate skill, and is made with easily accessible materials. This PBR is a valuable addition to a laboratory looking to expand their capacity for microalgal research without investing in very technical or expensive products. When choosing to acquire or build a PBR, researchers must consider the suitability of a design to their own culture conditions, financial position, and research questions.”

Minor Concerns:
Some minor comments, that the authors may consider for improving the paper quality, are given below.

1) Line 78: "contaminating inorganic salts", I suggest deleting "contaminating.

The word “contaminating” has been removed.

2) Line 79: "in our experience", information from the literature should be provided. The same for similar sentences.

We have rectified this by providing additional referenced information, particularly with regards to the volume of culture required for accurate AFDW determination. The number of cited articles for this manuscript has been increased from 14 to 27.

3) Some parts in chapters 4, 5, 7 and 8 are highlighted in yellow, why?

The journal asks that the portions of the protocol that are intended to be filmed be highlighted this way.