List of our responses to comments on an item-by-item basis

We sincerely thank the editor and the two reviewers for their valuable comments, which helped us to revise and improve the manuscript substantially. Our responses to the comments are given below.

Editorial comments

Please rephrase the Summary to clearly describe the protocol and its applications in complete sentences between 10-50 words: “The present protocol describes. …”. Here the word limit is exceeding.

We reduced the word limit to 40.

Please revise the text to avoid the use of any personal pronouns (e.g., "we", "you", "our" etc.).

We rephrased the manuscript. However, when our personal preference (and not an evidence-based method) was used, we sticked with “we”.

For in-text formatting, corresponding reference numbers should appear as numbered superscripts after the appropriate statement(s).

We changed the format.

Please ensure that the Introduction includes all of the following:

a) A clear statement of the overall goal of this method

b) The rationale behind the development and/or use of this technique

c) The advantages over alternative techniques with applicable references to previous studies

d) A description of the context of the technique in the wider body of literature

e) Information to help readers to determine whether the method is appropriate for their application.

We added the required information in the introduction.

JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials (including reagents, instruments, software, etc.). Please sort the Materials Table alphabetically by the name of the material.

We changed the manuscript accordingly.

Please adjust the numbering of the Protocol to follow the JoVE Instructions for Authors. For example, 1 should be followed by 1.1 and then 1.1.1 and 1.1.2 if necessary. Please refrain from using bullets or dashes.

We changed the manuscript accordingly.

The Protocol should contain only action items that direct the reader to do something. Please move the discussion about the protocol to the Discussion.

We changed the manuscript accordingly.

Please ensure that all text in the protocol section is written in the imperative tense as if telling someone how to do the technique (e.g., “Do this,” “Ensure that,” etc.). The actions should be described in the imperative tense in complete sentences wherever possible. Avoid usage of phrases such as “could be,” “should be,” and “would be” throughout the Protocol. Any text that cannot be written in the imperative tense may be added as a “Note.” However, notes should be concise and used sparingly.

We changed the manuscript accordingly.

Please note that your protocol will be used to generate the script for the video and must contain everything that you would like shown in the video. Please ensure you answer the “how” question, i.e., how is the step performed? Alternatively, add references to published material specifying how to perform the protocol action. There should be enough detail in each step to supplement the actions seen in the video so that viewers can easily replicate the protocol.

Please add more details to your protocol steps:

Step 1: Please include the patient inclusion/exclusion criteria.

We included point 1.1.2.

Line 149: Please mention how the anesthetization was carried out. Also, please mention how it is ensured.

Line 151: Please clarify how the dissection was carried out.

We added this information.

Line 155: Please provide additional details regarding how this step is done.

We added this information.

Line 160-172: Please subdivide this step into multiple discrete action steps clearly mentioning how each step is performed.

We subdivided the steps.

Please highlight up to 3 pages of the Protocol (including headings and spacing) that identifies the essential steps of the protocol for the video, i.e., the steps that should be visualized to tell the most cohesive story of the Protocol. Remember that non-highlighted Protocol steps will remain in the manuscript, and therefore will still be available to the reader. Please ensure that the highlighted steps form a cohesive narrative with a logical flow from one highlighted step to the next. Please highlight complete sentences (not parts of sentences). Please ensure that the highlighted part of the step includes at least one action that is written in imperative tense.

We highlighted the sections accordingly.

Please modify the Result section to include all the observations and conclusions you can derive from the Figures. The Results section should focus on the effectiveness of your technique backed up with data.

We added a paragraph in the result section.

Please include a title and a description of each figure and/or table. All figures and/or tables showing data must include measurement definitions, scale bars, and error bars (if applicable). Please include all the Figure Legends together at the end of the Representative Results in the manuscript text. Each Figure Legend should include a title and a short description of the data presented in the Figure and relevant symbols. The Discussion of the Figures should be placed in the Representative Results. Details of the methodology should not be in the Figure Legends, but rather the Protocol.

We changed the legends accordingly.

As we are a methods journal, please revise the Discussion to explicitly cover the following in detail in 3-6 paragraphs with citations:

a) Critical steps within the protocol

b) Any modifications and troubleshooting of the technique

c) Any limitations of the technique

d) The significance with respect to existing methods

e) Any future applications of the techniques

We adjusted the discussion section.

Reviewer #1

Major concern:

Please describe what approaches the surgeon can take once a drop in response is seen. Also, please describe what is a significant drop in response.

## We thank the reviewer for this important point. We added a paragraph (lines 308-314).

Minor concerns:

Line 84 - Stimulating at 1000 Hz complicates the picture as one would expect that the electrode will cross the tonotopic region of the cochlea for 1kHz during the insertion process. New systems now can stimulate at 2 frequencies (both AB and Cochlear).

That is correct. We added this information in lines 96-100.

Line 95 - Out of curiosity, how are you able to get methylprednisolone to patients 6 hrs prior to outpatient surgery?

Patients arrive at our hospital the evening before the procedure. Among other things, this has insurance-related reasons.

Line 204 - To make this protocol universally applicable, maybe conditioning of the electrode should be mentioned which is required for the Cochlear system.

We added the information in line 207-208.

Line 231 - What do you mean by 100 averages? I thought this was a continuous recording. Please clarify.

Thank you for your comment. In fact, we perform 100 epochs and form the average (mean) signal. “100 averages” was a spelling mistake. We changed this in line 262. However, by the time, continuous recordings are not possible with MED-EL implants (instead a step-by-step insertion is done, as described in lines 219-223). To measure 100 epochs and build the average signal, it takes about 30 seconds using a 9.6 ms measurement window.

Line 232 - Python is spelled incorrectly

We corrected accordingly.

Line 235 - Please describe SNR >1 or reference this. I have always seen a different description for significant signals.

Indeed, we use a different technique to calculate the SNR. As we work with raw data, we have all the single epochs available. When calculating the SNR, we work with the so-called +/- averaging method1. The +/- averaging method has the advantage that all consistent signal components can be eliminated and only the noise remains. We have determined the SNR value of 1 empirically. Hereby, an ECochG response is clearly detectable at an SNR level larger one. We added the method in line 266.

Line 271 - change to need "to" improve

We corrected the wording.

Reviewer #2

Major concerns:

In general, I feel like an overall introduction of ECochG is lacking. It would be helpful to explain in the beginning that ECochG signals are electrophysiological signals generated by the cochlea in response to sound (the sound part is completely missing from your introduction, hence the use of an acoustic stimulus may come as a surprise). Then explain that it consists of different components and introduce these. In the end, you show examples of cochlear microphonic signals but the reader may not know yet which information is represented in ECochG responses, so they do not know what cochlear microphonics are.

We thank the reviewer for this point and included the according information in lines 55-64.

Also, you are describing the measurement protocol for measuring ECochG using software from Med-EL. The procedure varies largely between manufacturers, e.g. you insert the implant in a step-wise manner while for other manufacturers, ECochG is measured continuously during insertion. The duration of the measurement and the flexibility of settings in the software also differs between manufacturers. I would explain the different procedures for measuring ECochG in your introduction, and then mention that your procedure describes ECochG measurements for recipients of Med-EL implants. Are you using a research version of the software or a clinical tool?

We have rewritten the manuscript (lines 78-83, lines 223-230) to include the information for the three common manufacturers (MED-EL, AB, Cochlear). Please also see our new table 1. For MED-EL, to record the data, we used Maestro 8.03 AS, which is a research software.

Line 58, 'Inner ear potentials (i.e. electrocochleography, ECochG) are increasingly measured during the implantation process (real-time measurements, rt) to monitor the inner ear function.'

Comment: Later on (line 270), you state that we are only beginning to understand the actual implications of ECochG changes. I therefore think it is too strong to state here that ECochG is increasingly being used to monitor the inner ear function. Atraumatic implantation still mostly depends on the surgeon's judgement and ECochG is currently mostly measured for research purposes.

We rephrased the sentence in lines 55-59.

There are groups that show good results with using ECochG for providing surgical feedback. These could be cited here.

We included references in line 56. Please also consider the new paragraph lines 308-314.

Line 62, 'changes of the ECochG signal are correlated with the residual function of inner ear function'. Comment: To me it is not clear what you mean with this sentence. Intra-OP ECochG changes correlate to pre-operative residual hearing status?

We complemented the sentence, lines 66-67.

Line 64, 'In up to 20% of cases, no interpretable signal can be derived'. Comment: This percentage depends highly in your inclusion criteria for measuring ECochG. I would include a statement like 'despite residual hearing of at least …'. Also, please specify which patient-specific factors you are referring to.

In our experience, the threshold of residual hearing cannot alone explain the success rate of ECochG. Some patients with very good residual hearing have no or only weak signals whereas other patients, who are borderline candidates for ECochG, have strong signals with large amplitudes (also see our revised table 2). Patient-specific factors are mainly the presence or absence of functioning hair cells. We added information to the sentence in line 69-71.

Line 80, 'We perform rt-ECochG measurements in patients where hearing preservation is the goal.' Comment: I would be more specific about the patients in whom hearing preservation is the goal. What are your audiological criteria (e.g. PTA)?

Please consider paragraph 1.1.2.

Line 82, 'Usually, we stimulate with a 500Hz pure tone.'

What are your stimulation settings? How many repetitions, repetition rate, recording window, nr. of averages etc. Here it could be helpful to show which flexibilities your software has and which settings you choose.

Thank you for your questions. We added the configuration settings in lines 101-105.

Line 95-98

Comment: Is point 4 related to ECochG or part of standard clinical practice regardless of whether ECochG is performed? In the prior case, I would state the relevance for measuring ECochG.

It is relevant for both. For ECochG measurements, an unimpeded stimulus transmission is key. Please consider the new paragraph lines 247-260.

Line 106-107

Comment: How did you obtain or sterilize the sterile sound tube and foam eartip. From our experience, this is not trivial.

We sterilized the sound tube and eartip (Etymotic, USA) by gas plasma sterilization.

Line 113-115

Comment: This depends on the setup in the OR and the space requirements. Whether the measurement process can be monitored well also depends on e.g. the placement of screens, which differs between ORs.

This is only a recommendation. In our experience, the communication between engineer and surgeon is crucial. The setup in the OR can have an influence on this. For this reason, we consider the figure to be a helpful illustrative example.

Line 177-178

Comment: Can you explain the relevance when measuring ECochG? I can imagine the anesthesiologist is monitoring many parameters and I am wondering why the blood pressure should be assessed specifically.

To minimize bleeding, if possible, the systolic blood pressure should be below 100mg Hg. We added this information line 188-189.

Line 200

Comment: Which value should the impedance be below to be low enough?

Impedance values are manufacturer-specific. As a rough guide, the impedance should be below 10 kΩ. We added this to the manuscript in line 212-213.

Line 205-209

Comment: why is this step-wise insertion related to the type of implant? In case this is because of the measurement software, I would explain that here and write in your introduction that ECochG can be performed continuously during insertion or in a step-wise manner. Then I would include a rationale for both and include references for both protocols. As the surgeon has to hold the implant still for your protocol, could you mention the duration of one measurement?

Please consider the new subsections 5.5.1. and 5.5.2. We included the requested references.

The duration of a step-wise measurement using a MED-EL implant is approximately 30 seconds for each step.

Line 225-236

Comment: I find the results section difficult to interpret in the overall context of your text. I have a few suggestions for improvement:

- 'for signal processing, we focused on cochlear microphonic signals'. Cochlear microphonic signals have not been introduced in your text so far. Maybe a short introduction of the different ECochG signal components in your introduction could help. I would also say you show an example of these signals and why you did so (because they show the best relation to post-OP hearing outcomes?), not that you focused on them for signal processing.

We have revised the introduction section (see comment above) as well as the results section. An example of our recordings is shown in figure 3. As outlined in the introduction, the cochlear microphonic signal is usually employed as it is the most robust signal with (usually) the largest amplitudes.

What is the reason you are post-processing your signals in Python? Bandpass filtering is normally done in the clinical software. Does the Med-EL software not have this option or did you decide to perform additional processing?

No, we obtain raw signals from the device and process the data later on.

You apply bandpass filtering, but you are not mentioning the cutoff frequencies.

Thank you for the question, we added the cut-off frequencies in lines 268-269 (100 Hz / 3 kHz).

Why are you using an SNR of one, do you have a reference of a rationale behind this? I think it is more common to use a higher SNR for considering a measurement as valid. Also, please explain how your SNR is calculated.

We use a different technique to calculate the SNR as we have access to the single epoch recordings. When calculating the SNR, we work with the so-called +/- averaging method1. The +/- averaging method has the advantage that all consistent signal components can be eliminated and only the noise remains. We have determined the SNR value of 1 empirically. In our experience, using this SNR, an ECochG response is (visually) clearly detectable.

To me it is not clear why you are using all these Python packages for signal processing, if you are only bandpass filtering your signal. Sklearn is a machine learning package and is not needed for this. Mentioning Python in your manuscript may confuse researchers from a more clinical background who could easily perform ECochG using standard settings, but are not familiar with the use of Python for post-processing. Since you are establishing a standard procedure, I would try to keep it simple. In case these Python scripts are provided together with the clinical software, please mention this.

Implementing the +/- averaging method1 required us to use these Python packages. Numpy and Scipy are used only for bandpass filtering. Sklearn (sklearn.model_selection.train_test_split) is used only for randomly splitting the individual measurement signals into two groups. We then use these two groups to calculate the SNR. We repeat this process 1000 times and calculate the mean SNR. We repeat the SNR calculation 1000 times as there is a small variance in each SNR which is caused by the small number of epochs (100).

Line 242-244 (Figure 2)

Comment: This figure of the ear could be replaced by a figure in which the earfoam is inserted in the ear canal and the ear is folded forward. This would be more representative of the ECochG setup.

We changed figure 2 accordingly.

Line 246-250 (Figure 3)

Comment: I would replace this figure with some ECochG responses that you measured during insertion, instead of afterwards, as your protocol describes intra-OP monitoring. Then, I would also explain what can be seen in the figure. For example, why does the amplitude change? What should the surgeon do when this happens?

Figure 3a recordings were taken during the electrode insertion (at different insertion depths). In figure 3b, ECochG recordings were recorded after full insertion from electrode 1, 4, 7, and 10 (counted from the most apical electrode). This study was purely observational. At this stage, we do not give a recommendation what to if a drop occurs. Drops of the ECochG signal may have different causes (e.g. traumatic event during the insertion, passing the 500 Hz region within the cochlea, interaction of two hair cell generators). Please also consider our new section lines 309-115.

Are you sure that the figure shows the cochlear microphonic signal and not just the condensation or rarefaction response? At least in earlier versions of the Med-EL software, cochlear microphonics could not be visualized directly.

Yes, you are right; the standard software only allows the selection of condensation or rarefaction. However, we developed a custom software, which enables to record with alternating stimulation.

Line 248-249, 'Please note that the numbering of electrodes for figure a & b starts at opposite ends.' Comment: Do you mean that for Figure A the apical electrode is EL 1 while for Figure B the basal electrode is EL 1? This does not seem very logical to me.

As outlined above figure 3a shows the insertion depth; figure 3b, displays the electrode position, counted from the most apical one (as suggested by the manufacturer). We added lines 288-290 for better explanation.

Line 252

Comment: Please explain this table. It cannot be interpreted without context. How should we look at the change in PTA with respect to the ECochG amplitude change? Are your post-operative measurements really post-operative or after insertion but in the OR? What should we learn from the SNR - is there much noise in the OR, is it related to the signal amplitude? Cochlear access (RW or cochleostomy) is unrelated to ECochG.

## We thank the reviewer for this point and have revised the table.

## The new table 2 shows the maximum SNR reached during the electrode insertion and where this maximum was reached. Further, we added the SNR at the peak electrode after full insertion. It should now be clear i) where the maximum amplitude was reached, and ii) if there was a drop of this amplitude when the electrode was fully inserted. This change of the SNR can now be compared to the change of the PTA. Furthermore, as outlined above, it becomes evident, that the measured SNR cannot be correlated to the pre-operative PTA (“the better the acoustic hearing before surgery, the higher the SNR of the ECochG recording”). We mentioned the type of surgical access as this is an important information in regards to hearing preservation. It has been shown that hearing preservation is much more likely in patients with round window access2.

Minor concerns:

Minor Concerns:

Line 87

Comment: between 80 and 85 dB HL?

We included this information.

References

1.	Drongelen, W. van. Chapter 4 - Signal Averaging. in Signal Processing for Neuroscientists (Second Edition) (ed. Drongelen, W. van) 59–80 (Academic Press, 2018). doi:10.1016/B978-0-12-810482-8.00004-7.

2.	Causon, A., Verschuur, C. & Newman, T. A. A Retrospective Analysis of the Contribution of Reported Factors in Cochlear Implantation on Hearing Preservation Outcomes. Otol Neurotol 36, 1137–1145 (2015).