Title: 	In vivo imaging of biological tissues with combined two-photon fluorescence and stimulated Raman scattering microscopy

Author: 	Wanjie Wu, Xuesong Li, Jianan Y. Qu and Sicong He,

Dear Editor,

Thank you for considering our manuscript entitled “In vivo imaging of biological tissues with combined two-photon fluorescence and stimulated Raman scattering microscopy” for publication in Journal of Visualized Experiments. We are most grateful to the reviewers for their time in reviewing our manuscript. We appreciate the reviewers’ valuable and constructive comments to improve the quality of our manuscript.

We have revised the manuscript carefully and all the concerns from the reviewers have been addressed in the revised manuscript or in the responses to reviewers’ comments. Enclosed please find the detailed responses (in black text) to the reviewer’s comments (in blue text) and the indications of revisions in the manuscript (highlighted in grey). All the corresponding changes in the revised manuscript have been indicated as red text for your and the reviewer’s convenience to locate the revisions.

Sincerely,

Sicong HE

Responses to Editorial comments

The authors agree with all the issues raised by the editor and all the concerns from the editor have been addressed in the revised manuscript or responses. We appreciate the editor’s suggestions to improve the quality of our paper. The responses and revisions based on comments received are as follows:

Comment 1: 	Please take this opportunity to thoroughly proofread the manuscript to ensure that there are no spelling or grammar issues.

Response:  	Thanks for the suggestion. Spelling and grammar issues were checked and corrected throughout the manuscript.

Comment 2: 	Please provide citations for the following lines: 85-91.

Response:  	According to Reviewer #2 and Reviewer #3’s suggestions that it is better to focus on one biological application, we removed the contents of adipose tissue imaging in the revised manuscript.

Comment 3: 	Please define all abbreviations upon first use. For example: OPO, OM, etc.

Response:  	Thanks for careful reading of our manuscript. We defined all abbreviations in the revised manuscript.

Revision:  	Page 3, line 92-93: “1.1.	Use an integrated optical parametric oscillator (OPO) connected with a mode-locked Ytterbium fiber laser as the picosecond (ps) laser source for SRS imaging.”

Page 3, line 96-98: “The Stokes beam is modulated at 20 MHz by a built-in electro-optical modulator (EOM) for SRS high-frequency (MHz) phase-sensitive detection.”

Page 6, line 228-231: “2.2.2.2. Place an alignment plate P0 before L3 (Figure 1). Use an M2 hex key to adjust the optical mirror 1 (OM1) so that the Stokes beam center passes the through-hole of the alignment plate at the laser output port. Use the IR scope to confirm the position of the beam spot at P0 during adjustment.”

Comment 4: 	The Protocol should contain only action items that direct the reader to do something.

Response:  	Thanks for the suggestion. We reorganized the protocol accordingly.

Comment 5: 	JoVE cannot publish manuscripts containing commercial language. This includes trademark symbols (™), registered symbols (®), and company names before an instrument or reagent. Please remove all commercial language from your manuscript and use generic terms instead. All commercial products should be sufficiently referenced in the Table of Materials. For example: Kwik-Sil, etc.

Response:  	We have removed all commercial language in the revised manuscript.

Comment 6: 	Please revise the text to avoid the use of any personal pronouns (e.g., "we", "you", "our" etc.).

Response:  	All personal pronouns were revised accordingly.

Comment 7: 	Please add more details to your protocol steps. Please ensure you answer the “how” question, i.e., how is the step performed?

Step 1.1.3: How was the output fed into the oscilloscope?

Step 1.2.1.1: How was the pump wavelength and power set? Please provide all associated steps.

Step 1.2.1.3: How was the positioning of I1 and I2 determined? How was the position of the laser beam checked?

Step 1.3.1.4: How were the imaging parameters set? What was the value used for the current experiment?

Step 1.3.1.5, 1.3.1.6: What are the various steps followed to tune the amplifier phase and to set the SRS signal to zero? Please mention all associated steps.

Line 284: Please shift the ethical statement to the beginning of the protocol.

Step 2.1.3, 2.2.1: How was proper anesthetization confirmed? Please specify the use of vet ointment on eyes to prevent dryness while under anesthesia.

Step 2.2.2: How was the mouse skin removed? Please mention all steps.

Step 3.1.19: Please mention the suture size.

Response:  	Thanks for the comments and suggestions. We added more details to the protocol steps in the revised manuscript.

Revision:  	Page 4, line 158-160: “Connect the output port of the photodetector to the input port of an oscilloscope using a coaxial cable with Bayonet Neill-Concelman (BNC) connector to monitor the laser pulse.”

Page 5, line 177-178: “Use OPO control software to set the wavelength of the pump beam to 796 nm, and set its power to the minimum (~20 mW) for optical alignment.”

Page 5, line 186-187: “2.2.1.3 Place two alignment plates (P1 and P2 in Figure 1) on the optical path. Place P1 behind PBS at a distance of about 10 cm and place P2 behind P1 at a distance of about 30 cm.”

Page 5, line 194-195: “2.2.1.5. Adjust mirror M1 to locate the ps laser beam center at the through-hole of P1. Use an IR scope to observe the position of the beam spot at P1 when adjusting the mirror M1.

Page 5, line 197-198: “2.2.1.6 Adjust mirror M2 to locate the ps laser beam center at the through-hole of P2. Use the IR scope to observe the position of the beam spot at P2 when adjusting the mirror M2.”

Page 6, line 257-258: “2.3.1.2 Set the wavelength of the pump laser to 796 nm and the power of the pump and Stokes beam to be 15 mW and 25mW after the 25× objective lens, respectively.

Page 6-7, line 264-265: “2.3.1.3 Place the olive oil sample on the stage and adjust the focus of 25× objective into the sample.”

Page 7, line 270-292: “2.3.1.4 Use the microscope control software to set the imaging parameters and use the lock-in amplifier control software to set a time constant value close to the pixel dwell time.

Note: the parameters for SRS imaging of olive oil are 512×512 pixels, 500 μm×500 μm field of view (FOV), 6.4 μs pixel dwell time, 10 μs time constant.

2.3.1.5 Scan the laser beam over the sample. Use the lock-in amplifier control software to tune the phase (0-180°) with a step size of 22.5° until the SRS signal intensity reaches the maximum.

Note: In this all-analog lock-in amplifier, the signal output is the In-Phase component which is dependent on the phase of the reference signal. The phase can be adjusted by the lock-in amplifier control software with 11-degree resolution, allowing to maximize the detected signal to within ~2%12. The lock-in amplifier gets out of phase once the sync signal is disrupted. The phase needs to be readjusted every time the sync signal is reestablished.

Step 2.3.1.6 Scan the sample with laser shutter closed. Use the lock-in amplifier control software to tune the offset value with a step size of 1mV until the average SRS signal is close to zero.

Note: the average SRS signal is estimated as the average intensity of all pixels in the SRS image, which is calculated automatically by the microscope control software. Offsets are useful for canceling unwanted phase-coherent signals.”

Page 8, line 346-347: “Perform toe pinching to ensure deep anesthesia of the mouse. Supplement with half the original dose of the anesthetics as needed.”

Page 8, line 352: “3.1.5 Apply ointment to the mouse eye to prevent corneal dryness.”

Page 11, line 447: “4.19 Suture the skin with #6-0 surgical suture to close the wound.”

Comment 8: 	Please ensure that the highlighting is not more than 3 pages including headings and spacing

Response:  	Thanks for the suggestion. We adjusted the highlighting contents to be less than 3 pages in the revised manuscript.

Comment 9: 	Figures: Please mention all abbreviations used in the figure legends.

Response:  	Abbreviations in the figure legend were revised accordingly.

Revision:  	Figure 1 caption: “Figure 1. The schematic diagram of the TPEF-SRS microscope system. Pump and Stokes beams are combined with a dichroic mirror (D1) in the picosecond (ps) laser. The ps beam and femtosecond (fs) beam are collimated and expanded/narrowed by a pair of lenses (L1, L2, and L3, L4 respectively) to match the 3 mm XY-scan galvo mirrors. The fs beam is rotated from horizontal to vertical polarization by a half-wave plate (HWP) and then combined with the ps beam by a polarizing beam splitter (PBS). The scanning mirrors and the rear pupil of the objective lens are conjugated by a telecentric scan lens L5 and an infinity-corrected tube lens L6. The laser beam is expanded by the scan and tube lens to fill the back aperture of the 25× objective. For stimulated Raman scattering (SRS) imaging, the backscattered pump beam collected by the objective is reflected by a PBS and directed to a large area (10mm×10mm) Silicone photodiode (PD). For two-photon imaging, the two-photon excited fluorescence (TPEF) signal is reflected by a dichroic beam splitter D2 to the photodetection unit. A current photomultiplier (PMT) module is used to detect the TPEF signal. L1-L10: lenses; OL: objective lens; D1-D3: dichroic mirrors; Fs1, Fs2: filter sets; M: mirrors; OM1, OM2: optical mirrors; P0-P2: alignment plates.”

Comment 10: Figure 3: Please provide scale bar for all microscope images.

Response:  	Scale bars were added to all microscope images in the revised manuscript.

Comment 11: Please also include in the Discussion the following in detail along with citations:

a) Critical steps within the protocol

b) Any limitations of the technique

Response:  	Thanks for the suggestions. We added discussions along with citations on the critical steps within the protocol and limitations of this technique in the revised manuscript.

Revision:  	Page 12, line 512-526: “For SRS imaging, though pump and Stokes beams are temporally and spatially overlapped inside the OPO, this overlapping can be disrupted after passing through the microscope system. Therefore, both spatial and temporal optimization of the colocalization of the pump and Stokes beams is necessary and critical to achieve optimal SRS imaging. Phase and offset adjustment are also important to get the maximal SRS signal output with the all-analog lock-in amplifier setup12. In addition, initial calibration of the temporal delay between the Stokes and pump beams at different wavelengths would be helpful for fast SRS spectra sweeping and imaging at various Raman shifts. For TPEF-SRS imaging, strict spatial overlapping of the ps and fs beams after the combination is important to avoid any FOV shift when switching the two imaging modes. Therefore, in this TPEF-SRS microscope, totally three beams are involved in the spatial colocalization. Firstly, the ps pump beam and fs beam are aligned, to make sure they both pass the center of the microscope cage system. Then using the pump beam as a reference, the Stokes beam position is adjusted accordingly to achieve strict spatial colocalization. Each alignment procedure requires several trials to reach the optimum.”

Page 12-13, line 530-541 “In this protocol, TPEF and SRS imaging at the same location are performed sequentially with 1 s interval. Motion artifacts during spinal cord imaging may cause an imperfect merge of the TPEF and SRS images, which limits the capability of this method for imaging highly dynamic processes or tissues largely affected by the breath and heartbeat of animals. One possible solution is to collect the ps laser excited two-photon fluorescence during SRS imaging9. However, this method is only applicable to the biological structures with strong fluorescence signals since the ps pulse has a much lower excitation efficiency compared to the fs pulse20. Alternatively, the motion problem can be solved by using a fs-SRS system21, 22, where the fs laser source allows exciting the SRS and TPEF signals simultaneously in an efficient way, at the expense of low spectral resolution of SRS imaging. As shown in Figure 3, this registration strategy works well when no significant motion occurs during the SRS and fluorescence imaging.”

Responses to reviewers’ comments

The authors agree with most of the issues raised by the reviewers and all the concerns from the reviewers have been addressed in the revised manuscript or responses. We appreciate the reviewers’ suggestions to improve the quality of our paper. The responses and revisions based on comments received are as follows:

Responses to the comments from reviewer #1:

General: 	A number of papers concerning SRS and TPE implementations have already reported in JOVE and the combination of these two techniques, using two different laser system, is not challenging and above all no significant news are proposed. Therefore, in my opinion, this paper doesn't deserve to be published in a journal whose aim is to bring to life details of cutting edge technologies. Concerning TPE implementation no new details in its implementation with respect to other ones are pointed out. For SRS, being the two difficult conditions to get in SRS implementation, i.e. spatial and temporal alignment, essentially provided by the laser system, the implementation presented is a simplified one. Details about the implementation of high-frequency modulation transfer method are missing. SRS is a shot noise limited technique, but this crucial point is not discussed into the paper and no evidence are reported that this condition is achieved. In addition the synchronization of the signal acquired by the phase-sensitive detection scheme, with the beam positioned onto the sample monitored by the scan head of the microscope is completely missing. This is a typical obstacle for SRS implementation. Again, no details about microscope are provided, if it is a home built or a commercial system. Of course this difference is important. If the used microscope is a commercial one, the only key experimental point would be colocalization of the ps and fs beams, but this is not enough for publication. Details about images time acquisition pixel dwell time and integration time should be reported to demonstrate that the techiniques are fast enough to allow in vivo imaging.

Response:  	We appreciate the reviewer's thorough assessment of our manuscript. We have followed the reviewer’s comments to revise our manuscript accordingly.

Comment 1:  In section 1.2.3 Setup of the detection optical path and electronics. The Filter should be linked to the list of material

Response:  	Thanks for the suggestion. All the filters have been listed in the Table of Materials with the same name as in the protocol and figure for cross reference.

Comment 2:  In section 1.3.1 Phase adjustment of the lock-in amplifier. The results obtained by lock in should not depend on phase. Why is this optimisation process necessary?

Response:  	Good question. The lock-in amplifier (LIA) used in our microscope system is an all-analog device designed based on the Ref below. This LIA is designed for high-speed SRS imaging and thus has no digital signal processing module to directly calculate the signal amplitude as other LIA device (e.g. SR844, Stanford Research Systems) does. Therefore, the output signal is the In-Phase component, which is dependent on the phase difference between the signal and reference. In our case, the phase of the reference is adjusted using the LIA control software so that the maximal SRS signal can be obtained.

Ref:

Saar, B.G., Freudiger, C.W., Reichman, J., Stanley, C.M., Holtom, G.R., Xie, X.S. Video-Rate Molecular Imaging in Vivo with Stimulated Raman Scattering. Science. 330 (6009), 1368–1370, doi: 10.1126/science.1197236 (2010).

Revision:  	Page 7, line 280-284: “In this all-analog lock-in amplifier, the signal output is the In-Phase component which is dependent on the phase of the reference signal. The phase can be adjusted by the lock-in amplifier control software with 11-degree resolution, allowing to maximize the detected signal to within ~2%12. The lock-in amplifier gets out of phase once the sync signal is disrupted. The phase needs to be readjusted every time the sync signal is reestablished.”

Comment 3:  In section 1.3.1.6 Scan the sample with laser shutter closed. Tune the offset value until the average SRS signal is close to zero. How this average SRS signal is estimated? Do you measure Noise with LIA or calculate noise from images?

Response:  	The average SRS signal was estimated as the average intensity of all pixels in the SRS image, which was calculated automatically by the control software when the SRS image was captured. We didn’t measure noise with LIA or from images.

Revision:  	Page 7, line 286-292: “Scan the sample with laser shutter closed. Use the lock-in amplifier control software to tune the offset value with a step size of 1mV until the average SRS signal is close to zero.

Note: the average SRS signal is estimated as the average intensity of all pixels in the SRS image, which is calculated automatically by the microscope control software. Offsets are useful for canceling unwanted phase-coherent signals.”

Comment 4:  The delay between pump and probe should not depend on wavelength. If not, explain why?

Response:  	Thanks for pointing out the issue. The delay between pump and probe laser depends on the optical path difference (OPD):

OPD=( nprobe-npump)×L

where nprobe and npump are the refractive index of optical components at the wavelengths of probe and pump lasers, respectively, while L represents the length of optical path or thickness of optical components in the microscope system. As the refractive index of glass changes with the light wavelength [Ref below], the OPD varies when we adjust the probe wavelength while maintaining the pump wavelength (1031nm). Therefore, the delay between the two beams depends on wavelength and needs to be calibrated prior to SRS imaging.

Ref:

1. Chartier, Germain. Introduction to optics. Springer Science & Business Media, 2005. (Chapter3.3, page 63-64)

Revision:  	Page 7, line 296-297: “Note: The delay between the pump and Stokes beam is related to the optical path difference, which is determined by the dispersion of optical components in the microscope system.”

Comment 5:  In discussion, you claim: As a result, though the backscattered signal collection by objective is less efficient than that by directly objective placing an annular photodetector in front of the SRS images (512 X 512 pixels ) of the spinal cord and adipose tissue with a high signal-to-noise ratio can 529 be acquired with 1~2 s integration time. The point is that if you use a PBS, you detect a scattered beam with a different polarization with respect to the incident beam. Please explain how this different polarization can influence the efficiency of SRS detected signal.

Response:  	Good question. Since the backscattered SRS signal has various polarizations, we used a PBS to distinguish the incident light and backscattered light. The polarization axis of PBS is parallel to the incident lasers to maximize the transmittance of excitation light. The SRS signal, backscattered from the tissue, has shifted polarization, and can thus be partially reflected by the PBS and collected by the photodiode. In contrast, the annular photodetector can collect most of backscattered SRS signals regardless of polarization, thus allowing higher collection efficiency.

Revision:  	Page 3, line 121-122: “Part of the back-scattered SRS signal is lost when passing through the PBS as a result of its randomly shifted polarization.”

Comment 6:  The aim of your work is to perform in vivo experiment, but your experimental set up imposes that TPEF and SRS imaging were performed sequentially at the same location. This a fundamental limitation for in vivo measurements and I think that you porposed solution is not a suitable relief.

Response:  	Good question. We used a motorized flipper and custom software to switch the imaging models. The total time for sequential TPEF imaging, model switching, and SRS imaging are about 3s. For in vivo imaging in brain or other tissues which are less affected by the breath and heartbeat of the mouse, this sequential imaging approach should not be a problem since the 1-s switch time  is short enough to avoid inter-frame motion artifacts. In contrast, spinal cord imaging is largely affected by the breath and heartbeat motion, and we indeed observe imperfect merge between SRS and fluorescence images, especially when the mice were not well anesthetized. One possible solution is to collect the ps laser excited two-photon fluorescence during SRS imaging simultaneously. However, because ps pulse has much lower excitation efficiency compared to the fs pulse, this solution is only applicable to the biological structures with strong fluorescence signal like the YFP-labeled axons in our case. Alternatively, if spectral resolution of SRS imaging is not highly demanded, a fs-SRS-TPEF system can be used to capture the SRS and fluorescence images simultaneously. Another solution is to use the ps laser excited fluorescence obtained during SRS imaging as a reference to register the fs fluorescence images. As shown in Figure 3, this registration strategy works well if no significant motion occurs during the SRS and fluorescence imaging. We have added the discussion in the revised manuscript.

Revision:  	Page 13, line 530-541: “In this protocol, TPEF and SRS imaging at the same location are performed sequentially with 1 s interval. Motion artifacts during spinal cord imaging may cause an imperfect merge of the TPEF and SRS images, which limits the capability of this method for imaging highly dynamic processes or tissues largely affected by the breath and heartbeat of animals. One possible solution is to collect the ps laser excited two-photon fluorescence during SRS imaging9. However, this method is only applicable to the biological structures with strong fluorescence signals since the ps pulse has a much lower excitation efficiency compared to the fs pulse20. Alternatively, the motion problem can be solved by using a fs-SRS system21, 22, where the fs laser source allows exciting the SRS and TPEF signals simultaneously in an efficient way, at the expense of low spectral resolution of SRS imaging. As shown in Figure 3, this registration strategy works well when no significant motion occurs during the SRS and fluorescence imaging.”

Comment 7:  In introduction, the first paragraph concerning comparison of SRS with CARS and Raman spectroscopy should be removed. This comparison has been reported too many times in literature (see for example several SRS review papers). A preliminary description of step by step implementation should be added, in which missing crucial points are pointed out and above all what are the novelty of this combination of techniques should be pointed out.

Response:  	Agree. We removed the contents on the comparison of SRS with CARS and Raman spectroscopy. More details on the implementation of the home-built microscope system were added in the revised manuscript.

Revision:  	Page 2-4, line 81-146: “

Note: All animal procedures performed in this work are conducted according to the guidelines of the Laboratory Animal Facility of the Hong Kong University of Science and Technology (HKUST) and have been approved by the Animal Ethics Committee of HKUST. Safety training for laser handling is required to set up and operate the TPEF-SRS microscope. Always wear laser safety goggles with appropriate wavelength range when dealing with laser.

- Instrumental setup of the TPEF-SRS microscope.

Note: The setup schematic of the home-built TPEF-SRS microscope system is shown in Figure 1. This part introduces major procedures for the setup of the TPEF-SRS microscope system

  - Use an integrated optical parametric oscillator (OPO) connected with a mode-locked Ytterbium fiber laser as the picosecond (ps) laser source for SRS imaging.

Notes: OPO outputs a Stokes beam (1031nm) and a pump beam (tunable from 780nm to 960nm) with 2-ps pulse duration and 80 MHz repetition rate. The Stokes beam is modulated at 20 MHz by a built-in electro-optical modulator (EOM) for SRS high-frequency (MHz) phase-sensitive detection.

  - Use a femtosecond (fs) Ti: sapphire laser as the laser source for TPEF imaging.

  - Use a pair of lenses (L1 and L2) to collimate and adjust the size of fs beam to be 3 mm.

  - Use a pair of achromatic doublets (L3 and L4) to collimate the ps laser beam and expand its diameter to be 3 mm.

  - the polarization of fs laser beam from p polarization to s polarization using a half-wave plate.

  - Combine the two laser beams with a polarizing beam splitter (PBS).

  - Add a pair of 3 mm XY-scan galvo mirrors behind the PBS for beam scanning.

  - Use a telecentric scan lens L5 and an infinity-corrected tube lens L6 to conjugate the scanning mirror and the rear pupil of the 25× objective lens. The laser beam is also expanded by the scan lens and tube lens to fill the back aperture of the objective.

  - Place a PBS or dichroic mirror (D2) between the tube lens and objective lens for SRS or fluorescence signal collection. PBS and D2 are switchable by using a motorized flipper.

Note: part of the back-scattered SRS signal is lost when passing through the PBS as a result of its randomly shifted polarization.

  - Use a pair of lenses (L7 and L9) to narrow the detection beam and conjugate the rear pupil of 25× objective lens with a photodiode sensor.

  - Use a pair of lenses (L8 and L10) to narrow the detection beam and conjugate the rear pupil of 25× objective with the detection surface of the photomultiplier (PMT).

  - Use a dichroic mirror (D3) to separate the detection path of fluorescence and SRS signals.

  - Place a filter set (Fs1) in front of the photodiode detector to block the Stokes laser and pass the pump beam only.

  - Place a filter set (Fs2) in front of the PMT detector to pass the target fluorescence signal only.

  - Connect the PMT to a current amplifier for signal amplification.

  - Connect the photodiode to a lock-in amplifier.

  - Connect the sync signal from the built-in EOM output to the reference input of the lock-in frequency for SRS signal demodulation.

  - Connect the PMT amplifier and lock-in amplifier outputs to the data acquisition module.”

Responses to the comments from reviewer #2:

General: 	In this manuscript, Wu and colleagues aim to illustrate two methods (1) how to setup a homemade system combining 2P fluorescence and SRS techniques and, (2) how to record images of biological tissues in living animals. Overall, these methods are of general interest for the engineers developing noncommercial systems and biologists, respectively.

My main concern comes from this dual goal in a limited space (the article by itself) and time (the expected video), mainly because important information, explanations and experimental details are lacking for setting up the microscope and recording the data. One option would have been to choose either method, but not both at the same time.

Response:  	We appreciate the reviewer's thorough assessment of our manuscript and positive opinion about our work. We have followed the reviewer’s comments to revise our manuscript accordingly.

Comment 1:  The described procedure is highly dependent on the design homemade installation. Consequently, the authors have to provide a detailed protocol for building this microscope, but the current version is not explicit enough. At a first step, one should expect a thorough explanation of the function of the different constituents and a description of the procedure to assemble them. This part should not necessarily be illustrated in the video, but only detailed in the text.

Response:  	Thanks for the reviewer’s suggestions. We added more detailed descriptions on the implementation of our home-built microscope system in the revised manuscript.

Revision:  	Page 2-4, line 81-146: “

Note: All animal procedures performed in this work are conducted according to the guidelines of the Laboratory Animal Facility of the Hong Kong University of Science and Technology (HKUST) and have been approved by the Animal Ethics Committee of HKUST. Safety training for laser handling is required to set up and operate the TPEF-SRS microscope. Always wear laser safety goggles with appropriate wavelength range when dealing with laser.

- Instrumental setup of the TPEF-SRS microscope.

Note: The setup schematic of the home-built TPEF-SRS microscope system is shown in Figure 1. This part introduces major procedures for the setup of the TPEF-SRS microscope system

  - Use an integrated optical parametric oscillator (OPO) connected with a mode-locked Ytterbium fiber laser as the picosecond (ps) laser source for SRS imaging.

Notes: OPO outputs a Stokes beam (1031nm) and a pump beam (tunable from 780nm to 960nm) with 2-ps pulse duration and 80 MHz repetition rate. The Stokes beam is modulated at 20 MHz by a built-in electro-optical modulator (EOM) for SRS high-frequency (MHz) phase-sensitive detection.

  - Use a femtosecond (fs) Ti: sapphire laser as the laser source for TPEF imaging.

  - Use a pair of lenses (L1 and L2) to collimate and adjust the size of fs beam to be 3 mm.

  - Use a pair of achromatic doublets (L3 and L4) to collimate the ps laser beam and expand its diameter to be 3 mm.

  - the polarization of fs laser beam from p polarization to s polarization using a half-wave plate.

  - Combine the two laser beams with a polarizing beam splitter (PBS).

  - Add a pair of 3 mm XY-scan galvo mirrors behind the PBS for beam scanning.

  - Use a telecentric scan lens L5 and an infinity-corrected tube lens L6 to conjugate the scanning mirror and the rear pupil of the 25× objective lens. The laser beam is also expanded by the scan lens and tube lens to fill the back aperture of the objective.

  - Place a PBS or dichroic mirror (D2) between the tube lens and objective lens for SRS or fluorescence signal collection. PBS and D2 are switchable by using a motorized flipper.

Note: part of the back-scattered SRS signal is lost when passing through the PBS as a result of its randomly shifted polarization.

  - Use a pair of lenses (L7 and L9) to narrow the detection beam and conjugate the rear pupil of 25× objective lens with a photodiode sensor.

  - Use a pair of lenses (L8 and L10) to narrow the detection beam and conjugate the rear pupil of 25× objective with the detection surface of the photomultiplier (PMT).

  - Use a dichroic mirror (D3) to separate the detection path of fluorescence and SRS signals.

  - Place a filter set (Fs1) in front of the photodiode detector to block the Stokes laser and pass the pump beam only.

  - Place a filter set (Fs2) in front of the PMT detector to pass the target fluorescence signal only.

  - Connect the PMT to a current amplifier for signal amplification.

  - Connect the photodiode to a lock-in amplifier.

  - Connect the sync signal from the built-in EOM output to the reference input of the lock-in frequency for SRS signal demodulation.

  - Connect the PMT amplifier and lock-in amplifier outputs to the data acquisition module.”

Comment 2:  Actually, the procedure to align and optimize the setup is unclear. For instance, line 136: "start the microscope control software": is it a homemade interface or a commercial one? Line 180: where are OM1 and OM2 in figure 1; Line 184, "remove the iris": Which one? Line 204: this note comes too late; Line 220-221: for what objective lens, 10X, 25X? Etc…

Response:  	Thanks for pointing out the issues. The microscope control software is a homemade interface developed using C# language.

OM1 and OM2 are two optical mirrors inside the optical parametric oscillator (OPO), which have been added in the revised Figure 1.

The iris mentioned in Line 184 was placed close to the OPO output. We described the iris as alignment plate which should be more appropriate and mark it as P0 in the revised Figure 1.

The content in Line 204 has been moved to the part 1 of the protocol.

The objective mentioned in Line 220-221 is 25×.

We revised the manuscript according to the reviewer’s comments and added detailed descriptions on the procedures of system alignment and optimization.

Revision:  	Page 5, line 184: “Note: the microscope control software is a homemade interface.”

Figure 1:  OM1 and OM2 were added in the figure.

Page 6, Line 236-238: “2.2.2.3 Remove the alignment plate P0 and use the M2 hex key to adjust the OM2 so that the center of the Stokes beam colocalizes with the reference mark.”

Page 4, Line 139: “1.15. Connect the PMT to a current amplifier for signal amplification.”

Page 4, Line 141-146: “1.16.	Connect the photodiode to a lock-in amplifier.

1.17.	Connect the sync signal from the built-in EOM output to the reference input of the lock-in amplifier for SRS signal demodulation.

1.18.	Connect the PMT amplifier and lock-in amplifier outputs to the data acquisition module.”

Page 6, Line 257-258: “Set the wavelength of the pump laser to 796 nm and the power of the pump and Stokes beam to be 15 mW and 25 mW after the 25× objective lens, respectively.”

Comment 3:  The procedure to analyze the date is really incomprehensible (Lines 273-275)

Response:  	Thanks for pointing out the issue. The data analysis procedures in Lines 273-275 are performed by clicking buttons on the OPO control software. To make it more comprehensible, we added a figure (Figure 2) to show the interface of the delay manager control. These procedures will also be recorded by video afterwards.

Revision:  	Figure 2 was added

Figure 2 caption: “Figure 2. Interface of the delay manager on the OPO control software. The red arrow indicates the checking box for applying the calibrated delay data.”

Page 7, Line 306: “2.3.2.1.	Open the delay manager dialog (Figure 2) in the OPO control software.”

Comment 4:  Moreover, laser safety procedures are lacking, especially when a femtosecond Ti:sapphire laser is used at not visible wavelengths.

Response:  	Thanks for pointing out this important safety issue. We wear laser goggles for eye protection when handling the laser. This important issue has been added in our revised manuscript.

Revision:  	Page 2, line 83-85: “Safety training for laser handling is required to set up and operate the TPEF-SRS microscope. Always wear laser safety goggles with appropriate wavelength range when dealing with laser.”

Comment 5:  If the main objective is to promote the use of this imaging system, more space can be obtained to detail the setup by choosing only one of the two biological models to illustrate and validate the interest of this method.

Response:  	Agree. We removed the second biological model of adipose tissue imaging, and focused on the use of microscope system in the revised manuscript.

Comment 6:  I assume that all surgical preparation is performed under anesthesia conditions that are not mentioned here.

Response:  	Thanks for the comment. Actually we mentioned the anesthesia procedures in Line 335-336. We added description on how to confirm proper anesthetization in the revised manuscript.

Revision:  	Page 8, line 346-347: “Perform toe pinching to ensure deep anesthesia of the mouse. Supplement with half the original dose of the anesthetics as needed.”

Comment 7:  Again, laser safety procedures are lacking, especially when a femtosecond Ti:sapphire laser is used at not visible wavelengths.

Response:  	Thanks again for pointing out this issue. We have added laser safety procedures in the revised protocol.

Revision:  	Page 2, line 83-85: “Safety training for laser handling is required to set up and operate the TPEF-SRS microscope. Always wear laser safety goggles with appropriate wavelength range when dealing with laser.”

Responses to the comments from reviewer #3:

General: 	This review introduces the detailed procedures of in vivo imaging of biological tissues with combined two-photon fluorescence and stimulated Raman scattering microscopy.

Response:  	We appreciate the reviewer's positive opinion about our work. We have followed the reviewer’s comments to revise our manuscript accordingly.

Comment 1:  This paper covers both spinal cord imaging and white fat tissue imaging, which belong to very different areas. In my opinion, I suggest the paper focuses on one application.

Response:  	Thanks for the suggestion. We removed the second application of adipose tissue imaging in the revised manuscript.

Comment 2:  We usually use 2850cm-1 to image CH vibration in lipids instead of 2863.5cm-1.

Response:  	Thanks for the comment. We used 2863.5 cm-1 to image the myelin based on its SRS spectrum in the C-H region of which the SRS peak is observed at 2863.5 cm-1. The difference between the 2850 cm-1 and 2863.5 cm-1 in the spectral domain is <1 nm, which is very close. For C-H vibration in lipids, the corresponding Raman peak varies among different lipid samples, determined by their specific molecular structures, as shown in the Ref below. Therefore, this difference in wavenumber may result from the different sample used. We added an explanation on the wavenumber to avoid any confusion in the revised manuscript.

Ref:

Czamara, K., Majzner, K., Pacia, M.Z., Kochan, K., Kaczor, A., Baranska, M. Raman spectroscopy of lipids: a review: Raman spectroscopy of lipids. Journal of Raman Spectroscopy. 46 (1), 4–20, doi: 10.1002/jrs.4607 (2015).

Revision:  	Page 10, line 400-402: “Note: Spinal cord SRS imaging is performed at the wavenumber of 2863.5 cm-1, which corresponds to the Raman peak of myelin sheaths at the carbon-hydrogen region based on the measured SRS spectra.”

Comment 3:  The work used a series of different laser power for different imaging experiments. Please explain how to decide the laser power in more details.

Response:  	Thanks for pointing out this power issue. There is a tradeoff between the signal intensity and tissue damage when we determine the laser power for optical imaging. For two-photon fluorescence imaging, strong fluorescence signal can be acquired with low laser power (10mW) due to the bright fluorescence labelling. However, it is not the case for label-free SRS imaging, which usually requires higher laser power to acquire good image contrast, especially for this epi detection mode used in this protocol. After test, we found that the current power condition (60mW Pump beam, 75mW Stokes beam) allows high imaging contrast with myelin sheaths clearly visualized and at the same time, tissue damage was not observed. In fact, the laser power we used is comparable to the previous studies on in vivo SRS imaging of myelin (Ref below). We added discussion on the laser power in the revised manuscript.

Ref:

Hu, C.-R., Zhang, D., Slipchenko, M.N., Cheng, J.-X., Hu, B. Label-free real-time imaging of myelination in the Xenopus laevis tadpole by in vivo stimulated Raman scattering microscopy. Journal of Biomedical Optics. 19 (8), 086005, doi: 10.1117/1.JBO.19.8.086005 (2014).

Revision:  	Page 10, line 402-404: “The laser power is determined to ensure high-resolution and high-contrast spinal cord images in both fluorescence and SRS imaging modalities. Tissue damage is not observed under the current imaging conditions.”