# Submission ID #: 64497

# Scriptwriter Name: Nilesh Kolhe

Project Page Link: https://review.jove.com/account/file-uploader?src=19671633

Title: Atomic Force Microscopy Cantilever-Based Nanoindentation: Mechanical Property Measurements at the Nanoscale in Air and Fluid

Authors and Affiliations: Ashton E. Enrriques1, Sean Howard2, Raju Timsina3, Nawal K. Khadka3, Amber N. Hoover4, Allison E. Ray5, Ling Ding4, Chioma Onwumelu6, Stephan Nordeng6, Laxman Mainali3,7, Gunes Uzer2, Paul H. Davis1,8

1Micron School of Materials Science & Engineering, Boise State University

2Department of Mechanical & Biomedical Engineering, Boise State University

3Department of Physics, Boise State University

4Energy and Environmental Science and Technology, Idaho National Laboratory

5Science and Technology, Idaho National Laboratory

6Harold Hamm School of Geology & Geological Engineering, University of North Dakota

7Biomolecular Sciences Graduate Program, Boise State University

8Center for Advanced Energy Studies

# Corresponding Author:

Paul H. Davis				pauldavis2@boisestate.edu

# Co-authors:

pauldavis2@boisestate.edu

ashtonenrriques@u.boisestate.edu

seanhoward@boisestate.edu

rajutimsina@boisestate.edu

nawalkhadka@boisestate.edu

amber.hoover@inl.gov

allison.ray@inl.gov

ling.ding@inl.gov

chioma.onwumelu@und.edu

stephan.nordeng@und.edu

laxmanmainali@boisestate.edu

gunesuzer@boisestate.edu

# Author Questionnaire

1. Microscopy: Does your protocol demonstrate the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or similar?    No

2. Software: Does the part of your protocol being filmed demonstrate software usage? Yes

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://www.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen captured video files to your project page by your script return date or as soon thereafter as possible.

3. Interview statements: Please select one.

☒ 	Interviewees self-record interview statements.

4. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: 11/18/2022 to 11/23/2022

When you are ready to submit your video files, please contact our Content Engineer, Devon Halley.

Protocol Length

Number of Steps: 22

Number of Shots: 32

# Introduction

- Introductory Interview Statements

REQUIRED:

  - Paul H. Davis: Atomic force microscopy, or AFM, cantilever-based nanoindentation can be used to determine the nanoscale mechanical properties of materials ranging in modulus from kilopascals to gigapascals in both air and fluid.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Ashton E. Enrriques: AFM cantilever-based nanoindentation enables co-localized topographical imaging and in situ quantitative mechanical property measurements with nanoscale precision and resolution on a wide range of materials in relevant environments.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 3.5.1

  - Sean Howard: AFM cantilever-based nanoindentation can be used to differentiate between healthy and diseased structures, tissues, or cells that exhibit divergent mechanical properties.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Paul H. Davis: Accurately determining the tip-sample contact area and force applied during cantilever-based nanoindentation requires careful calibration of the AFM probe, which is challenging but essential for quantitative nanoscale mechanical property measurements.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 3.8.2 and 4.1.3

# Protocol

Authors: Acquire screen capture videos for all shots labeled SCREEN and upload them to your project page: https://review.jove.com/account/file-uploader?src=19671633

- Sample Preparation and Instrument Setup

  - To begin, select an appropriate atomic force microscopy or AFM (A-F-M) probe for nanoindentation of the intended sample based on the medium, expected modulus, sample topography, and relevant feature sizes [1]. Load the probe onto the probe holder [2] and attach the probe holder to the AFM scan head [3].

    - WIDE: Talent showing an array of samples, AFM probe boxes with differing specifications, and probe holders. Then, choose a sample and corresponding appropriate probe.

    - Talent loading the chosen probe onto the probe holder.

    - Talent attaching the probe holder to the AFM scan head (with AFM visible in the frame).

  - Select an appropriate nanoindentation mode in the AFM software that affords user control of individual ramps [1].

    - SCREEN: Opening AFM control software and selecting nanoindentation mode

  - Align the laser onto the back of the probe cantilever, opposite the location of the probe tip and into the position sensitive detector or PSD (P-S-D) [1]. Center the laser beam spot on the back of the cantilever by maximizing the Sum voltage [2].

    - Talent aligning laser onto the probe cantilever and into the PSD (with SUM and Horizontal/Vertical Deflection voltage readouts visible on AFM case screen).

    - SCREEN: Laser beam spot is centered on the back of the cantilever/Sum voltage being maximized in the Setup window (over the reflective sample or alignment station).

  - Next, center the reflected laser beam spot on the PSD by adjusting the Vertical and Horizontal deflection signals to be as close to zero as possible, thereby providing the maximum detectable deflection range for producing an output voltage proportional to the cantilever deflection [1].

    - SCREEN: Reflected laser beam spot is centered on the PSD by adjusting deflection signals close to zero in the Setup window.

- Probe Calibration

  - Calibrate the deflection sensitivity or DS (D-S) of the probe-AFM system. To do this, set up and perform DS calibration indents on sapphire to achieve approximately the same probe deflection as the planned sample indents since the measured displacement is a function of the tip deflection angle and becomes nonlinear for large deflections. [1-TXT].

    - SCREEN: DS calibration indents are being set up on the sapphire using the “Touch” Probe Calibration method in the Setup window. Use the mouse to highlight Ramp Size, Ramp Setpoint, and # of Curves, then Engage on the sample surface. TXT: Probe displacement should be in V or nm

  - Repeat the ramp five times. Determine the DS in nanometers per Volt, or the inverse optical lever sensitivity in Volts per nanometer, from the slope of the linear portion of the in-contact regime after the initial contact point in the resulting force-displacement or FD (F-D) curve [1].

    - SCREEN:  User clicks “Ramp,” and DS or inverse optical lever sensitivity is determined from the slope of the FD curve (adjust horizontal cursors with the mouse).

  - Use the average of the values for maximum accuracy. If the relative standard deviation exceeds 1%, remeasure the DS, as sometimes the first few F-D curves are nonideal due to the initial introduction of adhesive forces [1].

    - SCREEN: Average and RSD of the DS values from the 5 ramps are calculated and displayed (highlight with a mouse).

  - If the probe cantilever’s spring constant, k (single letter ‘K’), is not factory-calibrated, calibrate the spring constant [1].

    - SCREEN: Spring constant is calibrated via the thermal tune method (option in “Touch” Probe Calibration method in Setup window or via “Thermal Tune” under “Calibrate” menu).

  - If the probe does not come with a factory-calibrated tip radius measurement, measure the effective tip radius, R (single letter ‘R’). If employing the blind tip reconstruction or BTR (B-T-R) method, image the tip characterization or “roughness” sample using a slow scan rate and high feedback gains to help optimize tracking of the very sharp features. Choose an image size and pixel density based on the expected tip radius [1-TXT].

    - SCREEN: Roughness of the sample is imaged using a slow scan rate and high feedback gains with the appropriate image size and pixel density selected. TXT: See text for pixel size details

  - Next, use AFM image analysis software to model the probe tip and estimate its end radius and effective tip diameter at the expected sample indentation depth [1].

    - SCREEN: Probe tip end radius and effective tip diameter at the expected sample indentation depth are determined in image analysis software (Tip Qualification analysis tool in NanoScope Analysis).

  - Upon completing the probe calibration, enter the DS, k, and R values in the software [1].

    - SCREEN: DS, k, and R values are entered in the software’s “Check Parameters” workflow window.

  - Finally, enter an estimate of the sample’s Poisson’s ratio to convert the measured reduced modulus to the actual sample modulus [1]. If employing a conical or conispherical (coni-spherical) contact mechanics model based on the tip shape and indentation depth, it is also necessary to enter the tip half-angle [2].

    - SCREEN: Estimate of the sample’s Poisson ratio being entered in the software.

    - SCREEN: Tip half angle is entered in the software for use in conical or conispherical contact mechanics models.

- Collecting Force-Displacement (F-D) Data

  - Navigate the sample under the AFM head [1] and engage on the desired region of interest [2]. Monitor the vertical deflection signal or perform a small initial ramp to verify that the tip and sample are in contact [3].

    - SCREEN: Sample surface being focused and moved under the AFM head to the region of interest (ROI) in the Navigate software window.

    - SCREEN: Check Parameters and Engage on the surface.

    - SCREEN: Vertical deflection signal is shown in the Engage workflow, then a small (50-200 nm) initial ramp is performed after switching to the Ramp workflow.

  - Adjust the AFM head position slightly upward and ramp again. Repeat until the tip and sample are just out of contact, as evidenced by a nearly flat ramp and minimal vertical deflection of the cantilever [1].

    - SCREEN: AFM head position being adjusted (enable “RealTime Status” icon in the toolbar) and ramps performed with a shot of a flat ramp and minimal vertical deflection of the cantilever.

  - Once no obvious tip-sample interaction is present, lower the AFM head by an amount corresponding to approximately 50% of the ramp size to ensure the probe tip will not crash into the sample while manually moving the AFM head. Ramp again, repeating until a good curve is observed [1].

    - SCREEN: Lowering the AFM head and acquisition of a good F-D curve.

  - Adjust the ramp parameters to optimize for the chosen instrument, probe, and sample. Select an appropriate ramp size depending on the sample and desired indentation depth [1-TXT]. Then, select an appropriate ramp rate. 1 Hertz is a good starting point for most samples [2]. Set the number of samples per ramp to achieve the desired resolution of the measurement [3].

    - SCREEN: Appropriate ramp size is selected. Ramp size: Total Z-piezo movement through one ramp cycle

    - SCREEN: Appropriate ramp rate is selected.

    - SCREEN: Number of samples/ramp being set.

  - Set the X (single letter ‘X’)-rotate to reduce the shear forces on the sample and tip by simultaneously moving the probe slightly in the X-direction; parallel to the cantilever, while indenting in the Z-direction; perpendicular to the cantilever. Use a value for the X-rotate equal to the offset angle of the probe holder relative to the surface normal [1].

    - SCREEN: X-rotate is set to 12° to move the probe slightly in the X-direction while indenting in the Z-direction.

  - Next, choose whether to employ a triggered or untriggered ramp; if a triggered ramp is chosen, set the trigger threshold to result in the desired indentation into the sample [1].

    - SCREEN: Select the triggered ramp, and the trigger threshold is set.

  - Perform an indent in the chosen location.

    - SCREEN: Indent being performed on the sample.

- F-D Curve Analysis

  - Select and load the data to be analyzed [1]. Input calibrated values for the spring constant, DS or inverse optical lever sensitivity, and probe tip radius, along with an estimate for the Poisson’s ratio of the sample [2].

    - SCREEN: Talent loading the data to be analyzed in the software with monitor/screen visible in the frame.

    - SCREEN: Values for the spring constant, DS, and probe tip radius, along with an estimate for Poisson’s ratio of the sample being entered.

  - Choose a nanoindentation contact mechanics model appropriate for the tip and sample [1].

    - SCREEN: Nanoindentation contact mechanics model appropriate for the tip and sample is selected.

  - Run the fitting algorithm. Check for proper fitting of the F-D curves; a low residual error corresponding to an average R2 (R-squared) near unity is typically indicative of a good fit to the chosen model. [1] If desired, spot-check individual curves to visually inspect the curve, model fit, and calculated contact points [2].

    - SCREEN: Fitting algorithm for the F-D curve is run, and aggregate results, including Descriptive Statistics and possibly histograms or box plots, are viewed. Mouse hovers over R2 values for individual fits and aggregate results/statistics.

    - SCREEN: A representative curve is visually inspected using the Find Graphs feature (right-click on fit data), emphasizing model fit and calculated contact point.

Authors: Acquire screen capture videos for all shots labeled SCREEN and upload them to your project page: https://review.jove.com/account/file-uploader?src=19671633

# Results

- Results: Recent Advances in Atomic Force Microscopy for Surface Property Studies

  - Near-ideal F-D curves [1] obtained in air on a resin-embedded loblolly pine sample [2] and in phosphate buffered saline solution on a mesenchymal stem cell nucleus are shown [3].

    - LAB MEDIA: Figure 7

    - LAB MEDIA: Figure 7 Video Editor: Please emphasize Figure 7A

    - LAB MEDIA: Figure 7 Video Editor: Please emphasize Figure 7A

  - With a silicon probe [1], the tip experienced significant wear relative to its initial pristine state [2] throughout imaging [3]. With each subsequent image, the tip becomes progressively more rounded [4].

    - LAB MEDIA: Figures 8A-D

    - LAB MEDIA: Figure 8A Video Editor: Please emphasize R-value, R = 11 nm

    - LAB MEDIA: Figures 8B-D Video Editor: Please emphasize the R-value of each figure

    - LAB MEDIA: Figures 8B-D Video Editor: Please emphasize ‘tips’ (top portions) from each figure

  - In contrast, with the diamond probe [1], the tip radius did not change within the BTR method's limits, highlighting the diamond's extreme wear resistance [2].

    - LAB MEDIA: Figures 8E and F

    - LAB MEDIA: Figures 8E and F Video Editor: Please emphasize the R-value of each figure

  - An AFM topography image covering multiple cells with a pseudo-3D (three-D) depiction [1] and corresponding modulus map of a loblolly pine sample are shown here [2].

    - LAB MEDIA: Figure 9A

    - LAB MEDIA: Figure 9B

  - AFM-derived modulus map [1] generated while obtaining an AFM image [3] shows the mineral inclusion in the centre of images is significantly harder than the surrounding organic matrix [3].

    - LAB MEDIA: Figure 10D

    - LAB MEDIA: Figure 10B

    - LAB MEDIA: Figure 10

  - The effect of probe tip radius and shape on the appearance of high aspect ratio features of a Bakken shale sample is shown [1].

    - LAB MEDIA: Figure 11

  - Cantilever-based nanoindentation [1] on mesenchymal stem cells [2] and isolated nuclei exhibited no statistical difference in the elastic modulus [3].

    - LAB MEDIA: Figure 12

    - LAB MEDIA: Figure 12B Video Editor: Please emphasize the ‘Cell’ bar

    - LAB MEDIA: Figure 12 B Video Editor: Please emphasize the ‘Nuclei’ bar

  - The morphology and mechanical properties of cholesterol-containing lipid bilayers studied by AFM are demonstrated here [1].

    - LAB MEDIA: Figure 13

# Conclusion

- Conclusion Interview Statements

  - Ashton E. Enrriques: Proper probe calibration and selection of appropriate ramp parameters and contact mechanics models are essential to accurate nanomechanical measurements.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 3.3.1 and 4.4

  - Nawal K. Khadka: In addition to enabling elastic modulus measurements, AFM cantilever-based nanoindentation can be used to probe the rupture strength or breakthrough force of phospholipid bilayers under physiologically relevant conditions.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: 3.8.1

  - Sean Howard: AFM cantilever-based nanoindentation has enabled investigation of the effects of structural knockouts, pharmaceutical treatments, and low intensity vibrations to simulate exercise on the mechanical properties of mesenchymal stem cell nuclei.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B roll: LAB MEDIA: Figure 7 and Figure 12B