|  | Absorbance (OD) | Interpolated concentration (pg/ml) | Percent yield |
| --- | --- | --- | --- |
| Un-spiked control | -0.2087 | Does not fall on curve |  |
| Spiked at 200 pg/ml | 2.371 | 264.2±104 | 132.1% |
| Spiked at 100 pg/ml | 1.134 | 167.5±31.2 | 167.5% |
| Spiked at 50 pg/ml | 0.3218 | 80.55±4.54 | 161.1% |

Table 3. Results of an unsuccessful spike and recovery experiment of 3 control participants. Percent yield for spiked samples all fell well above the upper limit of 120% yield, indicating that the assay may be detecting non-specific binding. After obtaining these results, we added an additional step in the protocol involving blocking the plate prior to the assay with blocking buffer.