Submission ID #: 65411

Scriptwriter Name: Shalini K. Vasan

Project Page Link: https://review.jove.com/account/file-uploader?src=19950403

Title: Design and Building of a Customizable, Single-Objective, Light-Sheet Fluorescence Microscope for the Visualization of Cytoskeleton Networks

Short Title:

The Short Title is correct. (Character limit with spaces: 80)

Authors and Affiliations:

Nathan Felcher1, Daisy Achiriloaie1, Brian Lee2, Ryan McGorty3, Janet Sheung1,2,4

1W.M. Keck Science Department, Scripps College

2W.M. Keck Science Department, Claremont McKenna College

3Department of Physics and Biophysics, University of San Diego

4W.M. Keck Science Department, Pitzer College

All author names and affiliations are correct (city/state/country information not included in video title page).

Corresponding Authors:

Janet Sheung 				jsheung@kecksci.claremont.edu

Email Addresses for All Authors:

nfelcher@scrippscollege.edu

dachiril0224@scrippscollege.edu

blee24@cmc.edu

rmcgorty@sandiego.edu

jsheung@kecksci.claremont.edu

Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

N/A

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://www.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen captured video files to your project page as soon as possible.

3. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: 4/28/23-5/2/23

DO NOT use this draft script for filming. Please wait until your script is finalized to begin the filming process.

When you are ready to submit your video files, please contact our Content Engineer, Devon Halley.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  29

Number of Shots:  55

Interviews

- Video 1: Author Interviews

Answers to these questions will become interview statements that you will deliver on camera.

- Answer at least 4 of the questions below. Up to 7 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud. The questions will not appear in the video.

- Limit the length of each statement to 50 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

What is the scope of your research? What questions are you trying to answer?

- Janet Sheung: Our work aims to advance knowledge on the relationship between transport dynamics, local structure, and mechanical properties in far-from-equilibrium materials -- using dense and three-dimensional in vitro cytoskeleton composites as our experimental system.

What are the current experimental challenges?

- Daisy Achiriloaie: Visualizing far-from-equilibrium materials such as cytoskeleton networks requires an ability to image dense, three-dimensional samples with comparable image quality throughout the volume of the sample while causing minimal damage via photobleachinga technically challenging task that fluorescence lightsheet micrsocopy is best suited for.

What advantage does your protocol offer compared to other techniques?

- Nathan Felcher: Fluorescence Light-sheet microscope systems provide excellent optical sectioning capabilities which allows for imaging of these thick, three-dimensional samples over long timescales with minimal photobleaching. In particular, the single objective light-sheet system we will present is compatible with traditional slide-mounted samples, making it a very versatile tool.

What research gap are you addressing with your protocol?

- Nathan Felcher: The overwhelming majority of single objective light-sheet guides are written for users with extensive optics experience. To make the complex single-objective light-sheet style setup more accessible, this detailed guide provides steps that a user with an entry level optics course depth of knowledge can follow.

Protocol Videos

Each video will include a section of your protocol and accompanying results, if applicable. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration. Professional voiceover artists will narrate the video.

- Red and italics are pronunciation guides (how the word will be spoken).

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product from that step in advance.

Protocol:

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that you/your videographer will capture.

Representative Results:

- The three-digit numbers (e.g., 2.3.1., 2.3.2.) are the figures/tables from your manuscript. These will not be recorded by the videographer.

- Please review the result section to make sure it logically follows the video.

- Please note that the video cannot include voiceover without an accompanying visual.

- Video 2: Imaging Set-Up and Volumetric Scans for Visualization Cytoskeleton Networks

Demonstrator:

Protocol

- To begin, ensure that the microscope layout is prepared on the surface of the optical table with all the distances accurately measured [1]. Then, mount the excitation laser on the table [2]. Set two irises to the intended height of the laser and use these irises to ensure that the beam is level and centered [3].

- Shot of the microscope layout on the surface of the optical table.

- Talent positioning the excitation laser.

- Iris height being adjusted.

- Aligning the laser beam to the irises

- Position the translation stage (TS1) under the location of Mirror 1 or M1 (M-one) and direct the laser along the path drawn on the table [1-TXT].

- Translation stage being placed under (NOTE FOR SCRIPTWRITER: There is no laser adjustment until after M1 has been mounted). TXT: Mirrors: M1,  (only M1 appears in this step)

- Use the pair of irises set to the exact height to define the desired exit beam path and guide the placement and alignment of each reflective element [1].

- Talent placing two irises to define the desired exit beam path (NOTE FOR SCRIPTWRITER: These irises define where M1 should direct the laser beam).

- Next, position M1 at the top of the translation stage or TS1 (T-S-one) [1]. Then, mount and align the dichroic over the table [2]. Similarly, mount the galvo and align it with the irises [3].

- Talent placing M1 on top of TS1.

- Dichroic being aligned on the table.

- Talent pointing at the aligned galvo and irises.

- Once the alignments are made, position M2 (M-two), then clamp M3 (M-three) to the table [1]. Adjust the height and position until the beam is roughly centered on both frosted glass alignment disks [2].

- M2 and M3 being positioned. (NOTE FOR SCRIPTWRITER: We believe that these should be two separate steps)

- Shot of the beam centered over the frost glass alignment disks.

- Adding supports to M3.

- Next, start mounting Lens 1 on the table. Adjust the tilt and lateral position until the beam is centered on the frosted glass plates above M3 [1].

- Talent positioning Lens 1 on the table with the beam visible in the frame.

- Position Lens two and check the collimation by using a mirror to bounce the beam onto a faraway surface. Use an index card or target to trace the beam and ensure that the beam does not change in size. [1]. Then, adjust the XY (X-Y) position of the pinhole with the XY mount and the axial distance with the 1D (one-dimensional) stage to maximize the transmission [2]. (NOTE FOR SCRIPTWRITER: The pinhole is mounted in an XY adjustable mount, which is then mounted on top of a small 1D translation which translates in the axial, or Z, direction).

- Talent testing the laser and aligning the path.

- Talent adjusting the pinhole with XY mount and 1D Z stage.

- Adjust Lens 4 axially to focus the excitation beam on the surface of the galvo and arrange Lens 3 onto the table, followed by SL1 (S-L-one) [1-TXT]. Adjust the axial distance of SL1 to form a collimated telescope with Lens 4, then position TL1 (T-L-one) parallel to SL1 [2-TXT].

- Shot of the galvo surface with the focused excitation beam. TXT: SL: Scan Lens

- Using shear plate to collimate L4 and SL1 (NOTE FOR SCRIPTWRITER: We changed collimation methods in video step 2.7.1 and no longer use the shear plate there because the beam is too small to read reliably. For a demonstration of using the shear plate to collimate two lenses, this becomes the next best step to do so).

- Talent pointing at aligned SL1, L4 and TL1 respectively. TXT: TL: Tube Lens

- Adjust the height of Objective 1 on the cage system until the beam forms an Airy disk on the ceiling [1], and then continue adjusting until the size of the disk is minimized [2].

- Talent adjusting the height of O1.

- Shot of the changing airy disk size.

- Place the square mirror on the sample stage of Objective 1 [1], and adjust the mirror axially until the size of the beam profile is minimized after the dichroic [2].

- Square mirror being placed on the sample stage.

- Mirror being adjusted with the beam profile visible in the frame.

Aligning the Emission Path and Testing System Quality

- Mount the alignment laser by sliding the cage rods into the empty holes of the two cage plates [1]. Use one kinematic mirror mount and one drop-down mirror to align the paths of the alignment and excitation beams [2].

- Cage rod being inserted into the plate holes.

- Talent adjusting the beam paths.

- Position the square mirror axially on the sample stage of objective 1 to minimize the beam profile after the dichroic [1].

- Square mirror being positioned on the sample stage of O1.

- Insert SL2 and TL2 in the emission path at the respective distances [1-TXT]. Adjust the XY knobs and tilt of Objective 2 so that the red alignment beam passes through the iris and frosted glass disk [2].

- Talent pointing at pre-inserted SL2 and TL2. TXT: Distance measurements: SL2: Measured from the galvo;  TL2: Measured from SL2

- Shot of the red alignment beam passing through the iris and frosted glass disk.

- Adjust the translation stage 2 or TS2 (T-S-two) until the beam forms a small Airy disk on the surface [1], and then continue adjusting TS2 to minimize the size of the Airy disk [2].

- Talent adjusting TS2.

- Shot of the airy disk size getting smaller.

- To optimize the galvo for tilt-invariant scanning, press the FSK (F-S-K) button on the waveform generator to select a triangle wave signal for the galvo and set it to a low frequency such as 1 Hertz [1]. Observe the alignment beam on the same faraway surface or wall [2].

- SCREEN: FSK button clicked, triangle waveform is selected, and frequency set to 1 Hz. (NOTE FOR SCRIPTWRITER: This is a hardware interface with physical buttons, not a screen).

- Shot of the alignment beam on the surface.

Authors: Please check if all the shots labeled as SCREEN will be demonstrated using screen captures. For all the screen capture shots, kindly create the screen capture videos, a screenshot summary, and upload the files to your project page: https://review.jove.com/account/file-uploader?src=19950403.

- Mount Objective 3 about 4 to 5 millimeters in front of Objective 2 at a 0-degree angle and adjust the height to match [1]. Position a frosted glass alignment disk in the shared focal plane between SL2 and TL2, as measured with a ruler [2].

- Talent placing O3 in front of O2 and adjusting the height.

- Frosted glass alignment disk being positioned between SL2 and TL2.

- Look through the back of Objective 3 and adjust the height and axial position to find the emission light [1], then adjust it axially until the emission light fills the back aperture [2-TXT]. Mount a frosted glass disk in the rough position of the camera sensor and align the center of the disk to the emission light exiting O3.

- Shot of the emission light filling the back aperture. TXT: Fine-adjust O3 to center the exit path of the emission light

- Height of the frosted glass disk being adjusted.

- Position TL3 behind objective 3 and adjust the tilt to align the outgoing light with the frosted glass disk [1]. Place the camera at the measured distance from the tube lens [2] and adjust objective 3 axially from the cage translation stage until the hole is in focus on camera [3-TXT].

- TL3 being adjusted behind O3.

- Talent placing the camera at the measured distance.

- SCREEN: Shot of the focus being adjusted.

(NOTE FOR SCRIPTWRITER: Imaging the frosted glass disk and imaging the PSFs of the beads are two different steps. Please add a step after 2.18 in which users mount the 3D bead sample for imaging, then fine adjust the position of O3 until PSFs are round and bright.)

- Re-adjust Objective 3 at a 30-degree angle from the optical axis of Objective 2 using the lines on the table as a guide [1].

- Talent pointing at the pre-adjusted O3 and the guide-lines marked on the table.

- Remount the positive grid test target at the same axial height, and illuminate the grid with the brightfield light [1-TXT]. Sweep the in-focus portion of the Field of View across the screen horizontally while the grid squares maintain a uniform size [2].

- Talent mounting the test target and illuminating the grid. TXT: Only one vertical section should be in focus

- SCREEN: In focus portion of the Field of view being adjusted.

- To align the oblique light sheet, position cylindrical lenses or CL so that the beam is focused into a horizontal sheet profile at the focal plane of CL3 [1-TXT].

- Talent aligning the oblique light sheet with the beam visible in the frame. TXT: Mount CL1; Rotate CL2 and CL3 to the same orientation as CL1

- Insert and position a slit in the vertical orientation at the focal plane between CL3 and Lens 3 [1]. At the camera sensor, confirm that the 0-degree light sheet appears thin and vertical [2].

- Slit being inserted and positioned between CL3 and L3.

- SCREEN: The 0-degree sheet appearing as a thin and vertical projection.

- Using the motorized translation stage control, translate M1 toward the cylindrical lenses to angle the light sheet [1-TXT].

- M1 being moved toward the cylindrical lenses. TXT: Set the angle to ~60°relative to the optical axis of O1

- Insert the pre-prepared rhodamine-labeled microtubule test sample onto the sample stage [1] and adjust it axially so that the dye is illuminated by the light sheet at five different depths between the center of the FoV and the right side of the screen [2]. Save each image [3].

- Talent mounting the sample on the stage.

- SCREEN: Stage depth being adjusted between FoV and the right side of the screen.

- SCREEN: Image being saved.

- Open the images in Fiji [1]. For each image, select the Line tool, and draw a horizontal line from the center of the FoV to the center of the light sheet [2]. To measure the displacement, go to Analyze | (then) Measure to see the length of the lines [3-TXT].

- SCREEN: Images being opened in Fiji.

- SCREEN: Line Tool selected and horizontal line being drawn from FoV to the light sheet.

- SCREEN: Analyze | Measure is clicked, and displacement is measured.

(NOTE FOR SCRIPTWRITER: The optimization of signal to noise ratio (by adjusting O3 stages) is a separate step. This requires the 3D bead sample to be mounted on the sample stage, which can only be done after the light-sheet angle has been set to 60 degrees. Please add a separate step for mounting the 3D bead sample and then optimizing the signal to noise ratio).

Acquiring Volumetric Scans

- After calibrating the instrument, mount the pre-prepared 3D (three-dimensional)  bead sample [1], and click the FSK button on the function generator to set a triangle wave [2].

- Talent mounting the samples on the stage.

- SCREEN: FSK button being clicked and triangle wave selected. (NOTE FOR SCRIPTWRITER: This is a hardware interface with physical buttons, not a screen).

- To find the sample, use the function generator to set the parameters starting from 400 mV (millivolts) peak-to-peak amplitude, 0 (zero) offset, 200 mHz (millihertz)  frequency and 100 ms (milliseconds)  exposure time [1]. Scroll in Z (Zee) manually until the sample plane is reached and optimize the z (Zee) setting [2].

- SCREEN: Function generator parameters being set. Authors: Please show all the parameters.

- SCREEN: Z setting being optimized.

- In the Micromanager program, select an exposure time, and open the Multi-Dimensional Acquisition window [1]. Use the Count box to choose the number of frames, which will set the total acquisition time, and on the function generator, set the frequency to create a full-volume scan [2].

- SCREEN: Micromanager program clicked, exposure time selected and Multi-Dimensional Acquisition window is opened.

- SCREEN: Number of frames and total acquisition time being set in the Count Box. Frequency being entered in the function generator tab.

- Once the parameters are set, record a time-lapse for one full scan of the volume [1].

- SCREEN: Time-lapse recording initiated.

Representative Results (Word limit: 100)

- The volumetric scans of the reconstituted microtubule network showed that the 3D structures grew dense toward the center, resulting in bright regions of fluorescence [1].

- LAB MEDIA: Figure 15.

- In imaging planes near the coverslip, confocal microscopy resolved single filaments around the periphery of the aster, with additional background toward the center due to out-of-focus fluorescence signals from above [1].

- LAB MEDIA: Figure 15. Video editor: Emphasize A.

- However, moving a few microns in z quickly reduced the quality of the images due to the out-of-focus dense sections of the aster [1].

- LAB MEDIA: Figure 15. Video editor: Emphasize the arrow and +Z above figure 15A.

- The single-plane illumination of the light sheet eliminated the out-of-focus signals, allowing comparable image quality between the planes [1].

- LAB MEDIA: Figure 15. Video editor: Emphasize B.