Submission ID #: 65562

Scriptwriter Name: Sritama Bose

Project Page Link: https://review.jove.com/account/file-uploader?src=19996458

# Title: Fentanyl Analog Screening using LC-TIMS-TOF MS/MS

Landing Page Title (not for video use): Harvesting the Orthogonality of LC and TIMS-TOF MS/MS to Screen Fentanyl Analogs

Authors and Affiliations:

Andrew R. Forero1, Lilian Valadares Tose1, Matthew Willetts2, Melvin A. Park2, Elisa N. Shoff3, Francisco Fernandez-Lima1,4

1Department of Chemistry and Biochemistry, Florida International University

2Bruker Daltonics Inc.

3Miami-Dade County Medical Examiner Department

4Biomolecular Sciences Institute, Florida International University

# Corresponding Authors:

Francisco Fernandez-Lima	(fernandf@fiu.edu)

# Email Addresses for All Authors:

Andrew R. Forero		(afore014@fiu.edu)

Lilian Valadares Tose		(lvaladar@fiu.edu)

Matthew Willetts		(matt.willetts@bruker.com)

Melvin A. Park			(mel.park@bruker.com)

Elisa N. Shoff			(eshoff@miamidade.gov)

Francisco Fernandez-Lima	(fernandf@fiu.edu)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  NO

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  YES

Authors: We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://www.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

Please upload all screen captured video files to your project page as soon as possible. https://review.jove.com/account/file-uploader?src=19996458

3. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: MM/DD/YYYY

When you are ready to submit your video files, please contact our Content Assistant, Utkarsh Khare

Current Protocol Length

Number of Steps:  24

Number of Shots:  50

# Interviews

- Video 1: Author Spotlight: An Efficient Methodology to Confidently Differentiate and Characterize Fentanyl Analogs

What is the scope of your research? What questions are you trying to answer?

  - Andrew Forero: The scope of our research is to develop a method to discriminate and characterize isomeric analogs. This research aims to investigate if a method can be developed to confidently differentiate a large quantity, especially structural isomers.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

What research gap are you addressing with your protocol?

  - Andrew Forero: This protocol incorporates four modes of characterization that have previously not been coupled. Using four orthogonal methods of identification for the analogs should improve the ability to differentiate isomers that are difficult to differentiate with current standard procedures.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

How will your findings advance research in your field?

  - Lilian Valadares Tose: Fentanyl is a big societal concern in the opioid epidemic. New analogs created in illicit labs are being found every day, and having a methodology that quickly and confidently helps determine which analogs are in a sample could improve the efficiency of drug law enforcement.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

# Protocol Videos

Authors: Please do not miss capturing the opening WIDE shots of each video.

Video 2: HPLC Method Development for Screening of Fentanyl Analogs

Demonstrator: Lilian Tose

Protocol

Begin by preparing the mobile phases for high-performance liquid chromatography or HPLC (H-P-L-C) [1]. For Mobile phase A or MPA (M-P-A), use 5 millimolar ammonium formate in water with 0.05% formic acid [2]. For mobile phase B or MPB (M-P-B), use 0.05% formic acid in a 1:1 (one-is-to-one) methanol-acetonitrile mixture [3].

WIDE: An establishing shot of talent adding water to the MPA bottle to fill it up to the desired level.

Talent placing the labeled bottle of MPA in the mobile phase slot of the HPLC.

Talent placing the labeled bottle of MPA in the mobile phase slot of the HPLC.

To begin the HPLC method development, open the liquid chromatography or LC software and create a new protocol [1]. Then, click the Edit Method button to edit the binary gradient [2]. Under the binary gradient tab, set the time stamps and concentrations of MPA and MPB [3].

SCREEN: The LC software is opened, and a new HPLC protocol is created. Authors: Considering the length limit, it would be difficult to demonstrate the whole protocol creation. What you can do is record the screen showing the already created protocol and point the cursor to the name of the new protocol that was created.

SCREEN: The Edit Method button is clicked to begin creating a new protocol.

SCREEN: In the binary gradient tab, the values of the time stamps and concentrations of MPA and MPB are being entered.

Set the stop time to 18 min (minutes), the flow rate to 0.400 mL/min (milliliters per minute), and the minimum and maximum pressure limits to 0 and 6000 psi (p-s-i), respectively [1].

SCREEN: The stop time is assigned as 18 min, the flow rate is assigned as 0.400 ml/min, the minimum pressure limit is set to 0 psi, and the maximum pressure limit is set to 6000 psi.

To initialize the HPLC, insert the LC column and column guards in the HPLC column section [1]. Carefully observe the minimum and maximum column pressures [2].

Talent attaching the LC column and column guards in the HPLC column oven.

SCREEN: The cursor is pointed at the minimum and maximum pressure displayed on the software.

Video 3: TIMS-TOF MS/MS Method Development and Calibration of Mobility and Mass for Fentanyl Analog Screening

Demonstrator: Andrew Forero

Protocol

To begin TIMS (tims) TOF (tawf) MS/MS (M-S-M-S) method development, open the timsControl (tims-control) application [1-TXT]. Under MS settings, set the scan beginning and ending values to 50 m/z (m by z) and 1800 m/z, respectively [2]. Select the ion polarity mode as positive and the scan mode as parallel accumulation-serial fragmentation [3].

WIDE: Talent sitting in front of a computer and opening the timsControl application. TXT: TIMS: Trapped ion mobility spectrometry; TOF: Time of flight; MS: Mass Spectrometry

SCREEN: The MS settings in the timsControl software is opened, and the scan value for the beginning is set to 50 m/z, and that for the ending is set to 1800 m/z.

SCREEN: The ion polarity mode is set to positive, and the scan mode is assigned as parallel accumulation-serial fragmentation.

Next, move to the tabs under Source, and set the end plate offset to 500 V (volts), capillary to 4500 V, nebulizer to 3.0 bar, dry gas to 10.0 L/min (liters per minute), and dry temp to 200 °C (degrees Celsius) [1]. In the syringe pump settings, assign the syringe as Hamilton 1 mL (milliliter), enable the option active, and set the flow rate to 80.0 µL/h (microliters per hour) [2].

SCREEN: In the tabs under Source, the end plate offset is set to 500 V, the capillary is set to 4500 V, the nebulizer value is set to 3.0 bar, the rate of dry gas is set to 10.0 L/min, and the dry temperature is set to 200 °C.

SCREEN: The syringe pump setting tab is opened, and the syringe is entered as Hamilton 1 mL, the active option is enabled, and the flow rate is set to 80 µL/h.

In the MS/MS tab under precursor ions, set the number of parallel accumulation-serial fragmentation ramps to 8, minimum charge to 1, and maximum charge to 5 [1].

SCREEN: Under the precursor ion option, the number of parallel accumulation-serial fragmentation ramps is set to 8, the minimum charge is entered as 1, and the maximum charge is entered as 5.

In scheduling settings, enable the precursor repetitions, and under active exclusion, check the box to enable and set the release to 0.40 min (minutes) after [1]. Do not adjust the collision energy settings and isolation width settings. Click to enable the TIMS Stepping Box [2].

SCREEN: The precursor repetitions option is enabled in scheduling settings, and under active exclusion, the box is checked, and the release is set to 0.40 min.

SCREEN: The cursor points to the collision energy settings and the isolation width settings. Then, the TIMS Stepping Box is clicked and enabled.

To perform calibration for both m/z and mobility domains, under the m/z calibration settings, select one of the preloaded tuning mix profiles in the reference list box [1].

SCREEN: In the m/z calibration settings, a preloaded tuning mix profile in the reference list box is selected.

To calibrate, load the syringe for the TOF with a tuning mix solution [1]. In the calibration mode on the right side of the window, select between linear, quadratic, and enhanced quadratic to achieve the best score and click calibrate to get as close to 100% as possible [2].

Talent drawing tuning mix solution from a labeled vial to the TOFsyringe.

SCREEN: The calibration types being set under the calibration mode linear, quadratic, and enhanced quadratic are being selected until the highest score close to 100% is achieved.

Next, calibrate for mobility following the steps demonstrated previously for calibrating m/z [1]. Once calibrated, save the MS method by selecting the Method tab on the top bar. In the drop-down menu, select Save As to generate a new MS method file [2].

SCREEN: In the mobility calibration settings, a preloaded tuning mix profile in the reference list box is selected.

SCREEN: The Method tab is selected, and from the dropdown menu, Save As is clicked. Then, a new file is generated.

Video 4: Data Independent Acquisition (dia) Parallel Accumulation-Serial Fragmentation Method for Screening of Fentanyl Analogs

Demonstrator: Andrew Forero

Protocol

Once a TIMS-TOF MS/MS data set has been collected in the data-dependent acquisition or dda (d-d-a) parallel accumulation-serial fragmentation mode [1-TXT], run the data-independent acquisition or dia (d-i-a) on the ion mobility software application. To do so, first, load the previously saved dda method [2].

WIDE: Talent sitting in front of a computer and opening the ion mobility software. TXT: TIMS: Trapped ion mobility spectrometry; TOF: Time of flight; MS: Mass Spectrometry

SCREEN: The ion mobility software is open, and the previously saved dda method is being loaded in it.

Change the mass spectrometry or MS setting from parallel accumulation-serial fragmentation to dia- parallel accumulation-serial fragmentation. Leave all other settings the same [1]. Under the MS/MS tab the dia-PASEF windows tab should appear, select edit windows. Using the open analysis button located at the top of the dia popup window, load the previously saved dda data set [2].

SCREEN: The setting being changed from parallel accumulation-serial fragmentation to dia- parallel accumulation-serial fragmentation.

SCREEN: The open analysis button at the top of the dia popup window is clicked, and the previously saved dda dataset is loaded.

When a heat map appears at the bottom left, displaying a polygon of windows running diagonally across the graph, click and drag to resize the polygon so that it fits the data in the heat map [1].

SCREEN: A heat map appears at the bottom left, which shows a polygon of windows running diagonally across the graph. The polygon is resized by clicking and dragging it to fit the data in the heat map.

Set the number of mobility windows in the per mass width and the overlap [1]. Click Calculate Windows on the bottom right of the popup to see the windows displayed with the new settings [2]. Once suitable, click Apply dia-PASEF (Pasef) Windows to Method [3].

SCREEN: The number of mobility windows in the per mass width and the overlap being set.

SCREEN: Calculate Windows is clicked on the bottom right of the popup, and the windows with the new settings are displayed.

SCREEN: The Apply dia-PASEF Windows to Method is clicked.

Video 5: Processing HPLC Ion Mobility TOF Data of Fentanyl Analog Screening

Demonstrator: Andrew Forero

Protocol

To process the HPLC ion mobility TOF data of fentanyl analog screening, begin by checking the calibration in the data analysis software [1-TXT]. To do so, right-click on the file name under the analysis box and select Properties. When the window labeled file name_analysis properties appears, select Calibration Status [2].

WIDE: Talent sitting in front of a computer and opening the data analysis software. TXT: HPLC: High-performance liquid chromatography; TOF: Time of flight

SCREEN: The file name under the analysis box is clicked, and Properties is selected. Then Calibration Status is selected when a window labeled file name_analysis properties appears.

From the drop-down box, select Instrument Calibration for the Mass Spectrometer and confirm that the error is no greater than 1 µg/mL (micrograms per milliliter) [1]. Then, select Initial Mobility Calibration and confirm that the error is no greater than 1 µg/mL [2].

SCREEN: Instrument Calibration for the Mass Spectrometer is selected from the dropdown box, and then the cursor is moved to the error value to show it is not greater than 1 µg/mL.

SCREEN: Initial Mobility Calibration is selected and the cursor is moved to the error value to show it is not greater than 1 µg/mL.

To prepare a chromatogram in the data analysis software, right-click on Chromatogram under the file name and select Edit Chromatogram. In type, click the drop-down box and select Extracted Ion Chromatogram [1]. Under filter, select All MS, and for scan mode, select All [2].

SCREEN: The Chromatogram option in data analysis software is right-clicked, and Edit Chromatogram is selected under the file name. Then, the drop-down box is clicked and Extracted Ion Chromatogram is selected.

SCREEN: All MS is selected, followed by the scan mode being assigned as ALL.

If using masses to extract an ion, insert the theoretical m/z of the molecule of interest [1]. If using a formula to extract an ion, insert the formula for the molecule as well as the ion forms of interest for the chromatogram [2]. Set the polarity to positive mode and add the ion to the list[3].

SCREEN: The theoretical m/z value of the molecule of interest is inserted.

SCREEN: The polarity is set to positive and ion is added to the list.

SCREEN:. The formula for the molecule and the ion forms of interest for the chromatogram is inserted. Polarity is already in positive mode.

For the mass spectrum, make sure the chromatogram is selected and right-click at the baseline of one edge of the compound peak and drag to the other edge of the peak [1].

SCREEN: The baseline on one edge of the compound peak is selected by right-click, and the cursor is dragged to the other edge of the peak.

To generate a Mobilogram, right-click on the left tab titled Mobilogram and select Edit Mobilogram [1]. When the window titled edit Mobilogram traces appears, begin editing like what was done with the chromatogram [2].

SCREEN: The left tab titled Mobilogram is selected by right-click, and then the Edit Mobilogram option is selected.

SCREEN: The edit Mobilogram window appears on the screen.

In the retention time input, add the retention time range of the peak of interest [1]. Once the parameters are selected, click Add, followed by Ok at the top right [2].

SCREEN: The retention time range of the peak of interest is added under the retention time input tab.

SCREEN: After selecting the parameters, Add and OK are clicked sequentially.

To generate the compound spectra, at the bottom of the spectrum view window, sequentially select Profile MS and Fragment MS [1]. At Spectrum view, right-click and select Copy Compound Spectra [2].

SCREEN: Profile MS and Fragment MS are selected at the bottom of the spectrum window.

SCREEN: Copy Compound Spectra is selected after right-clicking on Spectrum view.

To process the data, manually integrate the chromatogram and mobility peaks by clicking on Find. Then, select Compounds manually chromatogram or mobilogram [1]. Left-click and drag to highlight the desired peak to yield important information on the molecule of interest [2].

SCREEN: Find is clicked, and then compounds manually chromatogram or mobilogram is selected.

SCREEN: The desired peak is highlighted by left-clicking and dragging until information on the molecules appears.

- Authors: Please record all shots labeled as SCREEN using screen recording software, and upload the video files to your project page as soon as possible. `

- Representative Results

Screening of the four isomers in a particular fentanyl analog screening kit [1] revealed that the methoxy-acetyl fentanyl and fentanyl carbamate separated well in the LC and mobility domains [2] from the para-methoxy acetyl fentanyl and B-hydroxyl fentanyl [3]. However, a clear separation of the latter two was achieved in their fragmentation patterns [4].

LAB MEDIA: Figure 3.

LAB MEDIA: Figure 3. Video Editor: In panel 3A, emphasize the top-most and the bottom-most chemical structures with their names. In both panels 3B and 3C, emphasize the top-most and the bottom-most graphs.

LAB MEDIA: Figure 3. Video Editor: In panel 3A, emphasize the second and third chemical structures from the top, along with their names. In both panels 3B and 3C, emphasize the second and third graphs.

LAB MEDIA: Figure 3. Video Editor: In panel 3A, emphasize the second and third chemical structures from the top, along with their names. In panel 3E, emphasize the second and third graphs.

Screening of three isomers in another kit [1] showed that α (alpha)-methyl thiofentanyl and (±)-trans-3-methyl thiofentanyl separated in the LC domain but not in the mobility domain [2]. However, (±)-trans-3-methyl thiofentanyl and (±)-cis-3-methyl thiofentanyl separated in the mobility domain and not in the LC domain [3].

LAB MEDIA: Figure 4 A,B, and C.

LAB MEDIA: Figure 4 A,B, and C. Video Editor: In panel 4A, emphasize the top two structures and their names. In both panels 4B and 4C, emphasize the top two graphs.

LAB MEDIA: Figure 4 A,B, and C. Video Editor: In panel 4A, emphasize the bottom two structures and their names. In both panels 4B and 4C, emphasize the bottom two graphs.

This data indicated the complementarity of the orthogonal LC and mobility domains in improving the identification of the isomeric species [1].

LAB MEDIA: Figure 4. Video Editor: Please emphasize 4D.