# Submission ID #: 67066

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20455178

# Title: Zebra II as A Novel System to Record Electrophysiological Signals in Zebrafish

Landing Page Title (not for video use): Zebra II: Advanced ECG System for Multichannel Zebrafish Cardiovascular Studies

☐   The Landing Page Title is correct. (Character limit with spaces: 80)

# Please keep the same title for the main and landing page title: Zebra II as a Novel System to Record Electrophysiological Signals in Zebrafish.

# Authors and Affiliations:

# Ramses Seferino Trigo Torres1, Mao-Hsiang Huang2, Mohamed Benomar2, Tai Le1, Tim Etchells3, Xiaolei Xu4, Michael P. H. Lau3, Hung Cao1,2

# 1Department of Biomedical Engineering, UC Irvine

# 2Department of Electrical Engineering and Computer Science, UC Irvine

# 3Sensoriis Inc.

# 4Department of Biochemistry and Molecular Biology, Mayo Clinic, Rochester

☒   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# Hung Cao					hung.cao@uci.edu

# Email Addresses for All Authors:

Ramses Seferino Trigo Torres			rtrigoto@uci.edu

Mao-Hsiang Huang				maohsiah@uci.edu

Mohamed Benomar				mbenomar@uci.edu

Tai Le						tail3@uci.edu

Tim Etchells					tetchells@sensoriis.com

Xiaolei Xu					Xu.Xiaolei@mayo.edu

# Michael P. H. Lau				michaelphlau@comcast.net

# Hung Cao					hung.cao@uci.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   No

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  22

Number of Shots:  37

# Interviews

- Video 1: Author Spotlight: Title (Filled by scriptwriter during script finalization)

Ethics Title Card

This research has been approved by the Institutional Animal Care and Use Committee (IACUC) at University of California, Irvine

Videographer: Obtain headshots for all authors.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 50 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Hung Cao: Our lab develops novel systems to provide electrophysiological assessments in zebrafish with better performance and longer duration. Our tools have been used to study cardiac genes and carry out drug screening using zebrafish.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What are the current experimental challenges?

  - Ramses Seferino Trigo Torres: Electrophysiological studies in Zebrafish are limited by short recording times, less than 5 minutes per fish, which extends study duration and restricts sample size and experiment length.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Hung Cao: Our technique enables prolonged up-to-1hour recording of multi-channel ECG (expandable to EEG) while others are limited to one channel and 5 minutes.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What new scientific questions have your results paved the way for?

  - Michael Lau: How to utilize zebrafish cardiac optical data correlating with electrocardiography to study cardiac pathology and developmental abnormalities, especially in the context of pre-clinical drug development.

What research questions will your laboratory focus on in the future?

  - Michael Lau: How to leverage our expertise in optical, electrical and chemical sensing integrating AI to utilize zebrafish as a cost-effective animal model for drug development in the area of cardiovascular medicine, neurology, neuropsychology and pain medicine.

Videographer: Obtain headshots for all authors.

# Protocol Videos

Each video will include a section of your protocol and accompanying results, if applicable. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration. Professional voiceover artists will narrate the video.

- Red and italics are pronunciation guides (how the word will be spoken).

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product from that step in advance.

Protocol:

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Representative Results:

- The three-digit numbers (e.g., 2.3.1., 2.3.2.) are the figures/tables from your manuscript. These will not be recorded by the videographer.

- Please review the result section to make sure it logically follows the video.

- Please note that the video cannot include voiceover without an accompanying visual.

- Video 2: Amiodarone Treatment and Zebra II Device Set Up to Monitor ECG of Zebrafish for Cardiovascular and Pharmacological Research

Demonstrator: Ramses Seferino Trigo Torres

Ethics Title Card

Procedures involving animal subjects have been approved by the Institutional Animal Care and Use Committee (IACUC) at University of California, Irvine

Protocol

  - To begin, equally divide 40 fish into four groups and house them in separate tanks [1].

    - WIDE: Talent pouring the fish with water into the housing tank. Videographer: Please take a still image of the talent performing this action. Make sure that it is at least a half-body shot with the talent's face visible and zoom out so we have room for cropping.

  - Connect the Zebra II (2) device to a power source [1] and initiate operation with the Red Button located on the front panel [2].

    - Talent connecting the device to a power source.

    - Close-up of the talent pressing the Red Button on the front panel.

  - Navigate through the device’s touchscreen to choose the chamber to be configured [1].

    - Close up of navigating the touchscreen to select the chamber.

  - Adjust the low-pass filter setting to 0.1 Hertz and the high-pass filter setting to 500 Hertz [1] and set the gain to 10,000 [2].

    - Low-pass filter being adjusted to 0.1 Hertz and high-pass filter to 500 Hertz.

    - Gain setting adjusted to 10,000.

  - Lift the front lid to access the four chambers [1]. Prior to starting the experiment, immerse each chamber’s sponge in Tricaine solution and squeeze out any excess liquid [2].

    - Talent lifting the front lid to access the chambers.

    - Close-up of the talent immersing the sponge in Tricaine solution and squeezing out excess liquid.

  - Rotate the four-electrode array to ensure the sponge is accessible before introducing the fish [1].

    - Close-up of the talent rotating the four-electrode array.

  - For the treatment, dissolve 258.13 milligrams of Amiodarone in 1 liter of deionized water [1]. Insert a magnetic stirrer into the container and stir until the solution is completely dissolved and clear [2-TXT].

    - Talent adding 258.13 milligrams of Amiodarone to 1 L water.

    - Close-up of the solution being stirred.
Talent storing the solution in a refrigerator at 4 degrees Celsius. TXT: Store the solution at 4°C when not in use

  - Dilute the Amiodarone stock solution to prepare 70, 100 and 200 micromolar working solutions in fish water [1]. Submerge each fish individually in the prepared Amiodarone solution for exactly 5 minutes [2].

    - Talent adding stock solution into 41.25 milliliters of fish water.

    - Talent submerging each fish individually in the Amiodarone solution.

- Video 3: ECG Data Collection in Zebrafish using the Zebra II System for Biomedical Research

Demonstrator: Ramses Seferino Trigo Torres

Ethics Title Card

Procedures involving animal subjects have been approved by the Institutional Animal Care and Use Committee (IACUC) at University of California, Irvine

Protocol

  - To begin, treat the fish with appropriate concentrations of Amiodarone [1].

    - WIDE: Talent lifting the fish out of the solution.

  - After anesthetizing the fish, place it onto the sponge, ensuring its ventral side faces up [1].

    - Talent placing the fish on the sponge with its ventral side up.

  - Position the pre-set electrode array in a rhombus shape over the fish’s chest area to capture signals from four different points simultaneously [1]. Make sure to adjust each electrode carefully so that every pin makes contact with the fish’s skin [2].

    - Talent positioning the electrode array in a rhombus shape over the fish’s chest.

    - Close-up of the talent adjusting the electrode to ensure the pins make contact with the fish’s skin.

  - Then, rotate each electrode, lowering the pin until it makes contact with the fish’s underside [1]. If the pin is too high, turn the corresponding thumbscrew clockwise to lower it until it touches the fish [2].

    - Close-up of the talent rotating each electrode, lowering the pin to touch the fish's underside.

    - Close-up of the talent turning the thumbscrew clockwise to lower the pin until it contacts the fish. TXT: Perform the adjustment for the electrodes in all chambers

  - Close the chambers’ lid once all electrode arrays are correctly positioned over the fish in the operational chambers [1].

    - Talent closing the chambers' lid after positioning the electrode arrays.

  - To mark each chamber for recording, select it on the touchscreen [1], then initiate the recording by clicking the Record button on-screen [2].

    - Talent selecting each chamber on the touchscreen for recording.

    - Talent clicking the Record button on-screen.

  - Click the Stop button after 2 minutes [1] and then click the Save button from the prompt window that pops up [2]. Save the information either on a disk or to the cloud, depending on specific research needs [3].

    - Talent clicking the Stop button after 2 minutes.

    - Talent clicking the Save button on the prompt window.

    - Display showing options to save the information on a disk or to the cloud.

- Video 4: Troubleshooting the Four-Electrode Array Positioning System in Zebra II Device for Enhanced ECG Analysis of Zebrafish

Demonstrator: Ramses Seferino Trigo Torres

Ethics Title Card

Procedures involving animal subjects have been approved by the Institutional Animal Care and Use Committee (IACUC) at University of California, Irvine

Protocol

  - To begin, initiate electrocardiogram or ECG acquisition of the Amiodarone-treated fish [1].

    - WIDE: Talent acquiring the ECG of the fish.

  - For troubleshooting, press the Stop button on the device interface to halt the recording process [1] and identify the channels with inconsistent signals [2].

    - Talent pressing the Stop button on the device interface.

    - Display showing identified channels with inconsistent signals.

  - Lift the front lid for access to the electrode chambers [1] and identify the chamber and specific electrodes causing the issue [2].

    - Talent lifting the front lid for access to the chambers.

    - Close-up of the talent pointing to the problematic chamber and electrode.

  - Turn the top thumbscrew clockwise to adjust the electrode’s proximity to the fish, lowering the electrode closer to the fish [1]. If merely moving closer does not solve the issue, rotate the electrode to a different spot on the fish until a satisfactory signal is achieved [2].

    - Close-up of the talent turning the top thumbscrew clockwise to lower the electrode.

    - Close-up of the talent rotating the electrode to a different spot on the fish.

Data Extraction

  - Once data has been saved, retrieve it for analysis or send it directly to a cloud server for further processing [1]. Using the Browse button, navigate to the saved data and select the data sets to use [2].

    - Display showing options to retrieve data for analysis or send it to a cloud server.

    - Talent clicking the Browse button on the screen and selecting data sets.

  - Select the Extract button from the options [1] and attach a USB drive to the back part of the device to extract the data [2].

    - Talent selecting the Extract button on the screen.

    - Close-up of the talent attaching a USB drive to the back of the device.

  - Follow the device prompts to save the selected data sets onto the external drive for subsequent retrieval and analysis [1].

    - Saving the data onto the external drive.

Representative Results

  - The control group displayed a consistent heart rate around 120 Beats Per Minute and a QTc (Q-T-C) interval of 330 milli seconds, showing normal ECG patterns [1].

    - LAB MEDIA: Figure 3. Video editor: Highlight the peaks in A and bar in C for “control”.

    - LAB MEDIA: Figure 3. Video editor: Highlight the peaks in B and bar in D for “control”.

  - Amiodarone treatment resulted in a dose-dependent increase in QTc interval, extending from 365 milliseconds at 70 micromolar to 546 milliseconds at 200 micromolar [1], and a corresponding decrease in heart rate from around 105 Beats Per Minute at 70 micromolar to 84 Beats Per Minute at 200 micromolar [2].

    - LAB MEDIA: Figure 3. Video editor: Highlight the peaks in B and bar in D for “70 µM, 100 µM and 200 µM”.

    - LAB MEDIA: Figure 3. Video editor: Highlight the peaks in A and bar in C for “70 µM, 100 µM and 200 µM”.