Submission ID #: 67137

Scriptwriter Name: Pallavi Sharma

Project page Link: https://review.jove.com/account/file-uploader?src=20476833

Title: Radiofrequency Echographic Multi-Spectrometry for the Estimation of Femoral Bone Mineral Density and Fragility

Q: Filming location: Will the filming need to take place in multiple locations?   No

Authors: To ensure that your script can be filmed in one day, the Protocol section is restricted to 35 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  13

Number of Shots:  20

# Protocol

NOTE TO THE AUTHORS: 
Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

- To ensure that your protocol can be filmed in one day, the protocol is restricted to 20 shots.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

AUTHORS: Your protocol demonstrates stepwise software usage. So, we will need you to record the shots labeled as SCREEN using screen recording software. If you use a Mac, QuickTime X also has the ability to record the steps. Please create screen capture videos, a screenshot summary, and upload the files to your project page: https://review.jove.com/account/file-uploader?src=20476833

- Protocol

  - To begin, switch on the monitor by pressing the power button on the top left edge of the unit [1]. The operation will take a few seconds [2].

    - WIDE: Talent pressing the power button on the top left edge of the monitor.

  - Tap the icon of the software at the top right of the screen [1]. Insert the user identifier and password, then tap the Start button [2]. Press the power button on the device to switch on the main unit [3].

    - SCREEN: The software icon at the top right of the screen is being tapped.

    - SCREEN: The user enters the identifier and password and then taps the Start button.

    - Talent pressing the power button on the main unit.

  - Next, measure the weight and height of the subject using a standard scale with a stadiometer [1-TXT].

    - Talent measuring the subject's weight and height using a scale with a stadiometer. TXT: The subject should be dressed in undergarments for the procedure

  - Ask the subject to lay supine on a medical bed, with the lower limbs positioned in a neutral position [1].

    - Shot of the subject laid in supine on a medical bed with the lower limbs positioned neutrally.

  - For subject registration, fill in the registration form, including first name, surname, date of birth, gender, menopause age, ethnicity, weight, and height [1]. The software automatically calculates the age and body mass index [2].

    - SCREEN: User filling in the registration form fields, and the software displays automatically calculated age and body mass index.

  - Now, select Femur for the femoral scan [1]. A drop-down menu will automatically open. Select either Right or Left [2].

    - SCREEN: User selecting Femur and then, from the drop-down menu, selecting Right or Left femur.

  - Then, tap the Save button. After a few seconds, a live scan window will be shown [1].

    - SCREEN: User tapping the Save button. A live scan window appearing on the screen is being shown.

  - For scan acquisition, place the probe on the skin above the femoral neck, parallel to the head-neck axis, to visualize the proximal femur profile [1].

    - Talent placing the probe on the skin above the femoral neck, ensuring it is parallel to the head-neck axis.

  - Using the buttons on the left side of the touch screen, adjust the scan depth and the depth of the focus identified by the red line in the ultrasound image [1]. Ensure the target bone interface is in the ultrasound beam focal zone at about halfway through the depth of the B-mode image [2].

    - SCREEN: User adjusting the scan depth and depth of focus using the buttons on the left side of the touch screen.

    - LAB MEDIA: Figure 1B

  - Now, tap the Start button or press the space bar on the keyboard to initialize the scan [1-TXT].

    - SCREEN: User tapping the Start button. TXT: The subject should remain still during the 40 s acquisition

  - After the ultrasound image capture, wait for the software to automatically perform the analysis, showing the message: Analysis in progress - The processing could take a few minutes. Press the stop button to cancel the processing [1].

    - SCREEN: Software displaying the message: "Analysis in progress - The processing could take a few minutes. Press the stop button to cancel the processing."

  - Wait until the end of the analysis to get the report displayed on the screen [1].

    - SCREEN: The report is displayed on the screen.