- Submission ID #: 67155

- Scriptwriter Name: Poornima G

- Project Page Link: https://review.jove.com/account/file-uploader?src=20482323

- Title: Neuronavigated Focalized Transcranial Direct Current Stimulation Administered During Functional Magnetic Resonance Imaging

Landing Page Title (not for video use): Neuronavigated tDCS for Enhanced Electrode Placement Accuracy in fMRI Studies

☒   The Landing Page Title is correct. (Character limit with spaces: 80)

- Authors and Affiliations:

- Filip Niemann1*, Alireza Shahbabaie1*, Sven Paßmann1, Steffen Riemann1, Robert Malinowski1, Harun Kocataş1, Leonardo M. Caisachana Guevara1,2, Mohamed Abdelmotaleb1, Daria Antonenko1, Felix Blankenburg3,4, Rico Fischer2, Gesa Hartwigsen5,6, Shu-Chen Li7,8, Michael A. Nitsche9,10,11, Axel Thielscher12,13, Dagmar Timmann14, Anna Fromm1, Dayana Hayek1, Ann-Kathrin Hubert1, Andrew K. Martin15,16, Alexander Hunold17, Agnes Flöel1,18**, Marcus Meinzer1**

- 1 Department of Neurology, University Medicine Greifswald

- 2 Department of Psychology, University of Greifswald

- 3 Berlin School of Mind and Brain, Humboldt Universität zu Berlin

- 4 Neurocomputation and Neuroimaging Unit, Freie Universität Berlin

- 5 Wilhelm-Wundt-Institute for Psychology, Cognitive and Biological Psychology, Leipzig University

- 6 Lise Meitner Research Group Cognition and Plasticity, Max Planck Institute for Human Cognitive and Brain Sciences

- 7 Lifespan Developmental Neuroscience, Faculty of Psychology, Technische Universität Dresden

- 8 Centre for Tactile Internet with Human-in-the-Loop, Technische Universität Dresden

- 9 Department of Psychology and Neurosciences, Leibniz Research Centre for Working Environment and Human Factors

- 10 German Centre for Mental Health (DZPG)

- 11 Clinic of Psychiatry and Psychotherapy, Protestant Hospital of Bethel Foundation, University Hospital OWL, Bielefeld University

- 12 Danish Research Centre for Magnetic Resonance, Centre for Functional and Diagnostic Imaging and Research, Copenhagen University Hospital Amager and Hvidovre

- 13 Department of Health Technology, Technical University of Denmark

- 14 Department of Neurology and Center for Translational Neuro- and Behavioral Sciences (C-TNBS), Essen University Hospital

- 15 Department of Psychology, University of Kent

- 16 Kent Medway Medical School

- 17 Institute of Biomedical Engineering and Informatics, Technische Universität Ilmenau

- 18 German Center for Neurodegenerative Diseases (DZNE Site Greifswald)

- * and ** These authors contributed equally

☒   All author names and affiliations are correct (city/state/country information not included in video title page).

- Corresponding Authors:

- Filip Niemann		filip.niemann@med.uni-greifswald.de; filipniemann@gmail.com

- Alireza Shahbabaie		              Alireza.Shahbabaie@med.uni-greifswald.de

- Email Addresses for All Authors:

Alireza Shahbabaie			Alireza.Shahbabaie@med.uni-greifswald.de

Sven Paßmann			Sven.Passmann@med.uni-greifswald.de

Steffen Riemann			Steffen.Riemann@med.uni-greifswald.de

Robert Malinowski			Robert.Malinowski@med.uni-greifswald.de

Harun Kocataş			Harun.Kocatas@med.uni-greifswald.de

Leonardo M. Caisachana Guevara	leonardo.caisachanaguevara@uni-greifswald.de

Mohamed Abdelmotaleb		Mohamed.Abdelmotaleb@med.uni-greifswald.de

Daria Antonenko			Daria.Antonenko@med.uni-greifswald.de

Felix Blankenburg			felix.blankenburg@fu-berlin.de

Rico Fischer			rico.fischer@uni-greifswald.de

Gesa Hartwigsen			hartwigsen@cbs.mpg.de

Shu-Chen Li				shu-chen.li@tu-dresden.de

Michael A. Nitsche			nitsche@ifado.de

Axel Thielscher			axelt@drcmr.dk

Dagmar Timmann			dagmar.timmann-braun@uni-due.de

Anna Fromm			anna.fromm@med.uni-greifswald.de

Dayana Hayek			dayana.hayek@med.uni-greifswald.de

Ann-Kathrin Hubert		Ann_hubert@web.de

Andrew K. Martin			A.K.Martin@kent.ac.uk

Alexander Hunold			alexander.hunold@neurocaregroup.com

Agnes Flöel				Agnes.Floeel@med.uni-greifswald.de

Marcus Meinzer			Marcus.Meinzer@med.uni-greifswald.de

Filip Niemann		filip.niemann@med.uni-greifswald.de; filipniemann@gmail.com

Alireza Shahbabaie			Alireza.Shahbabaie@med.uni-greifswald.de

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   No

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  25

Number of Shots:  57 (29 SC)

# Interviews

- Video 1: Author Spotlight: Title (Filled by scriptwriter during script finalization)

Ethics Title Card

This research has been approved by the ethics committee of the University Medicine Greifswald

Videographer: Obtain headshots for all authors.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 50 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Alireza Shahbabaie:  We describe a new method of neuronavigated electrode placement for focal tDCS, that considers the head anatomy of individual study participants. We want to answer the question whether neuronavigated placement of tDCS electrodes results in superior placement accuracy compared to traditional placement methods, and whether this superior accuracy improves the behavioral and neural effects of the stimulation.

What are the most recent developments in your field of research?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What are the current experimental challenges?

  - Filip Niemann: One of the greatest challenges in tDCS experiments is to ensure accurate placement of electrodes on the scalp of each study participant in a way that current delivery is most effective. Therefore, researchers must ensure accurate placement of electrodes and verify potential displacement during fMRI.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Sven Paßmann: First, our approach optimizes  electrode positions on the scalp using realistic computer simulations of current flow for each study participant. Second, we use neuronavigated electrode placement, to ensure that electrodes are attached at individually optimized positions on the scalp. Third, we confirm the accuracy of electrode placement relative to the intended target regions, before and after each fMRI session.

How will your findings advance research in your field?

  - Marcus Meinzer: By improving placement accuracy of optimized electrode positions on the scalp, we ensure that the effects of tDCS are maximized for each study participant.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

What research questions will your laboratory focus on in the future?

  - Agnes Flöel: In the future, we plan to use neuronavigated tDCS to systematically investigate the neural mechanisms by which stimulation improves human brain function. This research will provide fundamental insights into „how“ tDCS modulates human brain function in health and disease. It will also help to maximize tDCS effects based on the characteristics of individual study participants in the future.

Videographer: Obtain headshots for all authors.

# Protocol Videos

Each video will include a section of your protocol and accompanying results, if applicable. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration. Professional voiceover artists will narrate the video.

- Red and italics are pronunciation guides (how the word will be spoken).

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product from that step in advance.

Protocol:

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

Representative Results:

- The three-digit numbers (e.g., 2.3.1., 2.3.2.) are the figures/tables from your manuscript. These will not be recorded by the videographer.

- Please review the result section to make sure it logically follows the video.

- Please note that the video cannot include voiceover without an accompanying visual.

- Video 2: Precise Electrode Positioning in Transcranial Direct Current Stimulation Using Neuronavigation and MRI Data

Demonstrators: Mohamed Abdelmotaleb, Harun Kocataş, Leonardo M. Caisachana Guevara

Ethics Title Card

Procedures involving human subjects have been approved by the ethics committee of the University Medicine Greifswald

Protocol

  - To begin, assemble all the required equipment for neuronavigation [1].

    - WIDE: Talent assembling equipment necessary for neuronavigation. Videographer: Please take a still image of the talent performing this action. Make sure that it is at least a half-body shot with the talent's face visible and zoom out so we have room for cropping.

  - Open the neuronavigation software and choose New Empty Project [1]. Load the participant's T1-weighted image and Save project [2]. To initiate the 3D head reconstruction, navigate to the Reconstructions section [3] and click on New... | Skin (new skin) to open a new window, then press the button to reconstruct the skin [4].

    - SCREEN: Selection of New Empty Project in the neuronavigation software.

    - SCREEN: Loading the T1-weighted image of the participant and selecting Save project.

    - SCREEN: Navigating to the Reconstructions section

    - SCREEN: Selecting New... | Skin and Pressing the button to reconstruct the skin.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=20482323

Also, please let us know if any of the screen capture shots are not performed on the computer so that we can modify the shots accordingly.

  - Configure five landmarks including nasion, left and right nostrils, and left and right preauricular pits [1]. Configure the electrode positions for the right dorsolateral prefrontal cortex with the x, y, and z coordinates of the anode in the Targets section [2]. Click on Add new and name the anode [3-TXT].

    - SCREEN: Configuring five anatomical landmarks on the neuronavigation system.

    - SCREEN: Entering the x, y, and z coordinates of the anode in the Targets section.

    - Clicking on Add new and naming the anode. TXT: Repeat the process for the three return electrodes

Neuronavigated Identification of Electrode Positions

  - Position the participant comfortably in a chair facing the tracking camera [1] and instruct them to put on the goggles with the subject tracker attached [2].

    - Participant taking position in a chair facing the tracking camera.

    - Participant putting on goggles with the subject tracker attached.

  - Then, go to the Sessions tab, and from the New dropdown menu, select Online Session [1]. Choose the Polaris tab and verify the visibility of the subject tracker and the pointer in the camera's field of view [2]. Ensure both devices are recognized when the red crosses change to green check marks in the Tools panel [3].

    - SCREEN: Selection of Online Session from the New dropdown menu in the Sessions tab.

    - SCREEN: Choosing Polaris tab, Verifying visibility of the subject tracker and pointer by moving them into the camera's field of view.

    - Cursor hovering over red crosses changing to green check marks in the Tools panel when devices are recognized.

  - Select the Registration section to register the five predefined landmarks [1]. Place the pointer perpendicularly to the skull with the sensors facing the camera and press the foot pedal to confirm the position [2].

    - SCREEN: Selection of Registration section in the software and registering the five predefined landmarks.

    - Talent positioning the pointer perpendicularly to the participant’s skull with sensors facing the camera and pressing the foot pedal.

  - Navigate to the Validation section and validate the landmarks after placing the pointer on the registered landmarks [1]. Check the two distance indices and move the pointer over the participant’s scalp to verify alignment with the crosshairs of the reconstructed scalp on the monitor [2-TXT].

    - SCREEN: Navigation to the Validation section and pointer placed on the registered landmarks.

    - Crosshair alignment check on the neuronavigation monitor. TXT: Sample additional points around extreme locations

  - Next, in the Refinement landmarks panel, click the Add button for each position [1]. Position the pointer on the target surface of the head and press the pedal to register the position [2-TXT].

    - SCREEN: Clicking Add button in the Refinement landmarks panel for each position.

    - Talent positioning the pointer on the target surface of the participant's head. TXT: Repeat until the distance between the crosshair and reconstructed skin is minimized

  - Then, select the Perform section and move the pointer to the approximate location of the dorsolateral prefrontal cortex [1]. Mark the electrode positions when the pointer aligns with the center of the green crosshair on the screen [2].

    - SCREEN: Selection of Perform section and alignment of the pointer with the dorsolateral prefrontal cortex.

    - Green crosshair aligning with the electrode position on the screen.

  - Move the participant’s hair away from the corresponding scalp area and mark the electrode positions using a skin marker or pen [1-TXT]. Apply a small amount of topical anesthetic cream to the marked electrode positions to reduce physical sensations during tDCS-fMRI (T-D-C-S-F-M-R-I) [2-TXT].

    - Talent moving hair away from the scalp and marking the electrode positions with a marker. TXT: Repeat the process for the three cathodes

    - Talent applying topical anesthetic cream on the marked electrode positions on the participant’s scalp. TXT: tDCS: Transcranial Direct Current Stimulation; fMRI: Functional Magnetic Resonance Imaging

Transcranial Direct Current Stimulation-Functional MRI

  - Insert the power plug into the same power strip as the scanner to use the multi-channel DC stimulator in normal mode [1]. Then, connect it to the outer box using the outer box cable and adapter [2] and connect the inner box cable to both the inner and outer boxes [3].

    - Talent inserting the DC stimulator power plug into the same power strip as the scanner.

    - Talent connecting the DC stimulator to the outer box using the provided cable and adapter.

    - Talent connecting the inner box cable to the inner and outer boxes.

  - Next, apply 1 millimeter of conductive gel evenly on the surface of all electrodes in the focal-tDCS setup [1] and use the electrode fill aid to standardize the gel thickness [2].

    - Talent applying conductive gel to the electrodes of the focal-tDCS 3x1 setup.

    - Shot of using the electrode fill aid to ensure standardized gel thickness.

  - Turn on the DC stimulator, followed by the Panel PC [1]. Double-click the DC-Stimulator MC (M-C) icon and click on Calibrate Stimulator to perform the calibration without the participant connected [2].

    - Talent turning on the DC stimulator and then the Panel PC.

    - SCREEN: Double-clicking the DC-Stimulator MC icon. Selecting the stimulation sequence from the Sequence setup dropdown menu.  Selecting Calibrate Stimulator.

  - Position the participant comfortably near the DC stimulator outside the MRI scanner room [1] and measure the widest part of the participant’s head which is the circumference from the forehead to the occipital bone [2].

    - Shot of Participant sitting near the DC stimulator outside the MRI scanner room.

    - Talent measuring the head circumference from the forehead to the occipital bone.

  - Then, place the electrodes in the spacer to ensure equal spacing of the cathodes around the center anode [1] and attach them to the marked scalp positions [2]. Use an optimal sized EEG (E-E-G) cap without plastic inserts to secure the electrodes [3].

    - Talent placing electrodes in the spacer for equal spacing.

    - Shot of attaching electrodes to the marked scalp positions.

    - Talent placing the EEG cap on the participant's head to secure the electrodes.

  - Connect the electrode cables to the inner box of the DC stimulator to perform an impedance check [1] and select the stimulation sequence from the Sequence setup dropdown menu [2].

    - Talent connecting electrode cables to the inner box.

  - To conduct the Impedance Check, press the respective button on the stimulator interface [1-TXT]. Then, disconnect the inner box from the outer box and insert the outer box into the waveguide of the scanner [2-TXT].

    - SCREEN: Pressing the Impedance Check button on the stimulator. TXT: If the impedance is ≤25 kΩ, proceed to the next step

    - Talent disconnecting the inner box and inserting the outer box into the scanner's waveguide. TXT: Guide the participant into the scanner room

  - Next, ask the participant to sit on the MRI examination table and re-connect the inner box to the outer box inserted in the waveguide [1]. Then, position the participant supine on the MRI examination table, positioning the head in the open head coil [2-TXT].

    - Participant sitting on the MRI table while talent reconnects the inner box to the outer box.

    - Participant lying in a supine position with head placed in the open head coil.
TXT: Use inflatable cushions to stabilize participant’s head

  - Lead the electrode cables through the lower part of the head coil before attaching the upper part and locking it in place [1]. Position the inner box next to the participant on the MRI examination table and move them into the scanner bore [2].

    - Talent guiding electrode cables through the lower part of the head coil and locking the upper part in place.

    - Talent placing the inner box on the MRI table and moving the participant into the scanner bore.

  - Leave the scanner room and inform the participant about the upcoming procedures via the communication interface [1].

    - Talent leaving the scanner room and using the communication interface to speak to the participant.

Concurrent tDCS fMRI

  - On the scanner Panel PC, register the new participant by navigating to Main Menu followed by Examination and Patient Registration [1]. After filling in the required fields [2], go to Program Choice and select the planned imaging protocol [3]. Click on Patient Orientation and select the Head First, Supine position [4]. From the Region of Examination and Laterality dropdown menus, choose Brain and click Examination to proceed [5].

    - SCREEN: Navigating to Main Menu > Examination > Patient Registration on the Panel PC.

    - SCREEN: Participant’s details in the registration form.

    - SCREEN: Selecting the planned imaging protocol in Program Choice.

    - SCREEN: Choosing Head First, Supine in the Patient Orientation menu.

    - SCREEN: Selecting Brain from the Region of Examination and Laterality dropdown menus, then clicking Examination.

  - Follow the on-screen instructions to acquire the planned scans [1]. First, acquire a PETRA (pet-ruh) scan to verify the electrode positions on the participant’s head [2].

    - SCREEN: Instructions appearing for acquiring the planned scans.

    - SCREEN: PETRA scan acquisition showing electrode positions on the participant's head.

  - Inform the participant that two 10-minute resting-state fMRI scans will be conducted [1] and instruct them to maintain their gaze on a fixation cross for the entire scanning period [2].

    - Talent communicating with the participant.

    - Shot of Participant focusing on a fixation cross during the scans.

  - To start the tDCS, press the Init Stimulation (initiate stimulation) button on the Panel PC [1]. Click the Release Start-Trigger button to begin the stimulation with a 10-second ramp-up before starting the functional imaging sequences [2-TXT]. After the functional scans and stimulation period, acquire a second PETRA scan while the electrodes are still attached to the participant’s head [3].

    - SCREEN: Pressing the Init Stimulation button to start tDCS stimulation.

    - SCREEN: Clicking the Release Start-Trigger button to initiate the 10-second ramp-up. TXT: Administer tDCS at 2 mA for 20 min

    - SCREEN: PETRA scan acquisition in progress.

  - Upon completion of the MRI session, disconnect the electrode cables from the outer box and turn off the DC stimulator [1]. Move the participant out of the scanner bore [2] and remove the cap and electrodes from the participant's head [3].

    - Talent disconnecting the electrode cables and turning off the DC stimulator.

    - The participant being moved out of the scanner bore.

    - Removing the EEG cap and electrodes from the participant's head.

Representative Results

  - A significant deviation between intended and actual electrode positions was found both before fMRI and after fMRI [1].

    - LAB MEDIA: Figure 6. Video editor: Highlight the black line arising from the “intended” position.

  - Electrode displacement between pre- and post-fMRI sessions was minimal, with an average drift of minus 0.08 millimeters [1].

    - LAB MEDIA: Figure 6. Video editor: Highlight the flat line connecting the "actual pre-fMRI" and "actual post-fMRI" positions.

  - No significant region-specific differences in electrode placement accuracy were found between right occipitotemporal cortex, left temporo-parietal cortex, and right dorsolateral prefrontal cortex [1].

    - LAB MEDIA: Figure 7. Video editor: Sequentially highlight the data under “intended”, "actual pre-fMRI" and "actual post-fMRI" positions.