# Submission ID #: 67518

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=20593038

Title: GC-MS Analysis of Deuterium Enrichment of Retinol in Serum and Estimation of Vitamin A Total Body Stores

# Authors and Affiliations:

Medoua Nama Gabriel

Centre for Food and Nutrition Research, IMPM

# Corresponding Authors:

# Medoua Nama Gabriel		(gmedoua@impm-cm.org; gmedoua@yahoo.fr)

# Email Addresses for All Authors:

# (gmedoua@impm-cm.org; gmedoua@yahoo.fr)

## Author Questionnaire

1. We have marked your project as author-provided footage, meaning you film the video yourself and provide JoVE with the footage to edit. JoVE will not send the videographer. Please confirm that this is correct.

√ Correct

2. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?   No

3. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  No

4. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: MM/DD/YYYY

When you are ready to submit your video files, please contact our Content Manager, Utkarsh Khare.

Current Protocol Length

Number of Steps:  16

Number of Shots:  38

# Introduction

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Medoua Nama Gabriel: This study presents a GC-MS protocol using deuterium-labeled retinol to estimate vitamin A body stores, improving biomarker accuracy for surveillance and deficiency interventions.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 2.9.1

What significant findings have you established in your field?

  - Medoua Nama Gabriel: GC-MS with deuterium-labeled retinol provides a precise, scalable biomarker of vitamin A body stores, transforming the monitoring of deficiency and excess.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: Figure 2

How will your findings advance research in your field?

  - Medoua Nama Gabriel: Our findings provide a validated GC-MS biomarker using deuterium-labeled retinol to accurately quantify vitamin A body stores beyond unreliable serum retinol measures.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 2.11.1

Ethics Title Card

The protocol was approved by the Ethics Committee of the Ministry of Public Health (No. 2015/02/550/CE/CNERSH/SP), and informed consent was obtained from parents/guardians.

# Protocol

- Extraction of Retinol from Serum, and GC-MS Analysis

Demonstrator: Click here to enter name of demonstrator(s)

Protocol

  - To extract retinoids in serum, allow frozen serum samples to thaw gently at room temperature [1]. Aliquot 400 microliters of serum into a 16 by 100-millimeter disposable culture tube [2]. Add 500 microliters of 0.85% saline solution, 100 microliters of the internal standard, and 5 milliliters of chloroform-methanol solution [3].

    - WIDE: Talent placing frozen serum samples on the bench to thaw at room temperature.

    - Talent aliquoting 400 microliters of serum into a 16 x 100 mm disposable culture tube using a pipette.

    - Talent adds saline solution to the culture tube.  Authors: Please keep the internal standard and chloroform-methanol solution in the frame as well.

  - Then, vortex the sample for 30 seconds to mix thoroughly [1]. Centrifuge the mixture at 1157 g for 10 minutes at 4 degrees Celsius [2]. Using a glass Pasteur pipette, carefully extract the chloroform lower layer and transfer it to a 13 by 100-millimeter disposable culture tube [3].

    - Talent placing the tube on a vortex mixer.

    - Talent placing the tube into a centrifuge.

    - Close-up of a glass Pasteur pipette extracting the lower chloroform layer from the centrifuged sample.

  - Now, dry the collected layer under nitrogen gas in a water bath set to 40 degrees Celsius [1]. Resuspend the residue in 100 microliters of ethanol [2]. After vortexing the sample, sonicate the sample for 30 seconds [3].

    - Talent drying the chloroform layer under a nitrogen gas stream in a water bath at 40 degrees Celsius.

    - Talent resuspending the dried residue in 100 microliters of ethanol.

    - The sample is being sonicated.

  - Then, transfer the resuspended sample into a crimp vial fitted with an insert [1]. Close the vial securely and label it appropriately [2-TXT].

    - Talent transferring the resuspended sample into a crimp vial with an insert using a pipette.

    - Talent sealing the crimp vial and attaching a label with clear identification details.  TXT: The sample is ready for HPLC retinol collection

  - Before injecting the serum samples, inject 70 microliters of the retinol working solution into the HPLC (H-P-L-C) system [1] and record the retention time of the retinol peak [2]. Set up the fraction collector to collect the eluted retinol fraction within a specified range in a 5-milliliter tube with a ground-glass stopper [3-TXT].

    - Talent injecting the retinol working solution into the HPLC system.

    - Retention time of the retinol peak displayed on the HPLC system.

    - Fraction collector set to collect eluted retinol fraction from 7.5 to 10.5 minutes, displayed in the system interface.   TXT: Collect the retinol fraction over a 3 min interval, from 7.5–10.5 min

  - Then, inject 70 microliters of each serum sample into an HPLC system equipped with a C18 column and a diode-array detector set to 340 nanometers [1]. Program the pump to deliver Mobile Phases A and B at a constant flow rate of 1 milliliter per minute [2].

    - Serum sample being injected into the HPLC system, showing the C18 column and diode-array detector settings at 340 nanometers.

    - The pump schedule displayed Mobile Phases A and B, with flow rate set to 1 milliliter per minute.

  - Now, dry the collected retinol fraction from HPLC under nitrogen gas in a water bath set to 40 degrees Celsius for at least 3 hours [1]. Add ethanol during drying to facilitate moisture evaporation [2]. Add 20 microliters of BSTFA (B-S-F-T-A) with 10% TMCS (T-M-C-S) to the tube [3].

    - Talent placing the retinol fraction in a water bath at 40 degrees Celsius under a nitrogen gas stream for drying.

    - Talent adding ethanol during the drying process.

    - Talent adding 20 microliters of BSTFA with 10% TMCS to the completely dry retinol fraction.

  - Now, place the tube containing the derivatized retinol sample in a dry block heater set to 70 degrees Celsius for 30 minutes [1]. Using a glass Pasteur pipette, transfer the reaction mixture into a crimp vial fitted with an insert [2], seal it tightly with a crimp cap using a cap crimper, and label the vial [3-TXT].

    - Talent placing the tube in a dry block heater and setting the temperature to 70 degrees Celsius for incubation.

    - Talent using a glass Pasteur pipette to transfer the reaction mixture into a crimp vial with an insert.

    - Talent sealing the crimp vial using a cap crimper and attaching a label.  TXT: Store the sample in a desiccator at 4 °C until analysis; Derivatize standard solutions for GC/MS

  - For GC/MS (G-C-M-S) analysis, inject 3 microliters of the derivatized retinol sample into the gas chromatograph using an autosampler equipped with a cool on-column injector [1-TXT].

    - Shot of the autosampler injecting 3 microliters of the derivatized retinol sample into the GC.   TXT: Injector connects to a 15m fused silica capillary column (0.25mm ID, 0.25µm DB-1)

  - Program the column oven and on-column injector temperatures to increase from 50 degrees Celsius to 285 degrees Celsius at a rate of 15 degrees Celsius per minute [1]. Set the GC/MS interface temperature to 285 degrees Celsius [2]. Use helium as the carrier gas to elute the trimethylsilyl derivative of retinol over approximately 12 minutes [3].

    - Temperature program for the column oven and injector, showing the increase from 50 degrees Celsius to 285 degrees Celsius at a rate of 15 degrees Celsius per minute.

    - GC/MS interface temperature setting at 285 degrees Celsius displayed on the interface.

    - Helium is being selected as the carrier gas

  - Using 0.5 torr methane negative ion chemical ionization, detect the GC eluate with a quadrupole mass spectrometer [1]. Set the ion source temperature to 150 degrees Celsius and configure the mass spectrometer to scan between 260 and 280 daltons [2].

    - Quadrupole mass spectrometer interface showing the methane negative ion chemical ionization settings with 0.5 torr pressure.

    - Ion source temperature set to 150 degrees Celsius and mass spectrometer configured to scan between 260 and 280 daltons.

- Data Analysis and Estimation of Vitamin A Stores

  - Process the chromatographic data using the GC/MS data analysis software. Extract all desired ions based on their mass-to-charge ratios [1], integrate the area under the peaks [2], and transfer the results to a spreadsheet file for further calculations [3].

    - GC/MS data analysis software interface showing chromatograms and extracted ion traces.

    - Manual integration interface for adjusting peak start and end points.

    - Spreadsheet showing transferred peak area data.

  - Calculate the sum of peaks for labeled retinol by adding the peak areas at mass-to-charge ratios of 274, 275, 276, 277, and 278 [1]. Then, calculate the sum of peaks for unlabeled retinol by adding the peak areas at mass-to-charge ratios of 268, 269, and 270 [2].

    - Spreadsheet showing calculations for labeled retinol by summing peak areas at m/z 274 to 278.

    - Spreadsheet showing calculations for unlabeled retinol by summing peak areas at m/z 268 to 270.

  - Calculate the enrichment of labeled retinol [1] and the amount of unlabeled retinol using the presented equations [2].

    - TEXT ON A PLAIN BACKGROUND
D = SD/(SH + SD).

    - TEXT ON A PLAIN BACKGROUND
H = 1 – D

  - To control the system, calculate a linear regression equation between the weight ratios of the calibration standards and the integrated areas for labeled retinol and non-labeled retinol [1].

    - Linear regression equation and associated graph generated from the calibration data.

  - Calculate the total liver stores of vitamin A using the Olson equation [1]. Then, calculate the total body stores of vitamin A using the presented Green equation [2].

    - TEXT ON A PLAIN BACKGROUND
Vitamin A Total Liver Stores = F x Dose x [S x a x (H/D – 1)]

    - TEXT ON A PLAIN BACKGROUND 
TBS = Fa × S × (1/SAp)

# Results

- Results

  - The injection of 3 microliters of a derivatized sample containing approximately 50 picomoles per microliter of calibrants showed no molecular ion [1] but showed a major fragment ion at mass-to-charge ratio 268 for retinol [2] and at 278 for D₈-retinol [3].

    - LAB MEDIA: Figure 1.

    - LAB MEDIA: Figure 1. Video editor: Highlight the tall peak labelled 268.0 in the lower graph

    - LAB MEDIA: Figure 2. Video editor: Highlight the tall peak labelled 278.0 in the lower graph

  - Injection of mixtures of retinol and D8-retinol showed 2 major fragment ions at mass-to-charge ratio 268 for retinol [1] and mass-to-charge ratio 276 for D8-retinol [2].

    - LAB MEDIA: Figure 3. Video editor: Highlight the peak labelled 268.0 in the lower graph

    - LAB MEDIA: Figure 3. Video editor: Highlight the peak labelled 276.0 in the lower graph

  - The calibration curve exhibited excellent linearity [1], with a high correlation coefficient [2].

    - LAB MEDIA: Figure 4. Video editor: Show the full scatter plot with the dotted trend line

    - LAB MEDIA: Figure 4. Video editor: Zoom in on the text showing R squared equals 0.975

  - The gas chromatography and mass spectrometry response obtained from serum samples from children who received an oral dose of 2 milligrams of retinol equivalents of D8-retinol showed the presence of retinol and D8-retinol [1], with 2 major fragment ions at mass-to-charge ratio 268 [2] and mass-to-charge ratio 274 [3].

    - LAB MEDIA: Figure 5.

    - LAB MEDIA: Figure 5. Video editor: Highlight the peak labelled 268.0 in the lower graph

    - LAB MEDIA: Figure 5. Video editor: Highlight the peak labelled 274.0 in the lower graph