# Submission ID #: 67520

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20593648

# Title: Equibiaxial Stretching Device for High Magnification Live-Cell Confocal Fluorescence Microscopy

# Authors and Affiliations:

# Anabel-Lise Le Roux1, Valeria Venturini1, Manuel Gómez1, Amy E. M. Beedle1,2, Xarxa Quiroga1, Xavier Menino3, Xavier Trepat1,4,5,6, Pere Roca-Cusachs1,6

# 1Institute for Bioengineering of Catalonia (IBEC), Barcelona Institute of Science and Technology (BIST)

# 2 Department of Physics, King's College London

# 3ICFO-Institut de Ciencies Fotoniques, The Barcelona Institute of Science and Technology

# 4Centro de Investigación Biomédica en Red en Bioingeniería, Biomateriales y Nanomedicina (CIBER-BBN)

# 5Institució Catalana de Recerca i Estudis Avançats (ICREA)

# 6University of Barcelona (UB)

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# Anabel-Lise Le Roux		aleroux@ibecbarcelona.eu

# Xavier Trepat			xtrepat@ibecbarcelona.eu

# Pere Roca-Cusachs		proca@ibecbarcelona.eu

# Email Addresses for All Authors:

Valeria Venturini 		vventurini@ibecbarcelona.eu

Manuel Gómez		mgomez@ibecbarcelona.eu

Amy Beedle			amy.beedle@kcl.ac.uk

Xarxa Quiroga			xarxaquiroga@gmail.com

# Xavier Menino			xmenino@icfo.net

# Anabel-Lise Le Roux		aleroux@ibecbarcelona.eu

# Xavier Trepat			xtrepat@ibecbarcelona.eu

# Pere Roca-Cusachs		proca@ibecbarcelona.eu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Our own epifluorescence microscope, we will record screen-videos

If Yes, can you record movies/images using your own microscope camera?

Yes

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, Matlab, we will record screen

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   No

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  18

Number of Shots:  60 (23 SC)

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Anabel-Lise Le Roux: We study how cells respond and adapt to mechanical stimuli in physiological and pathological context. For this we use the stretching device to apply external forces, a strain in this case, to cells.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What significant findings have you established in your field?

  - Manuel Gomez: We have found how cells and cell tissues react to specific mechanical stimuli, and how they mechanically respond to other biological cues.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Valeria Venturini : Our stretching device is very versatile compared to others as it offers the possibility to stretch cells cultured on polyacrylamide gels of different stiffnesses or to perform compression experiments.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research questions will your laboratory focus on in the future?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

Answers to these questions will not appear in the video but may be featured in our journal's promotional materials.

- Enter the full name of the author who will deliver the statement.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Answers will be mildly edited for clarity.

- Limit the length of each statement to 50 words or fewer.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Anabel-Lise: We diffuse a full protocol of our custom-made stretching device to the community, showing its reliability and versatility and how to build it from zero. We hope the community can benefit from this compiled information while becoming aware of the many use of the device, to advance research in cell mechanobiology.

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- PDMS Membrane Sandwiching in stretch-rings

Demonstrator: Valeria Venturini, Anabel-Lise Le Roux

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, place the PMMA rings on a sheet of white bench paper [1-TXT]. Carefully go around the edge of the PDMS-coated plate using a blade to ensure a clean cut along the entire circumference [2-TXT]. Use tweezers to lift the membrane at one edge, then gently peel it with the thumbs to evenly detach half of the membrane from the plate [3].

    - WIDE: Talent placing the PMMA rings placed on white bench paper. TXT: PMMA: Polymethyl Methacrylate

    - Talent using a blade to cut along the PDMS edge. TXT: PDMS: Polydimethylsiloxane

    - Talent peeling and lifting the membrane evenly along half of the plate circumference.

  - Use fingers to spread 96% ethanol evenly around the entire top surface of the PMMA ring [1]. Carefully remove the PDMS membrane from the plate by lifting it from the previously lifted edge [2], then place it symmetrically onto the ethanol-covered PMMA ring, ensuring no wrinkles while avoiding stretching [3-TXT].

    - Talent wiping ethanol on the PMMA ring with fingers.

    - Talent lifting the PDMS membrane from the previously lifted edge.

    - Close-up Shot of placing the membrane onto the ethanol-covered ring without wrinkles. TXT: Let the ethanol dry up and the PDMS membrane adhere to the PMMA ring

  - Then, take another empty PMMA ring and place it on top of the mounting ring [1]. Use metal clips to seal them together, sandwiching the PDMS membrane between the two rings [2]. Position the bottom half of the stretch ring onto the mounting support [3]. Place the PDMS ring sandwich onto the stretch ring and push it down firmly until fully seated, ensuring the gasket is correctly positioned [4]. Place the circular brass weights on top of the sample [5].

    - Talent placing an empty PMMA ring over the mounting ring.

    - Talent securing the rings together using metal clips.

    - Talent positioning the stretch ring onto the mounting support.

    - Talent pressing the PDMS ring sandwich into the stretch ring, ensuring the gasket is aligned.

    - Talent placing circular brass weights on top.

  - Next, place the top half of the metal stretch ring on top, aligning the screw holes [1]. Insert all screws and tighten them [2]. Cut the PDMS around the ring gently using a blade [3]. Cover the assembly with a Petri dish to protect the membrane from dust [4-TXT].

    - Talent aligning the top half of the stretch ring and screw holes.

    - Talent inserting and tightening the screws.

    - Talent cutting excess PDMS around the ring with a blade.

    - Talent covering the setup with a Petri dish to prevent dust contamination. TXT: Attach a polyacrylamide (PAA) gel on top of the PDMS membrane

- Stretching Device Operation, Strain Calibration and Resolution Assessment

Demonstrator: Manuel Gomez, Anabel-Lise Le Roux

  - Prepare a 1 to 50000 dilution of 0.2-micrometer fluorescent beads in phosphate buffer saline using serial dilution [1] and place a 10 to 30 microliter drop of the bead solution at the center of the PDMS [2]. Allow it to dry at a temperature no higher than 40 degrees Celsius to avoid post-curing confounding issues [3].

    - Talent mixing the fluorescent bead with water.

    - Talent pipetting a drop of the bead solution onto the PDMS center.

    - Close-up of the PDMS membrane being placed in an incubator.

  - Now, mount the stretching device onto the microscope stage holder and secure the stretching post onto the microscope [1]. Generously apply lubricant to the circular area of the post [2]. Mount the prepared stretch ring onto the stretching post and screw the top cap to seal the assembly [3]. Connect the vacuum controller to the vacuum source and the stretching post [4]. Then, connect the rear of the vacuum control box to the vacuum source followed by the front of the box to the stretch post [5].

    - Talent securing the stretching post on the microscope.

    - Talent applying lubricant to the circular area of the post.

    - Talent mounting and sealing the stretch ring assembly.

    - Talent connecting the vacuum controller to the vacuum source and stretching post.

    - Close-up of the vacuum controller connected to source and stretch post.

  - To perform strain calibration, choose a field of interest containing a high bead density and recognizable bead clumps for the user [1]. Acquire an image at rest of the selected field of view and save the position [2]. Apply stretch at a low vacuum level of negative 20 millibar, causing movement in all directions [3]. Now, refocus in the Z-plane (zee-plane) and move the stage in X-Y directions to locate the same region of interest and acquire an image in the stretched state [4].

    - SCREEN: Selecting field of view with high bead density and bead clumps.

    - SCREEN: Image acquisition interface showing the field of view saved at rest.

    - Talent pressing the ON button of the vacuum controller to apply low vacuum stretch at negative 20 millibar.

    - SCREEN: Cell moving out of the focused field of view

    - SCREEN: Adjusting focus in the z-plane and moving the stage in x-y to realign the imaged region.

    - SCREEN: Acquire the image of the stretched beads.

    - SCREEN: Go to the saved position at rest

    - SCREEN: destretch beads and refocus in z. The image of the beads after detretchech can be recorded again and is expected to match the initial image taken before stretch.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=20593648

  - Then, release the strain and repeat the steps with a higher vacuum level [1-TXT]. Store acquired images in a folder organized by sample [2]. Name the reference bead image acquired without stretch as reference_str1 (reference-S-T-R-1) and the images stretched at increasing vacuum levels as stretched_str1, stretched_str2, stretched_str3 (stretched-S-T-R-1-2-3) and so on [3].

    - SCREEN: Acquisition at higher vacuum levels. TXT: Repeat the strain calibration steps for at least 3 different stretch rings

    - SCREEN: Go to the folders being created for the acquired images by sample.

    - SCREEN: File naming convention for reference and stretched images.

  - To quantify the strain, open MATLAB and load the gel_strain (gel strain) code [1-TXT]. After setting the parameters, define the parent as well as the experiment folder paths and run the code [2].

    - Talent: User in front of his personal computer straing the analysis

    - SCREEN: MATLAB workspace with "gel_strain" code being opened. TXT: https://github.com/xt-prc-lab/Le_Roux_et_al_2024_JOVE

    - SCREEN: Defining folder paths in the code and running the "gel_strain" code.

  - Then, open stretched_str1 as requested by the code, followed by the reference_str1 [1-TXT]. Now, calculate the median strain matrix for each vacuum level [2]. After repeating this process for at least three samples, plot the median strain matrix versus vacuum level to obtain the calibration curve [3].

    - SCREEN: Selection of "stretched_str1" for processing and Selection of "reference_str1". TXT: Verify that the results folder appears with 4 subfolders

    - SCREEN: Open the sub-folders containing the plots and spreadsheet of the calculated strain matrix.

- Cell Seeding, Stretching and Image Acquisition

Demonstrator: Click here to enter name of demonstrator(s)

  - To coat the PDMS ring, sterilize it under ultraviolet light for 15 minutes [1]. Under a biosafety cabinet, rinse the plate under the hood with sterile PBS [2] and deposit a 100 to 200-microliter drop of the solution in the center of the ring [3]. Draw the perimeter of the drop with a permanent marker from the bottom side of the PDMS membrane and incubate it [4].

    - Talent placing the PDMS ring under ultraviolet light for sterilization.

    - Talent adding the PDMS membrane with sterile PBS.

    - Close-up Shot of pipetting a cell adhesion protein drop onto the membrane.

    - Close-up of the marker-drawn perimeter around the drop.

  - On the day of the experiment, trypsinize the cells and count them [1]. Seed an appropriate number of cells in a small volume of about 20 to 50 microliters with standard cell media to promote attachment [2] and incubate it in a carbon dioxide humidity-controlled incubator [3].

    - Talent placing the samp eina cell counter. TXT: If needed, transfect the cells of interest with fluorescent protein beforehand

    - Talent pipetting 20–50 microliters of cells onto the PDMS ring.

    - Talent placing the sample in an incubator.

  - After 20 to 30 minutes, gently add 500 microliters of media on top of the cells [1] and return the plate to the incubator to allow for further cell spreading [2]. Replace the media with carbon dioxide-independent media and immediately proceed with the stretching experiment [3].

    - Talent adding 500 microliters of media after incubation.

    - Talent placing the sample in an incubator.

    - Talent replacing media with carbon dioxide-independent media.

  - For a single stretch-release cycle and imaging, mount the stretch ring onto the microscope stage as demonstrated earlier [1]. After locating a cell of interest, record its position, and capture an image of the pre-stretched cell [2]. Apply the desired vacuum level corresponding to the required substrate strain based on the calibration curve [3].

    - Talent mounting the stretch ring onto the microscope stage.

    - SCREEN: Focus a cell of interest

    - SCREEN: Microscope interface showing the recorded position of the cell at rest.

    - SCREEN: Record an image of the cell at rest (before stretch application).

    - Talent:  Set Vacuum level according to calibration curve (6% here for fibroblasts) on the vacuum controller (record user setting the vacuum level)

    - Talent: Apply vacuum by turning he vacuum controller ON.

  - To locate the stretched cell, first refocus in the Z-plane, then move in the X-Y directions to center the cell in the field of view [1]. Record images during the desired duration [2]. Navigate to the cell’s original position and activate the microscope’s acquisition mode to ensure that image capture begins immediately upon refocusing [3]. Release the stretch and quickly refocus manually in the Z-plane [4]. Capture images of the released cell for the required time, manually adjusting the focus throughout the acquisition [5].

    - SCREEN: Show the cell moving out of the focal plane

    - SCREEN: Adjusting focus in z-plane and moving in x-y directions.

    - SCREEN: Image acquisition interface capturing cell under stretch.

    - SCREEN: Activation of image acquisition mode before refocusing.

    - SCREEN: Stretch release and immediate z-refocusing.

    - SCREEN: Image capturing of released cell with manual focus adjustments.

  - For post-processing, open Fiji and load the images acquired at rest and during strain [1]. Use the Fiji plugin template matching to select a region of the cell for alignment [2].

    - SCREEN: Fiji software with images being loaded.

    - SCREEN: Fiji plugin template matching interface showing cell alignment process.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 122.

- Please note that the video cannot include voiceover without an accompanying visual.

- Representative Results

  - Calibration ensured reproducibility and control over applied strain, confirming a nearly linear relationship between vacuum and strain up to 20% [1].

    - LAB MEDIA: Figure 3C. Video editor: Highlight the calibration curve showing the strain increasing with vacuum.

  - Radial displacement analysis demonstrated isotropy and homogeneity of strain over the PDMS membrane [1].

    - LAB MEDIA: Figure 3B. Video editor: Mark the row showing “radial displacement” distribution.

  - Fluorescence imaging of fibroblast cells under live-cell stretching revealed membrane evaginations upon stretch release [1], which reabsorbed within 3 to 5 minutes [2].

    - LAB MEDIA: Figure 4B and C. Video editor: Highlight the magenta-stained border in C for column “Stretch (magenta) versus initial state (green)”.

    - LAB MEDIA: Figure 4B and C. Video editor: Highlight the magenta-stained border in C for column “Stretch (magenta) versus relaxed state (green)”.

  - Restretching of fixed cells enabled imaging in a stretched state, preventing artificial folding artifacts [1].

    - LAB MEDIA: Figure 4D. Video editor: Highlight the image labeled “ Restretched fixed cells”.

  - PDMS membrane with a patterned grid facilitated multi-modal imaging, allowing localization of the same cells in fluorescence and scanning electron microscopy [1].

    - LAB MEDIA: Figure 4F.

  - Soft PAA gel coatings influenced cell adhesion and spreading under cyclic stretch, with fibroblasts forming focal adhesions on stiff substrates but remaining rounded on soft ones [1].

    - LAB MEDIA: Figure 4E.