# Submission ID #: 67538

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20599138

# Title: Mechanical Mapping of Spheroids Using Brillouin Spectroscopy

# Authors and Affiliations:

# Chloe B Rodgers1,*, Giedrė Astrauskaitė1,*, Rebecca E. Ginesi1, Greg Wardle2, Troy R Allen2, Oana Dobre1, Massimo Vassalli1, Matthew Walker1

# 1Centre for the Cellular Microenvironment, Advanced Research Centre, University of Glasgow

# 2LightMachinery Inc

*These authors contributed equally

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Matthew Walker                      matthew.walker@glasgow.ac.uk

# Email Addresses for All Authors:

Chloe B Rodgers                      chloe.rodgers@glasgow.ac.uk

Giedrė Astrauskaitė                giedre.astrauskaite@glasgow.ac.uk

Rebecca E. Ginesi                    rebecca.ginesi@glasgow.ac.uk

Troy R Allen                              tallen@lightmachinery.com

Greg Wardle                            gwardle@lightmachinery.com

Oana Dobre                              oana.dobre@glasgow.ac.uk

Massimo Vassalli                     Massimo.Vassalli@glasgow.ac.uk

Matthew Walker                      matthew.walker@glasgow.ac.uk

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps: 24

Number of Shots: 56 (27 SC)

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer up to 2 introduction and up to 3 conclusion questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Answer in full sentences, the questions will not be displayed in the video.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What questions will future research focus on?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Preparation of Spheroids

Demonstrator: Click here to enter name of demonstrator(s)

  - To begin, pre-treat the wells of the microwell plate by adding 500 microliters of anti-adherence rinsing solution to each well [1]. Centrifuge the plate at 1,300 g for 5 minutes at room temperature [2]. After eliminating air bubbles, remove the anti-adherence rinsing solution from the wells [3].

    - WIDE: Talent adding 500 microliters of anti-adherence rinsing solution into each well of the microwell plate using a pipette.

    - Talent placing the plate inside the centrifuge and setting it to 1,300 g for 5 minutes at room temperature.

    - Talent aspirating and removing the anti-adherence rinsing solution from the wells.

  - Next, add 1 milliliter of PBS to each well and aspirate it completely [1].  After repeating the PBS wash, determine the number of cells per spheroid and the number of spheroids per gel according to the experimental design [2-TXT].

    - Talent dispensing 1 milliliter of PBS into each well and aspirating it out.

    - Talent recording or noting the calculated cell number in a notebook. TXT: Calculate the number of cells required to obtain 1.2 × 10⁵ cells/well

  - Pellet the cells and resuspend them in 1 milliliter of supplemented DMEM [1]. Then add 1 milliliter of this cell suspension into one well of the microwell plate and mix well [2]. Centrifuge the plate at 100 g for 3 minutes at room temperature [3].

    - Talent resuspending the cell pellet in 1 milliliter of the described culture medium.

    - Talent adding 1 milliliter of the cell suspension into a microwell plate well.

    - Talent placing the plate into the centrifuge.

  - Observe the microwell plate under the microscope to confirm even distribution of cells and that they have settled at the bottom of the wells [1]. Incubate the cells at 37 degrees Celsius with 5 percent carbon dioxide for 24 hours to allow spheroid formation [2].

    - Talent placing the plate under a microscope.

    - Talent placing the plate into an incubator set to 37 degrees Celsius and 5 percent carbon dioxide.

  - To harvest the spheroids, gently and slowly remove the culture media without disturbing the spheroids [1]. Cut the PBS-FBS-coated pipette tip to widen the opening [2] and add 500 microliters of PBS to each well to dislodge the spheroids [3]. Collect the PBS containing the spheroids into a tube [4].

    - Talent aspirating and removing the culture media carefully from the microwell plate.

    - Talent cutting the pipette tip to create a wider opening.

    - Talent adding 500 microliters of PBS to each well and pipetting up and down to dislodge spheroids.

    - Talent transferring the PBS containing spheroids into a collection tube.

  - Then, transfer 50 microliters of the spheroid suspension into one well of a 96-well plate [1]. To aid counting, draw a cross on the bottom of the well using a permanent marker [2] and count the number of spheroids in this volume under the microscope [3].  Use the equation to calculate the total number of spheroids: [4]

    - Talent transferring 50 microliters of spheroid suspension into a 96-well plate.

    - Talent drawing a cross under the well with a permanent marker to assist with counting.

    - Talent viewing the sample under a microscope.

    - TEXT ON PLAIN BACKGROUND:

No. of spheroids=Volume of spheroid suspension μL50 μL × Spheroid count in 50 μL #.

- Setting up the Brillouin Microscope

Demonstrator: Click here to enter name of demonstrator(s)

  - Open the SpectraLok software and wait until all connected devices are detected [1-TXT]. Then, select the spectrometer camera [2].

    - SCREEN: Show SpectraLok software initializing and detecting all connected devices. TXT: Turn on and check all the devices beforehand

    - SCREEN: Cursor selecting the spectrometer camera from the left panel of the main window.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible : https://review.jove.com/account/file-uploader?src=20599138

  - Install the acrylic cube on the microscope sample holder with the clear side facing the objective lens [1]. Lower the objective lens as much as possible and center the cube beneath it [2]. Ensure that the optical path in the microscope is set so that the spectrometer camera is exposed by turning the rotary port selector to L [3].

    - Talent placing the acrylic cube onto the microscope stage, clear side facing the objective lens.

    - Close-up shot of the objective lens being lowered and aligned over the cube.

    - Talent adjusting the rotary port selector on the microscope to the L position.

  - In the SpectraLok Camera window, set the sensor exposure time to 500 milliseconds and the software gain to 100 [1-TXT].

    - SCREEN: Show the SpectraLok Camera window and the adjustment of exposure time to 500 milliseconds and software gain to 100. TXT: Unblock the laser shutter

  - Press Capture and slowly move the objective lens upward [1]. Continue raising the lens until the Brillouin peaks become visible and their intensity stabilizes, indicating that the focal spot is positioned within the acrylic cube [2].

    - SCREEN: Cursor clicking on Capture in the SpectraLok interface.

    - SCREEN: Display the transition from a saturated reflection pattern to distinct  illouin peaks, showing stable signal intensity.

  - Now, open the Pump Killer (PK) control window [1].

    - SCREEN: Show SpectraLok interface with the Pump Killer (PK) control window being opened.

  - Zoom in on one of the brightest Brillouin orders and adjust the pressure actuator position in increments of 100 micrometers by clicking Move Rel to minimize the Rayleigh peak intensity [1]. If the Rayleigh signal increases, click Reverse to change the adjustment direction [2].

    - SCREEN: Show zoomed-in view of the Brillouin order and Cursor clicking Move Rel to adjust the actuator position in 100 micrometer steps.

    - SCREEN: Cursor selecting Reverse when the Rayleigh signal increases.

  - Once the Rayleigh peak intensity decreases, fine-tune the etalon pressure in smaller increments of 10 micrometers [1]. In the camera window, zoom in on one of the Brillouin orders [2].

    - SCREEN: Fine-tuning pressure in 10 micrometer increments until optimal intensity is achieved.

    - SCREEN: Show zoomed-in view in the camera window highlighting a single Brillouin order.

  - Press S to visualize stripes across which the LUT is applied to the camera image [1]. Open the Settings window and press Quick Calibrate to automatically recalculate the optimal horizontal offset of the stripes [2-TXT].

    - SCREEN: Show the camera image with purple stripes visible across the Brillouin peaks.

    - SCREEN: Cursor clicking Settings, then selecting Quick Calibrate to adjust stripe alignment. TXT: If the sensor becomes saturated, reduce the exposure time

  - After verifying the stripe-alignment, adjust the collimator lens axes to maximize the coupling efficiency of light from the Pump Killer to the spectrometer [1]. Then, open the Spectrum window and click Unwrap to retrieve the spectrum [2].

    - SCREEN: Adjusting The Collimator Lens Axes.

    - SCREEN: Show Spectrum window being opened and Unwrap clicked to display the spectrum.

  - In the PK Control window, iteratively adjust the collimator axes 1 and 2 in increments of 0.0005 degrees to maximize the Brillouin signal [1]. In a homogeneous acrylic cube, aim for complete suppression of the Rayleigh signal and adjust until the Brillouin signal is clearly visible [2-TXT].

    - SCREEN: Cursor adjusting collimator axis values in 0.0005-degree increments in the PK Control window.

    - SCREEN: Final optimized view showing fully suppressed Rayleigh signal and distinct Brillouin peaks. TXT: Brillouin signal should be comparable to reference amplitude

  - To save the optimized parameters, click Refresh All and then Save All in the PK Control window [1]. Finally, in the main window, click Save in the File I/O (I-O) section to store the configuration [2].

    - SCREEN: Cursor clicking Refresh All and Save All in the PK Control window.

    - SCREEN: Cursor clicking Save under the File I/O section in the main interface.

- Mounting the Sample and Image Acquisition

Demonstrator: Click here to enter name of demonstrator(s)

  - After opening the scanning software, install the glass-bottom petri dish containing the hydrogel-spheroid sample in place of the acrylic cube [1-TXT].

    - Talent placing the glass-bottom petri dish with the hydrogel-spheroid sample onto the microscope stage. TXT: Incubator settings: 37 °C and 95% relative humidity

  - Remove the petri dish lid, add sufficient media to submerge the gel [1], and place a coverslip on top to prevent gel movement [2].

    - Talent removing the petri dish lid and adding media to fully submerge the gel.

    - Close-up shot of the coverslip being placed gently over the gel to prevent movement.

  - Locate the spheroid by first blocking the laser beam [1]. Switch the microscope optical path to the eyepiece and turn on the white light illumination [2]. Using the joystick, manually control the sample stage to locate and focus on the spheroid [3].

    - Talent rotating the wheel to block the laser beam.

    - Talent switching the optical path selector to the eyepiece and turning on white light illumination.

    - Talent using the joystick to move the stage and bring the spheroid into focus through the eyepiece.

  - Acquire a brightfield image of the spheroid and adjust the illumination brightness by rotating the control wheel [1]. Switch the optical path to R to expose the microscope camera [2] and set the exposure time to approximately 50 milliseconds or as needed to achieve a high-contrast image [3].

    - SCREEN: Show the camera view capturing a brightfield image of the spheroid and adjusting the illumination.

    - Talent rotating the optical path selector to R for camera exposure.

    - SCREEN: Display exposure time adjustment to 50 milliseconds to enhance image contrast.

  - Next, set the microscope port to L and capture the background with the laser still blocked [1]. Unblock the laser and depending on the desired signal-to-noise ratio, select an exposure time between 200 and 500 milliseconds [2-TXT].

    - Talent setting the port selector to L.

    - SCREEN: Show adjustment of exposure time within the 200–500 millisecond range based on signal quality. TXT: Re-optimize etalon pressure and Pump Killer–spectrometer coupling if disturbed

  - From the unwrapped spectrum, select the Brillouin peak fitting range [1]. Adjust the threshold value to approximately half the Brillouin peak intensity [2] and exclude any Brillouin peaks with amplitude below the threshold [3]. Apply wavelength correction to the laser according to all identified Brillouin peak pairs [4].

    - SCREEN: Show the unwrapped spectrum with the Brillouin fitting range selected.

    - SCREEN: Cursor adjusting the threshold slider to about half of the Brillouin peak amplitude.

    - SCREEN: Software disregarding lower-intensity peaks automatically.

    - SCREEN: Demonstrate wavelength correction being applied based on the identified Brillouin peak pairs.

  - Finally, choose the raster scan geometry by adjusting the scan width and step size [1] and start the scan to record the measurements [2].

    - SCREEN: Show the scan region setup with selected width and step size before starting the scan.

    - SCREEN: Display progress of the scan and acquisition.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 115.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - Representative images of a mesenchymal stem cell spheroid embedded in hydrogel showed distinct Brillouin Shift [1] and Brillouin Linewidth [2] maps, along with corresponding brightfield [3], nuclei [4], and actin filament fluorescence images [5].

    - LAB MEDIA: Figure 9A.

    - LAB MEDIA: Figure 9B.

    - LAB MEDIA: Figure 9C.

    - LAB MEDIA: Figure 9D.

    - LAB MEDIA: Figure 9E.

  - The low-resolution Brillouin Frequency Shift map was segmented into gel and spheroid regions [1] based on the contour manually traced from the brightfield image [2].

    - LAB MEDIA: Figure 10A.

    - LAB MEDIA: Figure 10B.

  - Ambiguous pixels resulting from down-sampling were excluded from analysis [1], while clear binary masks for the gel [2] and spheroid regions were created to extract BFS values [3].

    - LAB MEDIA: Figure 10B. Video editor: Highlight the “Excluded values” panel

    - LAB MEDIA: Figure 10C.

    - LAB MEDIA: Figure 10D.

  - Comparison of high- and low-resolution BFS maps demonstrated that coarse sampling provided mean BFS values comparable to high-resolution scans, indicating consistent measurement of spheroid bulk mechanics [1].

    - LAB MEDIA: Figure 11. Video editor: Highlight the corresponding BFS values “6.48…..” in the bottom row’s three images

NOTE to Authors:

- It's better if you upload each screen recording after naming each file according to the corresponding shot number (3-digit bullet, e.g, 2.2.1.mp4, 2.1.3.mp4, etc).

Each clip’s duration should be less than 25 seconds.

- But if you prefer to record all the steps together and upload a single screen capture file, that is okay.

In this case, you need to indicate the corresponding timestamp against each shot.

For example if the action described in shot 2.1.2 is present in the file name xxx.mp4 and is in between timestamp 00:30 to 00:45 then write as following in script above (shown in red font below).

Example, shot 2.1.2 SCREEN: performing ........... action.   xxx.mp4 00:30-00:45.

2.2.2 SCREEN: clicking on.................button... xxx.mp4 01:00-01:25

Please note that every shot (3-digit step) should have the best corresponding video clip of only 20 to 25 seconds and not more than that, so that the voice narration can match the duration.

If a particular shot takes minutes to complete, we can choose timestamps at the beginning and end of the process so that the total time is again within the 25-second limit.

Example: 3.1.1 SCREEN: Clicking on the 'analyze and plot' button, analysis in progress and the graph being plotted.  xxx.mp4 00:00-00:10 and 02:30-02:40