TITLE:

Surgical Technique for the Implantation of a Wireless Telemetry Device on Pigs’ Backs

AUTHORS AND AFFILIATIONS:

Angel Moctezuma-Ramirez, Jordana Goulet, Abdelmotagaly Elgalad

Center for Preclinical Surgical & Interventional Research, The Texas Heart Institute, Houston, Texas, USA

Email addresses of co-authors:

Angel Moctezuma-Ramirez		(amoctezuma@texasheart.org)

Jordana Goulet			(goulet@livemail.uthscsa.edu)

Corresponding author:

Abdelmotagaly Elgalad		(aelgalad@texasheart.org)

SUMMARY:

A new surgical technique was developed for implanting telemetry devices in pigs that allows continuous monitoring without discomfort. The device is placed on the back, minimizing irritation and infection risk.

ABSTRACT:

Telemetry devices are essential for preclinical cardiovascular research, allowing for continuous monitoring of physiological parameters like electrocardiogram, temperature, and respiration in experimental animals without causing disruption. Traditional methods, such as intra-abdominal placement, can cause discomfort, particularly in pigs due to their habit of lying prone. To address this issue, a novel surgical technique was developed for implanting wireless telemetry devices in pigs, aimed at improving animal comfort and maintaining signal quality. This study involved 40 Yorkshire pigs (20-30 kg). The telemetry device was implanted in the back. A 2 inch incision was made lateral to the spine, followed by blunt dissection to create a pocket under the latissimus dorsi fascia. The pocket was treated with a vancomycin solution to prevent infection before device insertion. The leads were secured, and the site was closed in layers. All pigs showed strong, continuous telemetry signals after surgery. Over the 180-day follow-up period, only two cases of infection occurred: One was managed with antibiotics, and the other required device removal. No significant bleeding or loss of telemetry signal was observed, and all animals survived to the end of the study. This new technique offers a reliable method for implanting telemetry devices in pigs, reducing discomfort and infection risk while providing high-quality physiological data. It holds promise for broader use in preclinical cardiovascular research, improving both animal welfare and data reliability.

INTRODUCTION:

Telemetry devices play a pivotal role in advancing the understanding of cardiovascular physiology, particularly in studies using an animal model of myocardial infarction. By enabling the continuous monitoring of physiological signals, including vital parameters and electrocardiogram (ECG) readings, these devices offer invaluable insights into the dynamic changes occurring during cardiac events. In preclinical studies, where animal models like pigs are instrumental in mimicking human cardiovascular responses, telemetry devices have become indispensable tools1. The ability to record ECG continuously in unrestrained, freely moving animals provides a nuanced perspective on the impact of myocardial infarction, allowing researchers to capture real-time alterations in cardiac function.

In studies involving animal models of myocardial infarction, implantable telemetry devices offer distinct advantages. They facilitate the measurement of critical parameters such as heart rate with minimal disruption to the animals’ natural behavior, eliminating potential confounding factors introduced by external devices2.

These devices often come with a manual with instructions for their surgical implantation, but different studies and animal models may require implantation techniques different from those in the manual. Traditional telemetry implantation sites, such as intra-abdominal or neck placement, present drawbacks in porcine models used in extended cardiovascular studies, including potential discomfort, interference with femoral access, restricted movement, and increased infection risk1-3. Therefore, an alternative approach was needed that was compatible with long-term monitoring in unrestrained pigs.

Compared to traditional intra-abdominal placement, which can cause discomfort, or neck placement, which has had lower success rates in pigs in some other studies, this dorsal approach aims to improve welfare while maintaining data quality.

The primary goal of this study was to develop and validate a novel surgical technique for implanting wireless telemetry devices in the backs of pigs to ensure animal comfort and reliability. The aim was to harness the benefits of continuous ECG monitoring over an extended period (up to 180 days). This technique was developed in response to the limitations observed with traditional implantation methods (e.g., intra-abdominal, inguinal, and neck placements), which in pigs are associated with substantial risk of discomfort, limited mobility, interference with vascular access, and infection. We developed this novel dorsal approach to enhance animal welfare, reduce postoperative complications, and maintain reliable signal quality during long-term monitoring in unrestrained animals.

PROTOCOL:

All animal procedures were approved by The Texas Heart Institute’s Institutional Animal Care and Use Committee (IACUC). Animals were implanted with the EMKA easyTEL+L-EEEETA digital system (https://www.emkatech.com/product/easytel-large-animals/).

- Preoperative preparations for experimental animals

  - Use Yorkshire pigs of either sex weighing between 20 kg and 30 kg. Do not feed the pigs 12-24 h before the start of anesthesia but allow access to water.

- Sedation and anesthesia

  - Administer Telazol (tiletamine and zolazepam; 2-6 mg/kg intramuscular [IM]) and atropine sulfate (0.02-0.05 mg/kg IM). Allow 10-15 min for the drugs to induce sedation, then transport the pig to the preparation room. Confirm adequate anesthetization by the absence of jaw tone and lack of response to toe pinch. Apply an ophthalmic ointment to both eyes to prevent corneal drying during anesthesia.

  - Place a face mask to administer oxygen and isoflurane at 3% until adequate jaw tone relaxation permits endotracheal intubation.

  - Intubate the pig with direct vision of the vocal folds. Connect the pig to a volume ventilator and maintain anesthesia by isoflurane inhalation (0.5%-3.0%) in O2 (40%-100%).

  - Place an intravenous (IV) catheter in an ear vein. Shave the area over the lateral auricular vein and disinfect it with 70% isopropyl alcohol followed by povidone-iodine solution. Insert a 20- to 22G catheter into the vein, using aseptic technique, then secure the catheter with medical tape and flush it with sterile saline to confirm patency.

  - Administer buprenorphine (Buprenex; 0.005-0.015 mg/kg IM), flunixin meglumine (1.1-2.2 mg/kg IV or IM), and ceftiofur (Naxcel; 3-5 mg/kg IM).

  - Shave the chest, upper back, lower back, and the distal portions of all four limbs, specifically the areas over the anterior aspect of the hooves, using electric clippers to fully expose the skin for electrode placement. Clean all shaved areas with sterile gauze soaked in 70% isopropyl alcohol, wiping in a circular motion from the center outward to remove debris and oils in preparation for sterile draping. Place adhesive ECG electrodes on all four limbs, ensuring firm skin contact to minimize signal artifacts.

- Surgical procedure

  - In the operating room, place the pig on the table in the ventral position. Maintain anesthesia with isoflurane inhalation (0.5%-3.0%). Start IV infusion of plasmalyte at 10-15 mL/h.

  - Surgical site preparation

    - Clean the intended surgical incision site on the left upper back and the surrounding shaved area (approximately 15 cm x 15 cm) with sterile gauze soaked in 70% isopropyl alcohol. Wipe from the center outward until no visible dirt or debris remains on the gauze. Allow the alcohol to air-dry completely.

    - Perform a sterile surgical scrub using povidone-iodine scrub solution and sterile gauze. Apply the scrub using firm, circular motions, starting at the incision line and moving outward in expanding circles. Continue scrubbing for at least 5 min, using fresh gauze as needed.

    - Apply povidone-iodine solution with sterile gauze or sponge sticks, again moving from the center outward. Allow this final paint layer to air-dry completely; do not wipe it off.

    - Isolate the prepared surgical area with sterile surgical drapes.

  - Identify the anatomical landmarks for the incision. These landmarks form a triangle comprising the spine (upper border), the scapula (left border), and the lower border of the trapezius muscle (right border; Figure 1).

  - Using a number 10 blade, make a 2 inch longitudinal incision approximately 2 inches laterally from the spine.

  - After the epidermis has been incised, use an electrocautery to advance through the subcutaneous tissue and then divide the trapezius muscle. Take special care to divide all of the fascial layers.

  - Once the inner fascial layer has been divided, use blunt dissection with the index and middle finger to create a pocket in this layer. Check that the tip of the scapula is within reach of the fingertips. After dissection, the tip of the scapula will be palpable at the extent of the pocket (Figure 2).

  - Check that the telemetry device can fit in the pocket with minimum tension on the skin incision. Remove the device from the pocket.

  - Mix vancomycin (1 g) with 50 mL of 0.9% saline and place approximately 10 mL of the mixture in the pocket for 1-2 min. Remove excess solution with a gauze pad.

  - Use a 14 Fr angiocath to place wires as shown in Figure 2.

    - Insert angiocath, going lateral to medial, outer to inner orientation. Direct the needle under the muscle toward the pocket. Advance the needle until only a third of its body remains visible. Remove the needle and thread the lead through the catheter, going from the tip of the catheter to the catheter hub.

    - Leave between 1 and 2 inches of lead exposed. Sufficiently exposed lead prevents it from being lost in the subcutaneous tract during placement of other leads. Remove the catheter.

    - Insert one colored lead (yellow, red, green, or blue), followed immediately by its corresponding contralateral transparent lead on the opposite side of the pocket. Insert each pair together, one pair at a time, to match their arrangement on the telemetry device and prevent confusion during placement. The colored and transparent leads form a differential pair for each ECG channel, enabling bipolar signal acquisition and reducing electrical noise. Repeat this process sequentially for all lead pairs. Insert the black ground lead last. This approach ensures consistent lead orientation and accurate signal transmission (Figure 3).

    - Once all leads are placed, pull the wires so the tips of the leads are barely under the skin. Ensure that no leads are visible at the skin level (Figure 3).

  - Collect all wires into a loop, using the hand as a guide, and tie them into a bundle with heavy silks.

  - Place the device and lead bundle in the pocket. Ensure that the device sits flat and is oriented correctly. One side of the device is labeled with the instruction: This side facing out.

  - Position the device within the pocket so that its lateral edge lies at least 1-2 cm medial to the palpable caudal edge of the scapula. Before proceeding, gently manipulate the pig's left forelimb through its range of motion to confirm that the device does not significantly impinge on or rub against the scapula.

  - Make sure device readings are transmitted to the receiver and computer (Figure 4).

NOTE: Successful transmission is defined as a transmission that accurately measures heart rate, assesses rhythm, and identifies QRS complexes. The raw signal may not be perfectly isoelectric or artifact-free.

  - Ensure hemostasis. Close the muscle layer with 1-0 Vicryl. Apply the vancomycin mixture on top of the muscle layer. Allow 1-2 min for absorption.

  - Close the subcutaneous layer with 2-0 Vicryl. Staple the incision and apply tissue glue.

- Post -operative care

  - Monitor and record vital signs—heart rate, respiratory rate, and temperature—hourly, beginning at the time of intensive care unit admission. Continuously assess the pig until extubation criteria are met, including spontaneous respiratory effort while on the ventilator, presence of a chewing reflex, and the ability to maintain sternal recumbency. Extubation should be based on these clinical indicators and supported by arterial blood gas measurements.

  - Continue hourly monitoring after extubation; reduce to every 4 h once the pig demonstrates physiological stability. Ensure the pig is never left unattended during recovery. A trained veterinary technician must remain present at all times throughout the recovery period.

  - House the pigs individually in adjacent pens that permit visual but not physical contact with conspecifics, and do not return the pig to group housing until it has fully recovered.

  - Administer postoperative analgesia, which include buprenorphine (0.05–0.1 mg/kg IM or PO) every 6–12 h or as directed by the attending veterinarian, and flunixin meglumine (1.1–2.2 mg/kg IV, IM, or PO) administered every 6–8 h as needed.

  - Administer antimicrobial prophylaxis beginning with Naxcel (ceftiofur; 3–5 mg/kg IM) once daily for 24 h, starting the day after surgery, followed by Baytril (enrofloxacin; 2.5–5.0 mg/kg PO, BID) for up to 14 days.

  - Maintain aseptic postoperative care by applying iodine spray to the incision site every 8 h until complete wound healing and staple removal at postoperative day 14. Apply triple antibiotic ointment to the incision site as needed.

RESULTS:

The data presented in Table 1 show the reliability and safety of the surgical technique for dorsal implantation of telemetry devices in pigs. Successful ECG transmission was recorded in all 40 pigs (100%), confirming that the positioning of the device and lead configuration consistently enabled effective long-term signal acquisition. The absence of major bleeding events (0%; defined as ≥10 mL of blood loss) indicates the safety of the fascial pocket dissection and careful handling of vascular structures. Infection occurred in 2 pigs (5%) early in the study and was managed with systemic antibiotics; protocol modifications, including intraoperative vancomycin irrigation and enhanced postoperative antiseptic care, were implemented thereafter. All animals survived to the 180-day endpoint, underscoring the long-term viability of the implantation method.

Table 2 compares the current results to previously published telemetry implantation strategies. Dorsal implantation achieved equivalent or superior signal reliability compared with implantation at other anatomical sites, such as the neck or abdomen, which may be less suitable in pigs due to behavioral tendencies, interference with femoral access, or higher complication rates. In contrast to neck implantation in pigs, which showed a 27.3% complication rate and only 72.7% signal success, the current method yielded complete signal acquisition with minimal complications, indicating its suitability for long-term monitoring studies in unrestrained large animal models.

Figure 4 shows representative ECG tracings recorded from a pig on postoperative day 30. The tracings display well-defined P waves, consistent QRS complexes, and minimal baseline drift, reflecting stable lead placement and signal fidelity. Four channels are shown, corresponding to paired bipolar leads placed as described in the protocol. The lower two rows in the figure show concurrent temperature readings and signal quality metrics from the device. These results indicate that the described implantation method provides high-quality ECG waveforms with sufficient resolution to analyze heart rate variability, rhythm, and conduction morphology throughout long-term follow-up.

FIGURE AND TABLE LEGENDS:

Figure 1: Anatomical landmarks for surgical incision site. This photograph shows the shaved and prepped upper back of the pig with the surgical site draped. The incision site is defined by anatomical landmarks forming a triangle: the spine (upper border), the scapula (left border), and the lower border of the trapezius muscle (right border). These landmarks are used to determine the optimal pocket location for telemetry device implantation.

Figure 2: Formation of the subfascial pocket and sequential insertion of telemetry leads. Top: Intraoperative photograph showing blunt dissection beneath the latissimus dorsi muscle to form a subfascial pocket for telemetry device implantation. Dissection is performed using the index and middle fingers, directed toward the tip of the scapula to ensure sufficient depth and coverage for the device and lead bundle. Bottom: Schematic representation of the lead insertion configuration. Colored leads (red, yellow, green, and blue) and their corresponding contralateral transparent leads are shown routed across the pocket in matched pairs, consistent with the telemetry device's terminal layout. The black lead represents the ground. Pairing leads in this sequence maintains bipolar alignment and minimizes confusion during insertion and signal acquisition.

Figure 3: Final lead positioning and device placement before surgical closure. Top: Final positioning of telemetry leads after subcutaneous routing. Leads are pulled until the tips rest just beneath the skin surface, ensuring reliable signal capture without external exposure or tension. All leads are grouped and secured with heavy silk ties in preparation for placement into the fascial pocket. Bottom: The telemetry device and secured lead bundle are seated inside the dissected subfascial pocket. The device is oriented with the labeled surface facing outward. The lateral edge of the device is positioned at least 1–2 cm medial to the caudal margin of the scapula, minimizing the risk of mechanical irritation during limb movement. Proper positioning is confirmed intraoperatively through palpation and passive limb mobilization.

Figure 4: Representative ECG tracings and telemetry signal data. The 10 s segment of continuous telemetry recording was made on postoperative day 30. From top to bottom: ECG signals from leads 1–4, core temperature, and signal quality. ECG traces show well-defined P waves, consistent QRS complexes, and minimal baseline drift, indicating stable lead positioning and high-fidelity signal acquisition. Temperature remains within physiologic range, and signal quality remains at 100%, reflecting uninterrupted telemetry transmission. The x-axis represents time in seconds (0–10 s), enabling temporal correlation across channels. The y-axis varies by signal type: ECG is shown in mV, temperature in °C, and signal quality as a binary scale (0 = no signal, 100 = optimal signal).

Table 1: Telemetry implantation outcomes.

Table 2: Comparison of telemetry implantation techniques.

DISCUSSION:

A device-implantation procedure is described as a new invasive telemetric method of continuous ECG monitoring. This method ensures reliable recordings in animal experiments. Although the user manual for the EMKA device4 describes a surgical procedure for implanting the device in the abdomen of large animals such as dogs and pigs, intra-abdominal placement was not considered desirable because pigs typically lie on their abdomen, resulting in pressure on the peritoneum that could cause pain or discomfort to the animal5. Inguinal placement was also not desirable because this study required unrestrained, freely moving pigs, so it was important that the device not limit movement or cause excessive irritation. Also, this study required percutaneous access to the femoral arteries, and inguinal placement of the device would have interfered with that. Implanting the device in the intramuscular layers of the neck, another common approach3, was also ruled out; the investigators’ previous experience suggests that pigs will scratch surgical areas they have access to, which increases their risk of infection.

Because these established surgical procedures would not work for the present study, a new protocol was created that placed the device on the pigs’ backs, between the scapula and the spine. This way, the pigs were free to move unrestrained without irritation, and the implanted device did not interfere with any of the activities the study required. This method of telemetric implantation will be useful in studies that require continuous physiological signals in unrestrained pigs, such as studies that induce myocardial infarction or arrhythmias. This approach allowed uninterrupted ECG telemetry in 100% of the pigs over 180 days. No signs of device-related irritation were observed. Behaviors commonly associated with discomfort, such as scratching or rubbing the surgical area, were not seen by veterinary staff. Postoperative care records did not note any concerning behavior. The infection rate was low (5%), supporting the suitability of this implantation site for long-term use.

Critical steps in this protocol include device positioning, application of antibiotics during the procedure, proper insertion of the angiocath, and confirmation of signaling. It is important to identify the anatomical landmarks and place the device in a location that will not irritate the scapula when the animal is moving and that is far enough from the spine to avoid injuring it. It is also important to make the pocket only as large as it needs to be to accommodate the device and wire bundle; this will limit infection risk and shorten recovery time. Another critical step is placing the leads with the angiocath as shown in Figure 3. It is crucial that the angiocath travels below the muscle layer but not low enough to make contact with the pig’s spine, which could cause paralysis. The last critical step before the pocket is closed is to confirm successful transmission from the device to the receiver.

A brief intraoperative signal verification was performed in all cases to assess the accuracy of the data from the implanted telemetry device. Specifically, heart rate values displayed by the device were compared to those recorded by the single-lead ECG monitor used during anesthesia. This check was performed before final surgical closure to confirm that the implanted leads were detecting physiological activity in agreement with real-time external monitoring. Postoperative validation of telemetry performance was based on continuous assessment of ECG waveform characteristics, including clearly defined P waves, consistent QRS complexes, stable amplitude, and absence of baseline drift. Additionally, telemetry-derived heart rate readings were routinely compared with manually measured heart rates during postoperative care to ensure ongoing signal reliability. No formal comparison with a standard 12-lead ECG was made, as the study protocol did not include sedation or re-instrumentation during follow-up. Nevertheless, signal quality was consistent with the fidelity benchmarks proposed by Krause et al.1, who defined valid telemetry as ECG tracings with clearly identifiable P waves, QRS complexes, and T waves; stable baseline; minimal artifact; and physiologically coherent heart rate output. These parameters are met in the representative 10 s telemetry tracing presented in Figure 4, where all waveform components are visible, the baseline is stable, and signal quality is uninterrupted.

This technique offers benefits over alternative research methods. Primarily, it prioritizes animal welfare by positioning the device on the animal's back, significantly reducing discomfort and allowing for natural movement, which is a substantial improvement over traditional intra-abdominal placement. Furthermore, the method consistently yields continuous physiological data, maintaining the level of reliability found in other established techniques1,6-10. Finally, the technique has a low rate of complications, indicating its low risk, and allows for effective data collection.

Table 2 compares this technique with other methods of telemetry implantation. Krause et al.1 implanted telemetry devices in pigs' necks, whereas Langin et al.6 implanted devices in the chest cavity of baboons. Markert et al.7 varied the implantation site depending on the animal model, placing devices between the external and internal abdominal oblique muscles along the left lateral flank in dogs and minipigs and in the abdominal wall in primates. Prior et al.8 and Ruppert et al.9 used intra-abdominal implantation sites in dogs and guinea pigs, respectively. Shiotani et al.10 implanted devices in the flank of guinea pigs.

The studies also varied in their focus and objectives. Like the current study, some aimed to develop novel surgical techniques to improve animal comfort and maintain signal quality. Markert et al.7 focused on developing a new telemetry system with adjustable catheters and true core temperature measurement. Langin et al.6 explored the use of telemetry in heterotopic thoracic cardiac xenotransplantation, a complex surgical scenario involving two independently working hearts.

All pigs that underwent this procedure had continuous signaling during the study. Continuous telemetry recordings were reviewed daily by study personnel to ensure signal fidelity and to monitor potential device failure. Each implant transmitted data in 6 h recording blocks, stored locally using the software, allowing retrospective analysis. Full ECG waveforms were visually inspected at random intervals throughout the study to confirm persistence of cardiac electrical activity. More importantly, no implant-related signal loss occurred in any of the 40 pigs. Occasional transient signal interruptions were associated with logistical issues, such as temporary disconnection of the receiver or pigs being moved beyond the receiver’s optimal transmission range (~10 feet), but these interruptions were infrequent and were promptly resolved. Despite adherence to the initial standard aseptic preparation protocol, surgical site infections occurred in the first two pigs (out of 40). These infections required intervention and highlighted the susceptibility of this model to such infections after this type of long-term device implantation, a risk consistent with prior institutional experience with porcine surgical models. Consequently, the protocol was modified for the remaining 38 pigs to enhance infection prevention. These modifications included enhanced pre-scrub cleaning with alcohol and, critically, two intraoperative applications of a vancomycin solution (1 g in 50 mL 0.9% saline) locally within the surgical pocket and over the muscle layer. This change was further motivated by the logistical requirement to transport these animals to a different facility days after the procedure; during this transport, stress could potentially increase susceptibility to infection, making robust prophylaxis essential.

The critical importance of antibiotic stewardship and the concerns regarding the prophylactic use of vancomycin were not discounted. Its inclusion was a corrective action prompted by the initial 5% infection rate and justified by the specific high-risk context: the known susceptibility of the model, the presence of a large subcutaneous implant, the long duration of the study, and the added stress of postoperative transport.

After these modifications were made, no infections occurred in the remaining 38 pigs over the 180-day study period, showing this comprehensive preventative strategy's empirical necessity and success in the specific circumstances of this study. This approach, developed under IACUC oversight, struck a necessary balance to ensure both animal welfare and the integrity of the study data. Researchers adopting this protocol should always prioritize meticulous aseptic technique and consider these context-specific risk factors, institutional guidelines, and local resistance patterns when determining appropriate antibiotic prophylaxis.

The protocol was also modified to describe how the excess length of the wires was managed. Initially, each wire was cut to the precise length from the device to the outer opening; however, this took a long time, and it was found to be faster and easier to bundle the excess wires and place them in the pocket. Neither method interfered with the transmission of the signal.

Overall, these results suggest that the surgical technique described in this manuscript is reliable and effective for implanting telemetry devices in pigs. The high rate of successful ECG transmission, the low complication rate, and the excellent survival rate support its potential for broader use in preclinical cardiovascular research.

ACKNOWLEDGMENTS:

Stephen N. Palmer, PhD, ELS, of the Department of Scientific Publications at The Texas Heart Institute, contributed to the editing of the manuscript.

DISCLOSURES:

None declared.

REFERENCES:

1	Krause, A. et al. Surgical implantation and functional assessment of an invasive telemetric system to measure autonomic responses in domestic pigs. Vet J. 207, 140-146 (2016).

2	Horning, M. et al. Best practice recommendations for the use of fully implanted telemetry devices in pinnipeds. Animal Biotele. 5 (1), 13 (2017).

3	Swindle, M., Smith, A. C., Helke, K. Recommendations for medical device implantation in swine. Israel J Vet Med. 68, 3-11 (2013).

4	National Centre for the Replacement, Refinement and Reduction of Animals in Research. Implanted telemetry: A valuable tool for preclinical research https://www.emkatech.com/implanted-telemetry-in-preclinical-research/ (2023).

5	Morton, D. B. et al. Refinements in telemetry procedures. Seventh report of the bvaawf/frame/rspca/ufaw joint working group on refinement, part a. Lab Animals. 37 (4), 261-299 (2003).

6	Langin, M. et al. Perioperative telemetric monitoring in pig-to-baboon heterotopic thoracic cardiac xenotransplantation. Ann Transplantat. 23, 491-499 (2018).

7	Markert, M. et al. A new telemetry-based system for assessing cardiovascular function in group-housed large animals. Taking the 3rs to a new level with the evaluation of remote measurement via cloud data transmission. J Pharmacol Toxicol Meth. 93, 90-97 (2018).

8	Prior, H., Mcmahon, N., Schofield, J., Valentin, J. P. Non-invasive telemetric electrocardiogram assessment in conscious beagle dogs. J Pharmacol Toxicol Meth. 60 (2), 167-173 (2009).

9	Ruppert, S., Vormberge, T., Igl, B. W., Hoffmann, M. Ecg telemetry in conscious guinea pigs. J Pharmacol Toxicol Meth. 81, 88-98 (2016).

10	Shiotani, M., Harada, T., Abe, J., Hamada, Y., Horii, I. Methodological validation of an existing telemetry system for qt evaluation in conscious guinea pigs. J Pharmacold Toxicol Meth. 55 (1), 27-34 (2007).