# Submission ID #: 67871

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=20700703

Title: Evaluating Tests of Cognition Using a Computerized Touch-Sensitive Tablet, Eye Tracking, and Functional Magnetic Resonance Imaging

# Authors and Affiliations:

Alexandra Pavel1*, Francis A. Fernandes1*, Sean Rose1, C. Fung1, Serenna Gerhard1, Fred Tam1, Corinne E. Fischer2,3, Nathan W. Churchill2, Tom. A. Schweizer2,4, Simon J. Graham1,5

1Physical Sciences Platform, Sunnybrook Research Institute, Sunnybrook Health Sciences Centre

2Keenan Research Centre for Biomedical Science, the Li Ka Shing Knowledge Institute, St. Michael’s Hospital

3Department of Psychiatry, Faculty of Medicine, University of Toronto

4Division of Neurosurgery, Faculty of Medicine, University of Toronto

5Department of Medical Biophysics, Temerty Faculty of Medicine, University of Toronto

*These authors contributed equally

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Simon J. Graham				Simon.Graham@utoronto.ca

# Email Addresses for All Authors:

Alexandra Pavel				alexandrapavel863@gmail.com

Francis A. Fernandes				F.Fernandes@mail.utoronto.ca

Sean Rose					sean11kr@hotmail.com

C. Fung						c29fung@uwaterloo.ca

Serenna Gerhard				Serenna.Gerhard@sri.utoronto.ca

Fred Tam					Fred.Tam@sri.utoronto.ca

Corinne E. Fischer				Corinne.Fischer@unityhealth.to

Nathan W. Churchill				Nathan.Churchill@unityhealth.to

Tom A. Schweizer				Tom.Schweizer@unityhealth.to

Simon J. Graham				Simon.Graham@utoronto.ca

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible. https://review.jove.com/files_upload.php?src=20700703

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  22

Number of Shots:  52

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research questions will your laboratory focus on in the future?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

Answers to these questions will not appear in the video but may be featured in our journal's promotional materials.

- Enter the full name of the author who will deliver the statement.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Answers will be mildly edited for clarity.

- Limit the length of each statement to 50 words or fewer.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 50 or fewer words.

Ethics Title Card

This research has been approved by the Research Ethics Board (REB) at the Sunnybrook Health Sciences Centre

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Experimental Setup for Recording Visuomotor Behavior and Brain Activity

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, secure the tablet to its frame and attach the MRI-compatible video camera [1]. Apply fresh blue tape across the tablet surface, ensuring the entire touch area is covered without major creases [2]. Remove any excess tape from the tablet edges [3].

    - WIDE: Talent attaching the tablet firmly to its supporting frame and connecting the MRI-compatible video camera.

    - Talent applying fresh blue tape smoothly across the tablet’s surface, carefully pressing down to avoid creases.

    - Talent trimming the tape around the tablet edges.

  - Bring the tablet, stylus, tablet link cable, and tablet video camera link cables into the magnet room [1]. Connect the tablet link and video camera link cables to the magnet-room side of the RF penetration panel [2]. Then, slide the MRI-compatible tablet clips into the patient table rails, securing two clips on each side to fasten the tablet system in place [3].

    - Talent carrying the tablet, stylus, and cables into the magnet room.

    - Talent connecting the tablet link and camera cables to the RF penetration panel.

    - Talent inserting tablet clips into both rails of the patient table and securing them.

  - Next, place the MRI-compatible projector behind the back end of the magnet, approximately 1 meter away from the magnet bore [1]. Mount the MRI-compatible rear projection screen inside the magnet bore, approximately 2 meters from the projector [2].

    - Talent placing the projector on a stable surface 1 meter behind the magnet bore.

    - Talent positioning and securing the projection screen inside the magnet bore at the correct distance from the projector.

  - Place the MRI-compatible eye-tracking camera between the projector screen and the outer edge of the magnet bore, ensuring the camera mount is flush with the bore edge [1]. Tighten the plastic screws on the mount to secure the eye-tracking system in place [2].

    - Talent positioning the eye-tracking camera inside the magnet bore between the screen and the bore edge.

    - Talent adjusting and tightening plastic screws on the camera mount to secure it firmly to the bore.

  - Now, prepare the patient table by installing the 64-channel head coil [1]. Instruct the participant to lie supine on the table and position their head fully into the coil [2]. Place padding around the participant’s head to restrict movement and ensure a secure fit [3]. Use the landmark laser to confirm that the head is centered within the head coil [4].

    - Talent placing the 64-channel head coil onto the MRI patient table.

    - Talent guiding the participant to lie supine and adjusting their head into the head coil.

    - Talent placing padding on both sides of the participant’s head.

    - Talent using the laser tool to verify head alignment within the coil.

  - Adjust the position of the head coil mirror until the participant can see the rear projection screen clearly and without obstruction [1].

    - Talent adjusting the mirror attached to the head coil while communicating with the participant to confirm visibility of the screen.

  - Then, place the tablet mount across the participant’s waist so that the touch-sensitive surface rests comfortably to support writing and drawing movements [1].

    - Talent positions the tablet mount across the participant’s waist and adjusts its angle for ergonomic access.

  - Now, place the tablet stylus in the participant’s dominant hand and instruct them to hold it like a pen [1]. Ask the participant to touch all four corners of the tablet surface with the stylus to assess range and comfort [2]. Adjust the mount position and add padding under the elbow if needed to improve comfort or access [3].

    - Talent placing the stylus into the participant’s hand and demonstrating how to hold it.

    - Talent observing as the participant touches all corners of the tablet.

    - Talent adjusting tablet mount and adding padding under the elbow to optimize comfort.

  - Using Velcro straps, fasten the tablet system to the patient bed [1]. Slowly slide both the participant and tablet system into the magnet bore, ensuring the tablet does not contact the bore edge and the cables remain untangled [2].

    - Talent tightly securing the tablet system to the patient table using Velcro straps.

    - Talent carefully slides the participant into the magnet bore while monitoring tablet alignment and cable management.

  - On the video camera computer, launch the Video camera.exe (Camera-E-X-E) program [1]. Then, create a new screen capture session for the participant [2].

    - SCREEN: Show the desktop screen as the user double-clicks the Video camera.exe icon to launch the program.
Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/files_upload.php?src=20700703

    - SCREEN: A New screen capture session is being created.

  - Following the EyeLink (Eye-Link) 100 Plus User Manual recommendations, configure pupil and corneal reflection thresholds, and calibrate and validate the eye-tracking camera [1].

    - SCREEN: Pupil and corneal reflection thresholds are being configured, and the eye tracking camera us being calibrated and validated

  - To adjust the eye-tracking camera to focus on the participant’s right eye, switch between camera views, fine-tune the lens, and modify the illuminator until the view is optimized [1]. Once the pupil threshold and corneal reflection values are acceptable, document the values and initiate the 9-point calibration [2].

    - SCREEN: Show the camera interface as the user toggles camera angles, adjusts the lens focus, and changes illumination settings to optimize the right eye view.

    - SCREEN: Display the software screen as the threshold and CR values are confirmed and recorded. The 9-point calibration is being started.

  - Then, validate the calibration [1]. Record both the average and maximum validation angle values before proceeding to the fMRI experiment [2-TXT].

    - SCREEN: Show the calibration interface to validate the calibration.

    - SCREEN: Display the readout panel showing average and maximum angle values being recorded. TXT: Repeat calibration until avg. error ≤ 0.5° & max error ≤ 1.0

  - Next, use the stimulus or response computer to begin calibration of the tablet touch surface [1]. Instruct the participant to use the stylus to consecutively touch and release the three targets that appear on the screen within the allotted time limits [2].

    - SCREEN: Show the user opening the calibration program on the stimulus/response computer.

    - SCREEN: Show the participant using the stylus to tap and release each one within the visible time markers.

  - Once calibration is complete, launch the referenced graphics editing application [1]. Instruct the participant to write their name to confirm accurate stylus tracking [2].

    - SCREEN: Show the desktop interface as the user opens the graphics editing application.

    - SCREEN: Show the participant writing their name with the stylus on the tablet.

  - Now, familiarize the participant with tablet-based writing by asking them to complete a self-paced training task used in essential tremor studies [1].

    - SCREEN: Show the training interface with prompts guiding the participant to sign their name.

  - To familiarize the participant with the TMT (T-M-T), guide them through a simplified training version of TMT-A and TMT-B, each with 12 items [1]. Then, proceed to alternate full-sized versions of TMT-A and TMT-B with rearranged items, using experimental task timing [2]. Monitor the participant’s performance throughout to ensure the tablet remains properly calibrated and the task is followed accurately [3].

    - SCREEN: Show the training version of TMT-A and TMT-B with 12-item sequences and the participant performing the task with a stylus.

    - SCREEN: Display the alternate full-length TMT-A and TMT-B sequences being completed by the participant.

    - SCREEN: Show the operator observing the stylus trace and checking for alignment or calibration issues.

  - Now, begin the eye-tracking session by selecting Start Recording in the Screen Recorder program on the video camera computer [1]. Then, on the stimulus or response computer, open the TMT-Run1_slow.ebs2 (T-M-T-Run-One-Slow-E-B-S-two) E-Prime script file using E-Run [2]. Input the participant ID and the session number when prompted by the E-Run script [3].

    - SCREEN: Show the user clicking Start Recording in the SR Research Screen Recorder software interface.

    - SCREEN: Show the stimulus/response computer as the user navigates to and opens TMT-Run1_slow.ebs2 in the E-Run application.

    - SCREEN: Participant ID and session number are being entered

  - During the run, assign one lab member to monitor the eye-tracking data to ensure a stable signal [1]. Simultaneously, have a second lab member observe the participant’s performance on the TMT task, ensuring that instructions are being followed and that no technical issues are present, such as unreliable projection or stylus tracking [2]. Ask the second lab member to also document any errors in TMT-A or TMT-B performance, along with the associated trial number [3].

    - Show a lab member reviewing the real-time eye-tracking data graph for signal stability on the recording computer.

    - Show another lab member watching the participant's stylus trace as it appears on-screen, noting any delays or mismatches.

    - SCREEN: Show a log being filled in with any detected TMT-A or TMT-B performance errors, along with trial numbers.

  - After the run is complete, stop the eye-tracker recording [1]. In the SR Research Screen Recorder software, perform a single-point drift check by selecting Drift Correct [2-TXT].

    - SCREEN: Show the user clicking to stop the recording in the Screen Recorder program.

    - SCREEN: Show the drift-check process being initiated by selecting Drift Correct, with the result displayed on screen. TXT: If error ≥ 2.0°, repeat 9-pt calibration until avg. < 1.0° & max < 2.0°

  - For Run 2, restart the eye-tracking recording session on the video camera computer [1]. On the stimulus or response computer, open the TMT-Run2_slow.ebs2 E-Run script file [2]. Enter the same participant ID and session number used in Run 1. Repeat the task instructions [3].

    - SCREEN: Show the user restarting the eye-tracking session by selecting Start Recording in the Screen Recorder software.

    - SCREEN: Display the E-Run interface as TMT-Run2_slow.ebs2 is opened.

    - SCREEN: Show the participant ID and session number being re-entered into the software.

  - Once the experiment concludes, perform one final eye-tracking validation to confirm calibration accuracy [1]. Then, click File followed by Close on the eye-tracking software to finalize and export the data [2]. Remove the participant from the magnet bore and begin equipment take-down [3]. At last, shut off all computers and equipment [4].

    - SCREEN: Show the the final calibration check.

    - SCREEN: Display the software interface as the user selects File > Close, triggering data export.

    - WIDE: Talent carefully guiding the participant out of the magnet

    - Talent removing the tablet system and related equipment.

    - Talent shutting off the computer and equipment

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 201.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - During the initial seconds of both TMT-A and TMT-B, the participant’s gaze searched for targets before initiating stylus movement [1]. Across all time intervals shown, gaze consistently preceded the stylus linking motion in both TMT-A and TMT-B [2].

    - LAB MEDIA: Figure 7C and D. Video editor: Highlight the red gaze path in the first panel (0s - 2.5s) before any blue lines appear.

    - LAB MEDIA: Figure 7C and D. Video editor: In each of the four time segments, highlight the red path leading slightly ahead of the blue line.

  - Completion time trended longer for TMT-B than for TMT-A [1]. Across all time intervals shown, gaze consistently preceded the stylus linking motion in both TMT-A and TMT-B [2]. Stylus pause latency showed a trend of being higher for TMT-B than for TMT-A [3].

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Completion time” and compare the values for TMT-A  and TMT-B.

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Speed” and show the similar values for TMT-A and TMT-B.

    - LAB MEDIA: Table 1. Video editor: Highlight the row “SPL” and show the higher average value for TMT-B compared to TMT-A.

  - No significant differences were found between TMT-A and TMT-B in linking period durations, non-linking period durations, total distance, extra distance travelled, or distance per link [1].

    - LAB MEDIA: Table 1. Video editor: Highlight the rows “Linking period”, “Non-linking period”, “D”, “EDT”, and “DPL”, showing similar values for both TMT-A and TMT-B.

  - Stylus force trended slightly higher during TMT-B compared to TMT-A [1], and blink count per trial was significantly greater during TMT-B [2].

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Stylus Force” and show the higher value for TMT-B compared to TMT-A.

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Blink Count” and emphasize the difference between TMT-B and TMT-A.

  - Blink rate remained significantly higher in TMT-B even after accounting for completion time differences [1]. The percentage of time spent in fixation was significantly lower during TMT-B than during TMT-A [2].

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Blink Rate” and show the value for TMT-B compared to TMT-A.

    - LAB MEDIA: Table 1. Video editor: Highlight the row “Fixation %” and compare the lower value for TMT-B with TMT-A.

  - Functional MRI revealed widespread positive brain activation during both TMT-A and TMT-B compared to visual fixation [1]. Brain activation included regions such as the cerebellum, parietal lobules, occipital gyri, and frontal areas [2].

    - LAB MEDIA: Figure 8.

    - LAB MEDIA: Figure 8. Video editor: Highlight the yellow and orange activation clusters across the different brain slices.

  - No significant brain activation differences were observed when contrasting TMT-B directly against TMT-A [1].

    - LAB MEDIA: Figure 8.