# Submission ID #: 67956

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=20726628

# Title: A Streamlined, Label-Free Real-Time 50% Tissue Culture Infectious Dose (TCID50) Assay using Impedance for Automated Viral Titer Quantification

# Authors and Affiliations:

# Jing Zhang, Yama A. Abassi, Nancy Li, Xiaobo Wang, Xiaoyu Zhang

# Agilent Technologies, Inc. La Jolla

# Corresponding Authors:

Xiaoyu Zhang		(xiaoyu.zhang@agilent.com)

# Email Addresses for All Authors:

Jing Zhang		(jing.zhang8@agilent.com)

Yama A. Abassi	(yabassi1998@gmail.com)

Nancy Li		(nancy.li@agilent.com)

Xiaobo Wang		(xbwang2009@gmail.com)

Xiaoyu Zhang		(xiaoyu.zhang@agilent.com)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

Current Protocol Length

Number of Steps:  22

Number of Shots:  42

# Introduction

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Xiaoyu Zhang: The scope of my research was to develop a method accelerating viral quantification, reducing hands-on time, enabling high-throughput CPE tracking.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

What research gap are you addressing with your protocol?

  - Xiaoyu Zhang: This protocol addresses the lack of a reliable, simple, fast, safe, and highly consistent method compared with conventional viral quantification approaches.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What advantage does your protocol offer compared to other techniques?

  - Xiaoyu Zhang: Our protocol offers fully automated, label-free, real-time, and objective viral quantification using impedance-based TCID50 analysis with integrated virology software.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

What are the current experimental challenges?

  - Xiaoyu Zhang: Current challenges include reliance on staining with toxic reagents in conventional TCID50 assays and subjectivity and analyst variability in label-free imaging methods.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

# Protocol

- Permissive Cell Seeding into a Biosensor Plate

Demonstrator: Jing Zhang

  - To begin, measure the impedance background of a biosensor plate before cell seeding [1]. Use a multichannel pipette to carefully transfer 50 microliters of prewarmed media into each well [2]. Use the reverse pipetting technique to minimize bubble formation and ensure consistent volume dispensation [3].

    - WIDE: Talent holding  an empty biosensor plate.

    - Talent dispensing prewarmed media into all wells of the biosensor plate using a multichannel pipette.

    - Shot of the reverse pipetting technique.

  - Place the plate into the plate station [1]. Then launch the system’s software and select the Virology module [2].

    - Talent placing the biosensor plate into the plate station.

    - SCREEN: 67956_2.2.2.mp4		00:03-00:08, 00:15-00:19

  - Enter relevant notes in the Exp Notes (Experiment-Note) tab [1] and input cell information including cell name and cell number in the Cell subtab of the Layout tab after selecting the corresponding wells [2].

    - SCREEN: 67956_2.3.1.mp4		00:03-00:13

    - SCREEN: 67956_2.3.2.mp4		00:02-00:16

  - Next, on the Schedule tab, click Add a Step to create Step 1 designated for plate background measurement [1-TXT]. Initiate recording by clicking Start [2].

    - SCREEN: 67956_2.4.1-&2.4.2.mp4.		00:01-00:09
TXT: Click Start to initiate recording

    - SCREEN: 67956_2.4.1-&2.4.2.mp4.

  - Remove the plate from the plate station [1] and place it in the laminar hood for cell seeding [2].

    - Talent removing the plate from the plate station.

    - Talent transferring it to the laminar hood.

  - Add 2 microliters of trypsin to detach the permissive cells from the T75 (T-Seventy-five) flask [1]. After counting the cells, adjust the cell suspension to the desired concentration [2].

    - Talent adding 2 µL trypsin to the T75 flask to detach the cells.

    - Talent adjusting the cell suspension concentration.

  - Pipette 100 microliters of the cell suspension to each well of the biosensor plate [1-TXT]. After incubating it for 30 minutes at room temperature, transfer the plate into the plate station inside incubator at 37 degrees Celsius with 5 percent carbon dioxide [2].

    - Talent dispensing 100 µL of the cell suspension into each well of the biosensor plate. TXT: TXT: Cells: HEK293A cells or MDCK cells

    - Talent placing the plate station into the incubator.

  - Navigate to the Schedule tab and add Step 2 using the default recording settings of 100 sweeps and a 15-minute interval [1].

    - SCREEN: 67956_2.8.1.mp4.		00:00-end

- Virus Serial Dilution, Inoculation, and Real-Time TCID50 Analysis

  - To prepare virus serial dilutions, calculate the total volume of medium needed for dilution [1-TXT]. Warm the medium in a 37-degree Celsius water bath for 30 minutes [2].

    - Shot of the required volume of medium. TXT: Medium: DMEM + 2% FBS + 1% penicillin /streptomycin

    - Talent placing the medium into a water bath to warm.

  - Rapidly thaw an aliquoted virus stock vial in the water bath for 1 to 2 minutes until all ice crystals have melted [1].

    - Talent thawing the virus stock vial in the water bath.

  - Transfer 20 microliters of virus stock into a 1.5-milliliter microcentrifuge tube containing 180 microliters of prewarmed media [1]. Pipette up and down to thoroughly mix the virus with the medium [2].

    - Talent transferring virus stock into the microcentrifuge tube containing media.

    - Talent pipetting the mixture up and down to mix thoroughly.

  - To equilibrate the virus dilutions, incubate the diluted virus samples at 37 degrees Celsius with 5 percent carbon dioxide for 20 minutes prior to inoculation [1].

    - Talent placing virus dilution tubes into the incubator.

  - Next, on the Layout tab of the software, select the wells designated for viral titer treatment [1]. Click the Treatment subtab and Enter the initial dilution value of 0.001 [2].

    - SCREEN: 67596_3.5.1-3.6.2.mp4.	00:01-00:06

    - SCREEN: 67596_3.5.1-3.6.2.mp4.	00:07-00:13

  - Enter 10 as the dilution factor and select the direction of dilution from left to right [1]. Then enter the virus inoculation volume and click Apply to save the treatment information [2].

    - SCREEN: 67596_3.5.1-3.6.2.mp4.	00:14-00:18

    - SCREEN: 67596_3.5.1-3.6.2.mp4.	00:19-00:24

  - Assign negative control wells by selecting the appropriate wells and checking Negative Ctrl (negative-control) under Well Type in the Treatment tab [1].

    - SCREEN: 67956_3.7.1.mp4.		00:01-00:08

  - To inoculate the virus, click Pause in the software to stop data acquisition [1]. Remove the plate from the plate station and place it in the laminar hood [2].

    - SCREEN: 67956_3.8.1.mp4.		00:00-end

    - Talent transferring the plate from the plate station to the laminar hood.

  - Using a multichannel pipette, carefully remove the culture media from each well [1]. Gently add 50 microliters of the virus dilution onto the top of the cell monolayer, avoiding disruption of the cells [2].

    - Talent aspirating culture media from the wells using a multichannel pipette.

    - Talent gently adding virus dilutions to the wells.

  - Place the plate back into the plate station [1]. On the Schedule tab, add Step 3 using the default recording settings to resume impedance recording [2].

    - Talent placing the plate back into the plate station.

    - SCREEN: 67596_3.10.2.mp4.	00:00-end

  - After 2 hours of virus inoculation, pause the recording [1] and remove the plate from the plate station to the laminar hood [2].

    - SCREEN: 67956_3.11.1.mp4.	00:00-end

    - Shot of plate being transferred to the laminar hood.

  - With a multichannel pipette, slowly add 150 microliters of prewarmed medium to each well, avoiding disruption of the cell monolayer [1]. Return the plate to the plate station [2] and resume recording by clicking Start [3].

    - Talent adding prewarmed medium gently to each well.

    - Shot of plate in the plate station

    - SCREEN: 67956_3.12.3.mp4.	00:00-end

  - To calculate the tissue culture infectious dose 50, navigate to the Data Analysis tab and add all assay wells to the plot chart [1]. In the Parameter section, select TCID50 (T-C-I-D-Fifty) from the dropdown list to initiate calculation [2].

    - SCREEN: 67956_3.13.1-&3.13.2.mp4.	00:00-00:05

    - SCREEN: 67956_3.13.1-&3.13.2.mp4.	00:06-00:12

  - Choose Bar Chart to calculate TCID50 at a specific time point [1] or select Time Dependent TCID50 to generate a TCID50 time course over the desired period [2].

    - SCREEN: 67956_3.14.1.mp4.	00:02-00:09

    - SCREEN: 67956_3.14.2.mp4.	00:02-end

# Results

- Results

  - In uninfected HEK293A (H-E-K-Two-Ninety-Three) cells, the Normalized Cell Index increased over the first 48 hours post-inoculation and then plateaued, indicating continued cell proliferation followed by confluence [1].

    - LAB MEDIA: Figure 2A. Video editor: Highlight the black trace

  - In Adv-GFP (Adenoviral-G-F-P)–infected HEK293A cells, the Normalized Cell Index dropped below the uninfected control approximately 50 hours post-inoculation and declined to zero by 100 hours, indicating virus-induced cytopathic effects [1].

    - LAB MEDIA: Figure 2A. Video editor: Highlight the red trace decreasing from around 50 hours until reaching zero.

  - Viral green fluorescent protein expression increased inversely as the Normalized Cell Index decreased during the first 100 hours post-inoculation in Adv-GFP–infected HEK293A cells [1].

    - LAB MEDIA: Figure 2A. Video editor: Highlight the green trace representing integrated green intensity increasing over time.

  - Live-cell imaging showed continued cell growth during the first 48 hours post-infection, followed by reduced confluency and increased green fluorescent protein signal as viral infection progressed [1], with loss of viable cells observed by approximately 100 hours [2].

    - LAB MEDIA: Figure 2B. Video editor: Highlight Brightfield and Brightfield + green images for 1-3

    - LAB MEDIA: Figure 2B. Video editor: Highlight Brightfield and Brightfield + green images for 4-5

  - Adv-GFP infection induced a dose-dependent decrease in the Normalized Cell Index, with higher virus concentrations causing earlier and steeper declines across replicate wells [1].

    - LAB MEDIA: Figure 3A. Video editor: Sequentially highlight the colored traces from 10-3 to 10-10

  - At a 10⁻⁸ dilution of Adv-GFP, the Normalized Cell Index exhibited high variability across replicate wells, indicating partial cytopathic effects [1].

    - LAB MEDIA: Figure 3A. Video editor: Show the pop up of 10-8 (right side ) then highlight the light blue traces labeled 10⁻⁸

  - The Adv-GFP TCID 50 increased over time and reached a plateau by approximately 150 hours post-inoculation [1].

    - LAB MEDIA: Figure 4. Video editor: Highlight the plotted points showing increasing values

  - In MDCK cells infected with influenza A virus, impedance measurements showed dose-dependent changes in the Normalized Cell Index compared to uninfected controls [1].

    - LAB MEDIA: Figure 5A. Video editor: Highlight multiple colored traces

Pronunciation Guide:

TCID50
Pronunciation link: https://www.merriam-webster.com/dictionary/TCID50
IPA: /ˌtiː.siː.aɪˈdiː ˈfɪfti/
Phonetic Spelling: tee·see·eye·dee fif·tee

Impedance
Pronunciation link: https://www.merriam-webster.com/dictionary/impedance
IPA: /ɪmˈpiːdəns/
Phonetic Spelling: im·pee·duhns

Quantification
Pronunciation link: https://www.merriam-webster.com/dictionary/quantification
IPA: /ˌkwɑːntəfəˈkeɪʃən/
Phonetic Spelling: kwahn·tuh·fuh·kay·shuhn

Permissive
Pronunciation link: https://www.merriam-webster.com/dictionary/permissive
IPA: /pərˈmɪsɪv/
Phonetic Spelling: per·miss·iv

Biosensor
Pronunciation link: https://www.merriam-webster.com/dictionary/biosensor
IPA: /ˈbaɪoʊˌsɛnsər/
Phonetic Spelling: bye·oh·sen·ser

Dispensation
Pronunciation link: https://www.merriam-webster.com/dictionary/dispensation
IPA: /ˌdɪspənˈseɪʃən/
Phonetic Spelling: dis·puhn·say·shuhn

Laminar
Pronunciation link: https://www.merriam-webster.com/dictionary/laminar
IPA: /ˈlæmənər/
Phonetic Spelling: lam·uh·ner

Trypsin
Pronunciation link: https://www.merriam-webster.com/dictionary/trypsin
IPA: /ˈtrɪpsɪn/
Phonetic Spelling: trip·sin

Inoculation
Pronunciation link: https://www.merriam-webster.com/dictionary/inoculation
IPA: /ɪˌnɑːkjəˈleɪʃən/
Phonetic Spelling: ih·nah·kyuh·lay·shuhn

Microcentrifuge
Pronunciation link: https://www.merriam-webster.com/dictionary/microcentrifuge
IPA: /ˌmaɪkroʊˈsɛntrəˌfjuːdʒ/
Phonetic Spelling: my·kroh·sen·truh·fyooj

Aliquoted
Pronunciation link: https://www.merriam-webster.com/dictionary/aliquot
IPA: /ˈælɪˌkwɑːt/
Phonetic Spelling: al·ih·kwot

Cytopathic
Pronunciation link: https://www.merriam-webster.com/dictionary/cytopathic
IPA: /ˌsaɪtəˈpæθɪk/
Phonetic Spelling: sy·tuh·path·ik

Adenoviral
Pronunciation link: https://www.merriam-webster.com/dictionary/adenovirus
IPA: /ˈædənoʊˌvaɪrəs/
Phonetic Spelling: ad·uh·noh·vy·ruhs

Confluence
Pronunciation link: https://www.merriam-webster.com/dictionary/confluence
IPA: /ˈkɑːnfluəns/
Phonetic Spelling: kon·floo·uhns

Proliferation
Pronunciation link: https://www.merriam-webster.com/dictionary/proliferation
IPA: /prəˌlɪfəˈreɪʃən/
Phonetic Spelling: պրuh·lif·uh·ray·shuhn

Confluency
Pronunciation link: No confirmed link found
IPA: /ˈkɑːnfluənsi/
Phonetic Spelling: kon·floo·uhn·see

Influenza
Pronunciation link: https://www.merriam-webster.com/dictionary/influenza
IPA: /ˌɪnfluˈenzə/
Phonetic Spelling: in·floo·en·zuh

Variability
Pronunciation link: https://www.merriam-webster.com/dictionary/variability
IPA: /ˌveriəˈbɪləti/
Phonetic Spelling: vair·ee·uh·bil·uh·tee