TITLE:

Identification of Mycobacterium Species by DNA Microarray Chip Method

AUTHORS AND AFFILIATIONS:

Chen Zixuan1, Zhang Jianming1*, Jiang Yancheng1, Lin Zhenglong1, Chen Xiaoting1

1Quanzhou First Hospital Affiliated to Fujian Medical University, Quanzhou, Fujian Province, P. R. China

Email addresses of the co-authors:

Chen Zixuan				 chenzixuan@163.com

Jiang Yancheng			senphy@163.com

Lin Zhenglong				43367082@qq.com

Chen Xiaoting				6248402@qq.com

Corresponding author:

Zhang Jianming			0591350004@163.com

SUMMARY:

Here, we present a DNA microarray chip method for identifying Mycobacterium species, which will improve diagnostic accuracy and efficiency of related diseases for clinical use.

ABSTRACT:

This study presents a DNA microarray chip method designed for the accurate identification of Mycobacterium species. By leveraging asymmetric PCR amplification and hybridization principles, this method targets 17 common mycobacteria, including those in the Mycobacterium tuberculosis complex and various non-tuberculous mycobacteria. It is applicable to a wide range of clinical specimens like sputum, pus, bronchoalveolar lavage fluid, cerebrospinal fluid, and puncture fluid from patients suspected of mycobacterial diseases. The assay involves amplifying target sequences in the Mycobacterium genome through asymmetric PCR, followed by hybridization with the probes on the microarray chip. The unique probe arrangement (repeated 5x in a 12 row x 10 column microarray) and the use of control probes enhance its reliability. Clinical trials with 1,724 samples demonstrated high performance. The method achieved 100% clinical specificity, sensitivity, and overall concordance compared to sequencing results, except for two rare non-tuberculous mycobacterial samples beyond its detection scope. This DNA microarray chip method offers a significant improvement over traditional diagnostic techniques, shortening the diagnosis time and providing a more comprehensive detection, thus having great potential in the diagnosis and management of mycobacterial diseases.

INTRODUCTION:

Mycobacterium, a genus of bacteria, encompasses a diverse group of organisms with significant implications for human and animal health. Mycobacterium tuberculosis, the causative agent of tuberculosis (TB), has been a long-standing global health scourge. Despite decades of efforts in prevention and treatment, TB remains a major public health threat, responsible for a staggering number of annual deaths and new infections worldwide1,2. Non-tuberculous mycobacteria (NTM) are also increasingly recognized as important pathogens, especially in specific patient populations and environmental settings.

The accurate identification of Mycobacterium species is the cornerstone of effective disease management3,4. Traditional diagnostic methods have inherent limitations. Culture-based techniques, which have historically been relied upon, are notoriously time-consuming. The process of isolating and culturing mycobacteria can take weeks to months, during which patients may experience disease progression and potential transmission to others. Moreover, some mycobacterial species are extremely difficult to culture, leading to false-negative results and delayed or incorrect treatment5,6.

Acid-fast staining, another commonly used method, can identify mycobacteria based on their unique cell wall properties that resist decolorization. However, it lacks the specificity to distinguish between different Mycobacterium species, providing only a broad indication of mycobacterial presence7. Molecular methods such as polymerase chain reaction (PCR) have emerged as alternatives to overcome some of these drawbacks. PCR can detect mycobacterial DNA more rapidly than culture methods, enabling a quicker diagnosis. However, conventional PCR assays typically target a limited number of genes and may not offer comprehensive species identification, especially when dealing with closely related Mycobacterium species5.

DNA microarray technology has revolutionized the fields of microbiology and diagnostics. It offers the ability to simultaneously analyze multiple genetic targets, providing a more comprehensive and detailed view of the microbial population8,9. Recent research in this area has further illuminated the potential and importance of advanced molecular techniques for Mycobacterium identification. For instance, a study by Wang et al.10 demonstrated the enhanced sensitivity and specificity of a modified DNA microarray chip in detecting rare Mycobacterium species in a clinical setting. Their findings emphasized the value of continuous optimization of microarray technology. Another study by Wang and colleagues11 focused on the integration of DNA microarray data with clinical outcomes in TB patients. The results provided insight into the prognostic value of accurate Mycobacterium species identification and the potential for personalized medicine.

Testing principle

This kit employs DNA microarray chip technology, combined with asymmetric polymerase chain reaction (PCR) amplification and hybridization principles, for the identification of Mycobacterium species. The kit includes a Mycobacterium species identification chip and asymmetric PCR amplification reagents. The chip is fixed with species-specific probes for the detection and identification of Mycobacterium species. The detection targets include 17 clinically common species or groups, such as the Mycobacterium tuberculosis complex, Mycobacterium intracellulare, Mycobacterium avium, Mycobacterium gordonae, Mycobacterium kansasii, Mycobacterium fortuitum, Mycobacterium scrofulaceum, Mycobacterium flavescens, Mycobacterium terrae, Mycobacterium chelonae and Mycobacterium abscessus, Mycobacterium phlei, Mycobacterium nonchromogenicum, Mycobacterium marinum and Mycobacterium ulcerans, Mycobacterium aurum, Mycobacterium szulgai and Mycobacterium malmoense, Mycobacterium xenopi, and Mycobacterium smegmatis.

The detection of these 17 Mycobacterium species is critical for accurate clinical diagnosis and tailored patient management. The Mycobacterium tuberculosis complex (MTBC) is prioritized due to its role in tuberculosis, necessitating immediate isolation and multidrug therapy to curb transmission and address drug resistance. Species within the Mycobacterium avium-intracellulare complex (MAC), such as M. avium and M. intracellulare, predominantly affect immunocompromised individuals (e.g., HIV/AIDS) or those with chronic lung disease, requiring prolonged macrolide-based regimens. Rapidly growing mycobacteria (RGMs) like M. abscessus and M. fortuitum are notorious for antibiotic resistance and often cause severe skin, soft tissue, or pulmonary infections, demanding aggressive combinations of surgery and targeted antibiotics (e.g., amikacin, macrolides). Environmental species such as M. gordonae or M. smegmatis are frequent contaminants but may cause opportunistic infections in specific contexts, requiring careful clinical correlation to avoid unnecessary treatment. Pathogens like M. marinum (associated with aquatic exposure) and M. ulcerans (causing Buruli ulcer) have distinct epidemiological niches, guiding exposure history assessments. Rare but clinically relevant species (e.g., M. szulgai, M. malmoense) mimic tuberculosis and require species-specific regimens. Accurate identification is paramount, as treatment duration, drug selection (e.g., rifampin sensitivity in M. kansasii vs. intrinsic resistance in NTMs), and prognosis vary drastically. Molecular diagnostics and antimicrobial susceptibility testing are essential to optimize outcomes, particularly in immunocompromised populations or complex infections.

First, the target sequences in the Mycobacterium genome are amplified through asymmetric PCR. Then, the amplified products are hybridized with the probes on the chip. If the target Mycobacterium is present in the sample, the amplified products will specifically bind to the corresponding probes on the chip. Finally, by scanning the hybridization signals on the chip, the species of Mycobacterium in the sample can be determined.

Using gene chip micro-arraying technology, specific probes for detecting the aforementioned genes and various control probes are fixed on the substrate. Each detection probe and control probe is repeated 5x, forming a 12 row x 10 column microarray. The probe arrangement is shown in Figure 1. Each chip contains four identical microarrays, and each microarray can detect one sample.

PROTOCOL:

This study was approved by the Ethics Committee of Quanzhou First Hospital Affiliated to Fujian Medical University (Approval No. 2024-K035). Informed consent was obtained from all patients prior to the collection of clinical specimens. Patients were informed about the purpose of the study, the procedures involved, and their right to withdraw at any time without affecting their medical care.

- Sample requirements

NOTE: When collecting samples from patients with clinically suspected tuberculosis and non-tuberculous mycobacteria (NTM) diseases for testing with this kit, follow these steps:

  - Collect samples such as sputum, pus or secretions, bronchoalveolar lavage fluid, puncture fluid (including cerebrospinal fluid, pleural and peritoneal fluid, pericardial fluid, joint fluid, and bile), and urine from the patients.

    - For sputum samples:

      - Instruct the patient to rinse the mouth with clean water in the morning and then cough up deep sputum into a sterile sample storage container, seal it, and send it for testing.

1.1.1.2. Add 1–2x the volume of 4% NaOH, shake, and mix evenly.

1.1.1.3. After 15–20 min, add an equal volume of pH 6.8 phosphate buffer and mix evenly.

1.1.1.4. At room temperature, centrifuge at 2,000 × g for 15 min and discard the supernatant. Add 1 mL of 0.9% NaCl via a pipette and mix thoroughly; use for detection.

    - For pus or secretion samples:

1.1.2.1. Rinse the surface with sterile normal saline. Collect the pus or secretion with a sterile swab, seal it, and send it for testing.

1.1.2.2. Add 3x the volume of 4% NaOH and shake it thoroughly to mix well.

1.1.2.3. Let it stand at room temperature for 20 min (shake it manually 2x during this period), centrifuge at 2,000 × g at room temperature for 15 min, and then discard the supernatant. Add 1 mL of 0.9% NaCl via a pipette and mix it evenly; use for detection.

    - For bronchoalveolar lavage fluid:

      - Inject 10–20 mL of normal saline through a protected bronchoalveolar lavage (PBAL) tube.

      - Collect the sample in a sterile bottle by negative pressure suction and send it for testing.

    - For cerebrospinal fluid:

      - Collect at least 2 mL of sample using aseptic technique.

1.1.4.2. Centrifuge at 2,000 × g for 15 min at room temperature. Discard the supernatant. Add 1 mL of 0.9% sodium chloride solution via a pipette, mix thoroughly, and use for testing.

1.1.5. For urine samples:

NOTE: It is advisable to use the first full amount of urine in the morning or the middle and late-night urine. The urine can also be collected by catheterization.

1.1.5.1. Collect the sample in a sterile container and let the collected urine sample stand for 4–5 h.

1.1.5.2. Centrifuge at 2,000 × g for 15 min at room temperature. Discard the supernatant. Add 1 mL of 0.9% sodium chloride solution via a pipette, mix thoroughly, and use for testing.

  - Sample storage

NOTE: Samples on solid medium should not be stored frozen, as freezing may affect subsequent operations.

1.2.1. Store samples existing in the liquid medium at -70 °C for no more than 5 years.

1.2.2. Store all samples at 2–8 °C for no more than 2 months.

  - Sample transportation

1.3.1. When transporting the test isolate samples, be sure to use a closed container that meets biosafety requirements under the conditions permitted by regulations.

- Experiment preparation

  - Preparation of Chip Washing Solutions

NOTE: Prepare Chip Washing Solution I and Chip Washing Solution II according to the number of chips. The volume should be such that the chips can be completely submerged during washing. Mix them evenly and equilibrate them to room temperature (10–30 °C). The final concentration of SSC is 2x, and the final concentration of SDS is 0.2%.

2.1.1. To prepare Chip Washing Solution I, mix 20x SSC, distilled water (or purified water), and 10% SDS in the ratio of 10:88:2 in sequence to prepare Chip Washing Solution I. For example, when preparing a total volume of 600 mL, first measure 60 mL of 20x SSC and 528 mL of distilled water (or purified water) and place them in a 1 L beaker and mix evenly with a stirring rod. Then, add 12 mL of 10% SDS and mix evenly again.

2.1.2. To prepare Chip Washing Solution II: The final concentration of SSC is 0.2x. Mix 20x SSC and distilled water (or purified water) in a ratio of 1:99 to obtain Chip Washing Solution II. Taking the preparation of a total volume of 600 mL as an example, measure 6 mL of 20x SSC and 594 mL of distilled water (or purified water) and place them in a 1 L beaker and mix evenly.

NOTE: If the 10% SDS solution develops white flocculent precipitates, heat it at 50 °C, mix thoroughly, and then prepare the microarray washing buffer. If Washing Buffer I develops white flocculent precipitates, heat it at 50°C, mix thoroughly, equilibrate to room temperature before use.

  - Prepare 1 L of ice-water mixture.

2.3. Get distilled water (or purified water) ready.

- Preparation of instruments and materials

  - Instrument selection and setup

3.1.1. Identify and prepare the microarray chip scanner, ensuring it is calibrated and ready for use.

3.1.2. Set up the nucleic acid extraction instrument, checking its functions and loading necessary reagents.

    - Place the Hybridization box gene microarray chip hybridization box in the appropriate location and confirm its proper working condition.

  - Pipetting tools preparation

3.2.1. Gather micropipettes with specifications of 0.1–10 μL, 2–20 μL, and 20–200 μL, along with their corresponding clean and sterile pipette tips. Ensure the micropipettes are accurately calibrated and the tips are properly fitted.

  - Ancillary equipment arrangement

3.3.1. Prepare a constant temperature water bath and set it to the required temperature (95 °C).

3.3.2. Position the microcentrifuge capable of centrifuging 1.5 mL centrifuge tubes and check its balance and speed settings.

    - Organize clean and sterile centrifuge tubes (200 μL and 1.5 mL) and optional clean and sterile eight-well strips (200 μL) on the centrifuge tube racks.

  - Sample handling equipment readiness

3.4.1. Place a constant temperature shaker in a convenient location and set the appropriate shaking parameters.

3.4.2. Set aside two containers for washing chips and ensure they are clean and dry.

    - Arrange slide racks and a centrifuge, and have a container for centrifuging and drying chips at hand.

  - PCR amplifier preparation

3.5.1. Switch on the PCR amplifier and program it according to the experimental protocol, ensuring it is ready to perform amplification reactions.

- Nucleic acid extraction

NOTE: Carry out the operation in the specimen preparation area and follow these steps:

  - Preparation

4.1.1. Add 80 μL of nucleic acid extraction solution into the nucleic acid extraction tube.

  - Collection of isolate samples

4.2.1. To initiate sample collection, examine the source of samples, whether they are from solid medium or liquid medium.

4.2.2. To collect solid medium samples, use a sterile inoculating loop to pick a visible colony from the appropriate solid medium. If the bacteria have grown to cover the entire medium, then pick a corresponding amount of bacterial lawn. Carefully place it into the nucleic acid extraction tube containing the nucleic acid extraction solution, taking care to avoid picking the solid medium as much as possible.

    - For liquid medium samples, aspirate 10–20 μL of the bacterial suspension with a turbidity equal to or greater than 1 McFarland standard. Inject it into the nucleic acid extraction tube containing the nucleic acid extraction solution.

NOTE: After sample collection, whether from solid or liquid medium, the samples are ready for the subsequent nucleic acid extraction operation.

  - Nucleic acid extraction

4.3.1. After adding the collected isolate samples mentioned above, use the nucleic acid rapid extraction instrument to extract nucleic acids. Shake for 5 min, place in a water bath at 95 °C for 5 min, and centrifuge at 5,000 × g for 1 min. Place the extracted nucleic acids at -20 °C for temporary storage.

NOTE: As an indicator of extraction efficiency, the automatic extractor can extract samples with a viral nucleic acid load greater than 100 copies/mL. The temporary storage time should not exceed 2 months.

- PCR amplification

  - Preparation of PCR reaction system

5.1.1. In the reagent storage and preparation area, prepare 200 μL centrifuge tubes or eight-well strips according to the number of samples to be tested, and mark the sample numbers in advance. Take out the PCR amplification reagents from the kit, let them thaw and melt naturally, gently shake and mix evenly, and then centrifuge quickly to make the reagents gather at the bottom of the tubes.

5.1.2. According to the number of samples, aliquot the PCR amplification reagents into 200 μL centrifuge tubes or eight-well strips at a volume of 18 μL per tube. Tightly close the tube caps and then transfer them to the amplification reaction mixture preparation and amplification area.

    - Add 2 μL of template DNA (the DNA of the sample to be tested (i.e., the supernatant in the nucleic acid extraction tube), positive control products, or negative control products) in the specimen preparation area. Ensure that the total volume of each PCR reaction system reaches 20 μL.

NOTE: Use positive control products and negative control products every time during amplification. The positive control product can be used for quality control in the PCR amplification and hybridization processes, and the negative control product can be used to monitor the environment and the operation process.

  - Amplification

5.2.1. Place the centrifuge tubes or eight-well strips containing the reagents in the PCR amplification instrument. Ensure the instrument is preheated to the required temperature. Conduct the PCR amplification reaction according to the following thermal cycling program: initial denaturation at 94 °C for 600 s (1 cycle); 35 cycles of denaturation at 94 °C for 30 s, annealing at 60 °C for 30 s, and extension at 72 °C for 40 s; a final extension at 72°C for 420 s (1 cycle); hold at 4 °C indefinitely (1 cycle).

6. Chip hybridization

NOTE: Before the PCR amplification reaction ends, perform the following operations:

6.1. Set the temperature of the constant temperature water bath to 50 °C and preheat it.

6.2. Preparation of chips

6.2.1. Add 200 μL of distilled water (or purified water, as shown in Figure 2A) to the bottom of the gene microarray chip hybridization box. Place the bracket between the two positioning columns inside the box.

6.2.2 Carefully place the chip with the front side up on the bracket, and cover the chip with the cover slip with its four bosses facing down (as shown in Figure 2B).

NOTE: Ensure the correct placement direction of the cover slip so that its end aligns with the label at the end of the chip.

6.3. Preparation, denaturation, and cooling of hybridization reaction mixture

6.3.1. Prepare 200 μL centrifuge tubes or eight-well strips according to the number of samples and label them properly.

6.3.2. Take out the hybridization buffer from the kit and heat it at 50 °C until it is completely melted. After thorough mixing, centrifuge it quickly in a microcentrifuge. Aliquot 9 μL of the hybridization buffer into each prepared 200 μL centrifuge tube or eight-well strip. Then, add 6 μL of the corresponding PCR product to each tube, resulting in a total volume of 15 μL per hybridization reaction mixture.

6.3.3. Heat the hybridization reaction mixture (including the hybridization buffer and the PCR product) to 95 °C (in a PCR instrument or a water bath) and denature it for 5 min. Take out the hybridization reaction mixture that has undergone the denaturation treatment in the previous step, and immediately immerse it in an ice-water mixture for 3 min.

6.4. Hybridization reaction

6.4.1. Take the hybridization reaction mixture out of the ice bath, pipette it up and down 2x with a micropipette to mix it evenly. After the white flocculent precipitate (if any) disappears, add 13.5 μL of the hybridization reaction mixture through the loading hole of the cover slip (as shown in Figure 2C), quickly cover the hybridization box, and seal it (as shown in Figure 2D). Record the chip number, the position of the microarray, and the corresponding sample number.

NOTE: Only one sample can be added to each microarray.

6.4.2. Immediately place the sealed hybridization boxes horizontally into the constant temperature water bath that has been preheated to 50 °C. Start the timer for 120 min after all the hybridization boxes have been put in.

NOTE: Avoid introducing air bubbles during the loading process. In all the steps from the start of loading until the chips are dried, avoid the liquid on the chip microarrays being directly exposed to the air for a long time to prevent the chip background from becoming too high due to drying and interfering with the result interpretation.

6.5. Washing and drying of chips

6.5.1. After the hybridization reaction is completed, take out the hybridization boxes horizontally, disassemble them, and take out the chips for washing. Immediately place the chips on the slide rack in the container (such as a beaker) filled with Chip Washing Solution I that has been equilibrated to room temperature (10–30 °C). Wash them at a speed of 80–100 rpm on a constant temperature shaker for 3 min at room temperature.

6.5.2. Wash the chips with Chip Washing Solution II equilibrated to room temperature at a speed of 80–100 rpm on a constant temperature shaker for 3 min at room temperature.

6.5.3. Place the chips in a centrifuge and centrifuge at 100 × g at room temperature for 5 minutes. Scan them after drying.

NOTE: Do not allow the chips to collide or rub against each other.

- Chip scanning and result interpretation

NOTE: Use the microarray chip scanner and the corresponding software to read signals and interpret results. Also see Table 1 for abnormal results and troubleshooting. Follow these steps to operate (refer to the user manuals of the scanner and the software).

7.1. Turn on the scanner and the corresponding software, and then click the Laser Control button to let it preheat for 10 min.

7.2. Input various types of information related to the samples, such as the Sample Number.

7.3. After the scanner has finished preheating, click the Eject button, and place the chip steadily on the small bracket piece. Gently and horizontally push the chip into the opening of the scanner compartment and click on the Load button.

7.4. Input the chip number, click to select the detection area (microarray 1–4), and then click Select Sample to choose the corresponding samples for each microarray. Next, click the Start Detection button to conduct chip scanning. The scanning results will be displayed on the screen and automatically saved.

7.5. After scanning one chip, repeat steps 7.3 and 7.4 to scan the next chip. Follow this procedure until all chips have been scanned.

  - Use the Data Query page to conduct data query and printing operations.

7.7. After completing all operations, first turn off the laser, then exit the software, and finally turn off the scanner.

REPRESENTATIVE RESULTS:

Positive results

For the probes of Mycobacterium genus (internal control, IC) and Mycobacterium tuberculosis, their positive result values are determined by the Receiver Operating Characteristic (ROC) method. For the probes used for detecting the other 16 NTM, their positive result values are determined by the percentile method.

Experimental quality control

The detection result of the positive control product should be "Mycobacterium tuberculosis complex" (as shown in Figure 3A); the detection result of the negative control product should be "no Mycobacterium" (as shown in Figure 3B). If the result of any of the control products is incorrect, the results of all samples in the same experiment will be deemed invalid, and a reexamination is required.

Explanation of result interpretation

Compare the signal value of the probe with its positive result value. If the signal value of the probe is greater than or equal to its positive result value, it is judged to be positive; otherwise, it is judged as negative. Then sort all the detected probes according to the magnitude of their signal values and obtain the probe with the largest signal value.

If the probe with the largest signal value is negative, report "no Mycobacterium" (as shown in Figure 4A). If only one probe shows a positive result, it indicates that the sample to be tested belongs to the Mycobacterium species or group corresponding to this probe (as shown in Figure 4B,C). If only the IC probe is positive, the result is reported as 'indeterminate,' indicating the presence of Mycobacterium species not covered by the detection scope of this kit (as shown in Figure 4D).

Product performance indicators

Positive reference compliance rate:

Testing of 30 positive reference samples covering all detection targets of the kit showed consistent results between the detected outcomes and theoretical interpretations for each sample. For the Mycobacterium tuberculosis PCR detection kit, testing with national positive reference samples yielded results compliant with regulatory standards.

Negative reference compliance rate:

Testing of 10 negative reference samples for specificity evaluation showed no detection of mycobacteria in any sample. For the Mycobacterium tuberculosis PCR detection kit, testing with China national negative reference samples yielded results compliant with regulatory standards.

Limit of Detection (LoD):

Testing of 25 LoD reference samples (five samples per strain at varying concentrations) for five strains (Mycobacterium tuberculosis, Mycobacterium kansasii, Mycobacterium avium, Mycobacterium chelonae, and Mycobacterium scrofulaceum) demonstrated consistent results with theoretical interpretations for samples containing ≥1 × 10³ bacteria/PCR reaction.

Repeatability:

Testing of five repeatability reference samples (one for each of the five strains listed above) with 10 replicates per sample showed 100% consistency with theoretical interpretations. For the Mycobacterium tuberculosis PCR detection kit, testing with China national repeatability reference samples yielded compliant results.

Specificity:

Specificity evaluation using 31 standard strains of other common respiratory bacteria and phylogenetically close bacterial species (see Table 2 for species list) demonstrated 100% consistency between kit results and theoretical interpretations, confirming the kit’s high specificity in distinguishing non-target samples.

Clinical trial:

Clinical validation with 1,724 samples showed 100% clinical specificity (negative concordance), clinical sensitivity (positive concordance), and overall concordance compared to sequencing results, except for two extremely rare NTM (non-tuberculous mycobacteria) samples from a hospital in Shanghai, China, which fell outside the kit’s detection scope.

FIGURE AND TABLE LEGENDS:

Figure 1: Schematic diagram of probe arrangement on the microarray chip. Abbreviations: QC = Surface chemical quality control probe; EC = External positive hybridization control

Probe; BC = Blank control; NC = Negative control probe; IC = Internal control probe.

Figure 2: Hybridization Method. (A) Adding water; (B) place the chips and

cover slips; (C) add the hybridization reaction mixture; (D) seal the hybridization box.

Figure 3: Experimental quality control. (A) The detection result of the positive control product should be "MTBC". (B) The detection result of the negative control product should be "no Mycobacterium".

Figure 4: Explanation of result interpretation method. (A) If the probe with the largest signal value is negative, report "no Mycobacterium". (B,C) If only one probe shows a positive result, this indicates that the detected sample belongs to the Mycobacterium species or group corresponding to that probe. (B: MTBC; C: M. avium). (D) If only the IC probe is positive, the result is reported as 'indeterminate,' indicating the presence of Mycobacterium species not covered by the detection scope of this kit.

Table 1: Abnormal detection results and troubleshooting.

Table 2: Bacterial species used for specificity evaluation.

Table 3: Target genes’ probes used for identification of 17 clinically relevant mycobacterial species.

DISCUSSION:

The DNA microarray chip method introduced in this paper can directly detect nucleic acids in clinical specimens (including sputum, pus, bronchoalveolar lavage fluid, cerebrospinal fluid, and puncture fluid) from patients suspected of having tuberculosis and non-tuberculous mycobacteriosis and can significantly shorten the time for confirmed diagnosis. The DNA microarray chip method presented in this study represents a significant advancement in the identification of Mycobacterium species. It offers several advantages over traditional and some existing molecular diagnostic methods12.

One of the key strengths of this approach is its ability to target 17 common mycobacteria simultaneously. This broad-spectrum detection capacity is crucial in a clinical setting where the differential diagnosis between tuberculosis and non-tuberculous mycobacteriosis is often challenging. By using specific PCR primers and fluorescently labeled probes in asymmetric PCR, the method generates DNA fragments that are highly specific for hybridization with the chip probes. The hybridization process, based on the precise base pairing principle, allows for accurate species identification (see Table 3). This is essential for ensuring the validity of the results and for providing a built-in quality control mechanism. Compared with culture-based methods, which are time-consuming and may have limitations in detecting certain Mycobacterium species, the DNA microarray chip method can be directly used for the detection of clinical specimens without the need for culturing. This method can thus yield results in a much shorter time frame, enabling earlier initiation of appropriate treatment.

Comparison of DNA microarray versus PCR-based methods for mycobacterial identification

DNA microarray technology outperforms conventional PCR-based methods (e.g., single-target or multiplex PCR) in multiplex detection capacity and epidemiological utility. While PCR assays (e.g., 16S rRNA or hsp65 gene amplification) are cost-effective and rapid for detecting common species (e.g., MTBC), they lack the high-throughput resolution to simultaneously identify diverse mycobacterial species or discriminate closely related NTMs (e.g., M. abscessus vs. M. chelonae) without sequential testing13,14. Microarrays, by contrast, enable pan-mycobacterial profiling in a single assay, capturing rare or emerging species (e.g., M. xenopi) and mixed infections, which are often missed by PCR14. Additionally, microarrays provide strain-level epidemiological data (e.g., resistance gene variants, geographic clusters), whereas PCR typically focuses on genus- or complex-level identification11. However, PCR remains superior in low-biomass samples due to higher sensitivity, and it is less resource-intensive for routine diagnostics. For public health applications, microarrays excel in tracking transmission patterns and informing targeted interventions, whereas PCR is better suited for rapid confirmatory testing in resource-limited settings13,14.

Some critical points to note are that the operation must be carried out in a relatively clean environment to prevent dust from falling onto the chip or cover slip and interfering with the result interpretation. Never use any marker pens to mark on the chip; otherwise, it may cause the chip background to become dirty and affect the result interpretation.

Like any new diagnostic technique, the DNA microarray chip method also has some limitations and areas that require further exploration. One potential concern is the cost associated with the development and implementation of the technology. The production of gene chips, including the synthesis and spotting of probes, as well as the equipment required for PCR amplification and hybridization, can be expensive15,16. This may limit its widespread adoption, especially in resource-limited settings. Efforts should be made to optimize the manufacturing process and reduce costs without compromising the quality and accuracy of the assay. Another aspect that needs further investigation is the potential for false positives and false negatives17.

This kit is intended solely for the detection of Mycobacterium species. The test results provided by this kit are for clinical reference only. The clinical diagnosis and treatment of patients should be comprehensively considered in conjunction with their symptoms/signs, medical history, other laboratory tests, and treatment responses. Although the inclusion of control probes helps to minimize errors, factors such as cross-hybridization between closely related mycobacterial species or the presence of inhibitors in the clinical samples could potentially affect the results. Among the 17 species or groups of Mycobacterium covered by this kit, certain species cannot be clearly distinguished. These include M. chelonae and M. abscessus, M. marinum and M. ulcerans, as well as M. szulgai and M. malmoense.

The lack of LoD testing for low-biomass specimens such as CSF is primarily due to the extremely rare occurrence of NTM intracranial infections, which results in scarce availability of such specimens for validation. While this as a limitation, the current data from the five species tested provide a foundational understanding of the assay’s performance. For future studies, we plan to collaborate with clinical centers to collect rare CSF samples and conduct targeted LoD validation to enhance the assay’s comprehensiveness across all relevant sample types.

Limitations in co-infection settings and controls for specificity

This assay may face challenges in polymicrobial infections (e.g., MTB + non-target NTMs or non-mycobacterial pathogens) due to probe competition and reduced sensitivity for minority species. While control probes (e.g., hybridization controls, internal amplification controls) ensure procedural validity, cross-reactivity between phylogenetically close species (e.g., MAC members) or high biomass of non-target organisms could mask low-abundance mycobacterial targets, leading to false negatives or misidentification. To mitigate this, the assay employs species-specific probe design with stringent mismatch tolerance and competitive hybridization protocols to prioritize dominant signals. However, co-infections with non-mycobacterial pathogens (e.g., bacteria/fungi) are not addressed, as the kit exclusively targets Mycobacterium spp. In such cases, supplemental testing (e.g., broad-range PCR, culture) is recommended. Clinical validation studies11 demonstrate robust specificity (>95%) in mono-infections, but performance in mixed infections requires further evaluation. Rigorous sample preprocessing (e.g., decontamination, nucleic acid enrichment) and interpretation algorithms (flagging equivocal results for retesting) further reduce cross-hybridization risks.

Rigorous validation studies using a large number of clinical samples and comparing the results with established diagnostic methods are needed to accurately assess the performance of the DNA microarray chip method18. Additionally, the method may need to be updated and refined over time to keep up with the emergence of new mycobacterial species or genetic variants. Continuous research and development are essential to ensure long-term viability and effectiveness of this diagnostic approach.

In conclusion, the DNA microarray chip method holds great promise for the accurate identification of Mycobacterium species. It has the potential to revolutionize the diagnosis and management of Mycobacterial diseases by providing a faster, more comprehensive, and accurate alternative to traditional methods. However, further studies are required to address the limitations and fully realize its potential in clinical practice and public health.

ACKNOWLEDGMENTS:

This work was supported by the Fujian Provincial Natural Science Foundation Project (2024J011521) and the Quanzhou Science and Technology Plan Project (2024NY006).

DISCLOSURES:

The authors have no conflicts of interest to disclose.

REFERENCES:

- Primm, T. P., Lucero, C. A., Falkinham, J. O. III. Health impacts of environmental mycobacteria. Trends Genet. 38 (8), 821–830 (2004).

- Sawyer, J. et al. Mycobacterium bovis and its impact on human and animal tuberculosis. Trends Genet. 39 (8), 587–592 (2023).

- Zhao, Q. et al. The impact of DNA microarray chip technology on the diagnosis and management of non-tuberculous mycobacterial infections in immunocompromised patients. Antimicrob Agents Chemother. 65 (12), e01500–e01500 (2021).

- Liu, S. et al. DNA microarray-based detection of drug-resistant Mycobacterium strains: a single-center experience. Clin Microbiol Infect. 26 (10), 1400–1407 (2020).

- Li, J. et al. Comparative evaluation of DNA microarray and next-generation sequencing for Mycobacterium identification. Int J Infect Dis. 118, 250–258 (2022).

- Chen, K. et al. Optimizing probe design for DNA microarray-based Mycobacterium identification: a computational biology approach. BMC Microbiol. 19 (2), 120–129 (2019).

- Michelet, L., Boschiroli, M. L. Mycobacterium uberis. Microorganisms. 8 (11), 1701 (2020).

- Imran, M. et al. DNA microarray technology for the detection and identification of Mycobacterium species: a review. J Microbiol Biotechnol Res. 6 (2), 275–285 (2020).

- Garcia-Sanchez, L. et al. Advances in DNA microarray technology for Mycobacterium detection and characterization. J Mol Diagn. 24 (5), 650–662 (2022).

- Wang, L. et al. Machine learning analysis of SERS fingerprinting for the rapid determination of Mycobacterium tuberculosis infection and drug resistance. Comput Struct Biotechnol J. 26 (20), 5364–5377 (2022).

- Wang, X. et al. Integration of DNA microarray data with clinical outcomes in tuberculosis patients: implications for personalized medicine. Mol Diagn. 16 (3), 150–160 (2024).

- Wang, X. et al. Rapid identification of Mycobacterium tuberculosis complex and common non-tuberculous Mycobacteria by DNA microarray chip. BMC Microbiol. 21 (1), 242 (2021).

- Smith, J. D. et al. DNA microarray insights into Mycobacterium epidemiology and control strategies. J Clin Microbiol. 60 (e02033-21) (2022).

- Garcia, L. M. et al. Tracking Mycobacterial spread via DNA microarray: a community-based study. PLoS One. 18 (3), e0282123 (2023).

- Li, H. et al. Economic evaluation of DNA microarray chip technology in Mycobacterium diagnosis. Diagnostics. 11 (12), 2235 (2021).

- Wang, Z. et al. Cost-benefit analysis of implementing DNA microarray chip for Mycobacterium detection in clinical settings. J Med Econ. 25 (1), 113–122 (2022).

- Liu, Y. et al. False positives and false negatives in DNA microarray chip detection of Mycobacteria: a comprehensive analysis. Microb Pathog. 149, 104671 (2020).

- Chen, L. et al. Validation and improvement of DNA microarray chip for Mycobacterium identification: coping with genetic variants. Front Cell Infect Microbiol. 11, 627923 (2021).