# Submission ID #: 68238

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=20812638

# Title: Workflow Using a Cryogenic Coincident Fluorescence, Electron, and Ion Beam Microscope for Targeted Milling of Cells

# Authors and Affiliations:

# Jue Wang1,2, Anthony V. Sica2, Grant J. Jensen3, Peter D. Dahlberg2,4

# 1Division of Biology and Biological Engineering, California Institute of Technology

# 2SLAC National Accelerator Laboratory

# 3Department of Chemistry and Biochemistry, Brigham Young University

# 4Department of Photon Science and Structural Biology, Stanford University

# Corresponding Authors:

Peter D. Dahlberg 		pdahlb@slac.stanford.edu

# Email Addresses for All Authors:

Jue Wang 			jpwang@slac.stanford.edu

Anthony V. Sica 		asica1@slac.stanford.edu

Grant J. Jensen 		grant_jensen@byu.edu

Peter D. Dahlberg 		pdahlb@slac.stanford.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Yes, all done

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done

Current Protocol Length

Number of Steps:  23

Number of Shots:  41

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Jue Wang: My research develops optical microscopy methods to preserve rare cellular structures in cryogenic lamella during focused ion beam milling, enabling their detailed analysis by cryogenic electron tomography to uncover native biological organization.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.1

What are the most recent developments in your field of research?

  - Jue Wang: Optical microscopy has recently been integrated into cryoFIB-SEM systems, but rarely at the same focal position as the FIB. This necessitates registration-based guidance approaches to preserve fluorescently labelled structures in final lamella.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:2.19

What are the current experimental challenges?

  - Jue Wang: Accuracy of Fluorescence-based lamella targeting is limited by registration approaches, which suffer from registration error, diffraction-limited resolution, refractive index mismatch, and milling-induced motion.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What significant findings have you established in your field?

  - Peter Dahlberg: Simultaneous optical microscopy and FIB milling removes many of these sources of guidance error by enabling real-time feedback for precise milling that preserves fluorescently labeled structures for cryo-electron tomography within the final lamella.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What new scientific questions have your results paved the way for?

  - Jue Wang: Our research makes it possible to see cellular structures that were previously challenging to directly visualize by cryo-ET, shedding light on molecular mechanisms behind a wide spectrum of key macromolecular machines.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.4

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

- High-Precision Targeting for Cryo-Correlative Light and Electron Microscopy

Demonstrator: Jue Wang

  - To begin, manually cool the transfer station using liquid nitrogen once the system temperature nears cryogenic levels [1].

    - WIDE: Talent pouring liquid nitrogen into the transfer station while checking the temperature gauge.
Videographer’s Note: I included several angles of the liquid nitrogen being poured into the transfer station, as well as the temperature gauge, which is across the room from the transfer station. I've slated them with the additional info

  - Load one clipped Autogrid into liquid nitrogen to cool it and then into the sample cassette designed for the cryogenic light microscopy system [1]. Ensure that the Autogrid is positioned between the glued coverslip and the spacer [2]. 
NOTE: The VO has been edited to accommodate the added action in 2.2.1

    - Talent placing a clipped Autogrid into the cassette.
Videographer’s Note: For 2.2.1, before placing the clipped autogrid into the cassette, the talent places the tools in the liquid nitrogen. This is to cool down the tools.

    - Shot of the Autogrid correctly seated between the glued coverslip and spacer.

  - Tighten the ring screw to close the cassette [1], and insert it into the slot on the transfer station [2].

    - Talent tightening the ring screw on the cassette.

    - SCREEN: Talent inserting the closed cassette into the transfer station slot.
NOTE: Shot by videographer

  - Attach the transfer station to the cryogenic light microscopy system [1]. Use the internal chamber camera to guide the sample cassette onto the sample stage [2].

    - Talent connecting the transfer station to the system interface.

    - SCREEN: Display of the internal chamber camera feed guiding the sample cassette accurately onto the sample stage.
NOTE: Shot by videographer

  - In the light microscopy software, move the sample stage from the Loading position to the predefined three-beam position [1]. In the FIB-SEM (Fib--Sem) software, set the to 100X (Hundred-Ex) and click on the top-left window [2].

    - SCREEN: 2.5.1-(1).mp4	00:08-00:24, 01:34-01:40

    - SCREEN: 2.5.2-(1).mp4 	00:00-00:09

  - In the Alignment tab of the light microscopy software, click on the z and select the play function to acquire a SEM atlas of the grid [1]. Identify a square area with broken carbon film and double-click on it to move the sample stage to that location [2].

    - SCREEN: 2.6.1-(1).mp4	00:00-00:10

    - SCREEN: 2.6.2.mp4	00:00-00:22

  - Return to the software interface, activate the FIB and acquire an image [1]. Adjust the magnification using the drop down manu on the top left corner [2]. Press Play to display the FIB image [3-TXT]. Then set the step size using the slide bar [4].

    - SCREEN: 2.7.1_fib_added.mp4.mp4	00:30-00:53

    - SCREEN: 2.7.1_fib_added.mp4.mp4	00:55-01:07

    - SCREEN: 2.7.2_updated-(1).mp4		00:00-00:03
TXT: Set step size using slide bar

  - Now click +z (plus-Z) and -z (minus-Z) to adjust the stage z position until the horizontal line present in the ion and the electrode images coincide with the same feature [1]. Select the SEM tab in the light microscopy software [2]. Use the plus or minus X and Y buttons to align a common feature between the SEM and ion images [3].

    - SCREN: 2.7.2_updated-(1).mp4	00:12-00:42, 01:19

    - SCREEN: 2.7.3-(1).mp4	00:00-00:05

    - SCREEN: 2.7.3-(1).mp4	00:28-00:34, 01:57-02:05

  - Now turn on the fluorescence microscope [1]. Select FLM (F-L-M) tab, set reflected light imaging conditions and select play [2]. Then align the reflected light image with the SEM (sem) image. [3].

    - Talent pressing a button on a computer screen to switch on the microscope.

    - SCREEN: 2.8.2-(1).mp4	00:00-00:12

    - SCREEN: 2.8.2-(1).mp4	00:13-00:21, 00:46-00:53

  - Zoom into the image and click the rectangle option to draw focus of the ion beam [1]. Press Start Patterning in Display 2 (two) to mill a small pattern using the focused ion beam [1].

    - SCREEN: 2.9.1-(1).mp4	00:00-00:13

    - SCREEN: 2.9.1-(1).mp4	00:19-00:42, 00:58-00:59

  - Confirm that the milled pattern is visible in both the focused ion beam and fluorescent microscope views [1-TXT].

    - SCREEN: 2.10.1-(1).mp4	00:35-00:50
TXT: If necessary, repeat alignment

  - In the light microscopy software, move to the Localization tab and activate the appropriate conditions to set up the fluorescent channels [1]. Assign each fluorophore to a separate channel and include an additional channel for brightfield [2].

    - SCREEN: 2.11.1-(1).mp4	00:00-00:19

    - SCREEN: 2.12.1-(1).mp4	00:00-00:12

  - Adjust the LED (L-E-D) power and exposure time for each fluorescent channel to identify the target of interest [1-TXT]. Once the parameters are satisfactory, acquire a pre-milling fluorescent z-stack by stepping through the axial direction [2]. Then, mark the target of interest as a region of interest [3].

    - SCREEN: 2.12.1.mp4	00:14-00:30, 00:55-01:04
TXT: Use 1.2 s and 30 mW when imaging SiR-tubulin

    - SCREEN: 2.12.2-2.12.3-(1).mp4	00:00-00:10, 00:25-00:30, 01:0-01:30

    - SCREEN: 2.12.2-2.12.3-(1).mp4	01:40-01:45

  - Switch to the FIB-SEM software and draw a lamella at the region of interest [1].

    - SCREEN: 2.13.1-(1).mp4	00:00-00:10

  - Perform rough milling to remove the bulk material approximately 3 micrometers above and below the region of interest while keeping the fluorescence microscope on to monitor signal changes [1].

    - SCREEN: 2.14.1-(1).mp4 	00:00-00:04
and   2.14.1-(1)_complement.mp4	00:00-00:04	

Video Editor: Please play the files one after the other

  - Open the Python-based interferometric toolkit from the terminal to monitor fluorescent signal during lamella thinning [2]. Thin the ROI (R-O-I) to approximately 2 µM (micrometers) using the cleaning cross-section mode [1].

    - SCREEN: 2.16and17-(1).mp4	00:02-00:06,

    - SCREEN: 2.15.1_correct-(1).mp4	01:10-01:26

  - For a target with an axial extent of over 300 nanometers, in the Python GUI (G-U-I), select the target in the fluorescence micrograph [1]. The software will automatically begin plotting brightness as a function of time with each new frame [2].

    - SCREEN: 2.16and17-(1).mp4		00:25- 00:40

    - SCREEN: 2.16and17-(1).mp4		00:41-00:58

  - Monitor fluorescence intensity through the Python toolkit [1]. Mill the sample using cleaning cross-section mode to remove any remaining support film from below the region of interest [2-TXT].

    - SCREEN: 2.16and17-(1).mp4		01:00-01:14
 Video Editor: Here, emphasize the panel on the right with the red curve

    - SCREEN: 2.16and17-(1).mp4		01:25-01:35
TXT: Fluorescence is enhanced when support film is removed
Video Editor: Here, emphasize the panel on the left with the red box

  - Observe for sudden drops in fluorescence intensity as cues to switch milling direction [1]. If a sharp decrease is observed when milling from the bottom up, stop and begin milling from the top down [2-TXT].

    - SCREEN: 2.16and17-(1).mp4		01:35-01:42

    - SCREEN: 2.15.1_correct-(1).mp4	02:45-03:00
TXT: If structure is removed too quickly, reduce milling current to slow down target removal

  - After completing top-down milling, finalize the lamella by milling from the bottom up in cleaning cross-section mode until the desired thickness is reached [1]. Then capture a post-milling fluorescence image of the lamella [2].

    - SCREEN: 2.19.1-(1).mp4	00:00-00:30
Videographer’s Note: I also recorded the talent removing the rod, in case it's needed

    - SCREEN: 2.19.2_final_micrograph-(1).mp4		00:07-00:17
Fluorescence image acquisition showing the lamella post milling.

  - Transfer the sample to a CryoTEM (Cry-o-Tem) [1].  collect an atlas overview of the grid at 182 X to cover the entire grid [2-TXT].

    - Shot of the sample in a CryoTEM.

    - SCREEN: Atlas overview of the entire grid captured at 182 X in the CryoTEM interface. TXT: Use Tomo5 software
NOTE: Shot deleted by authors

  - Then acquire a montage series of images of the final lamella at 0 degree tilt at appropriate intermediate magnifications [1-TXT].

    - SCREEN: Montage imaging window displaying stitched images of the lamella at intermediate magnification and 0 degree tilt. TXT: Magnifications: 6500X to 11,000X
NOTE: Shot deleted by authors

  - Load the post-milling multi-channel fluorescence image and the CryoTEM projection image into the custom projective transformation software for region of interest registration [1]. Now select 8 to 10 pairs of reference points that are visible in both fluorescence and electron microscopy images to compute projective transformation [2]. Use the overlaid fluorescence and transmission electron microscopy image as a guide to select an imaging area within the camera frame size for high-magnification tilt-series collection [3].

    - SCREEN: 2.22-(1).mp4	00:04-00:28, 00:46

    - SCREEN: 2.22-(1).mp4	01:07-01:21, 01:50-01:53,02:19-02:20
AND

SCREEN: 2.22-(1).mp4	03:08-03:12, 03:55-04:00
Video Editor: Please play both files side by side in a split screen

    - SCREEN: 2.22-(1).mp4	04:30-04:36

  - Then select appropriate magnification and acquisition parameters for the tilt-series, depending on the biological objective [1].

    - SCREEN: 2.23.1-(1).tif

# Results

- Results

  - Differentiated macrophage cells labeled with SiR (S-I-R) -tubulin displayed a single bright fluorescent punctum approximately 1 micrometer in diameter, corresponding to the MTOC (M-T-O-C) [1-TXT], along with fibril-like structures radiating towards the cell periphery, consistent with the microtubule network [2].

    - LAB MEDIA: Figure 1A (left image)  Video editor: Highlight the single bright red dot in each cell, marked by yellow  arrowheads 
TXT: MTOC : Microtubule Organizing Center

    - LAB MEDIA: Figure 1A (right image)  Video editor: Highlight the red fibril-like strands marked by yellow arrowheads

  - As the sample was milled to approximately 800 nanometers in thickness, the singular fluorescent MTOC punctum resolved into two distinct speckles, each approximately 700 nanometers in diameter [1].

    - LAB MEDIA: Figure 1B Video Editor: Please sequentially show the columns “pre-milling” to “800 nm” then highlight the 800 nm column images

  - Fluorescence intensity increased after removal of non-fluorescent bulk material and the absorptive support film [1], followed by a sharp decrease as the lamella was further thinned from 2 micrometers to below 200 nanometers [2].

    - LAB MEDIA: Figure 1 C. Video editor: Highlight the middle region of the timeline where the curve rises, marked by the label “Enhancement due to the removal of the absorptive support film”.

    - LAB MEDIA: Figure 1D. Video editor: Emphasize the downward trend in the red lines labeled 180 nm and the lighter red line labeled 500 nm.

  - Correlated fluorescence and electron microscopy imaging enabled precise targeting of structures during milling [1]. Reconstructed tomograms and segmented models reveal that the centrioles are composed of microtubule triplets [2].

    - LAB MEDIA: Figure 2A(a–c). Video editor: Highlight the areas pointed at by the yellow arrows

    - LAB MEDIA: Figure 2A (d). Video editor: Highlight the areas pointed at by the yellow arrows.

  - In cases where only one fluorescent speckle was retained in the final lamella, a single centriole was visualized in the corresponding tomogram [1].

    - LAB MEDIA: Figure 2B Video editor: Highlight the areas pointed at by the yellow arrows and the inset

Pronunciation Guide:

- Cryogenic
Pronunciation link: https://www.merriam-webster.com/dictionary/cryogenic
IPA: /ˌkraɪ.əˈdʒɛn.ɪk/
Phonetic Spelling: kry·uh·jen·ik

- Coincident
Pronunciation link: https://www.merriam-webster.com/dictionary/coincident
IPA: /koʊˈɪn.sɪ.dənt/
Phonetic Spelling: koh·in·suh·duhnt

- Fluorescence
Pronunciation link: https://www.merriam-webster.com/dictionary/fluorescence
IPA: /ˌflʊrˈɛs.əns/
Phonetic Spelling: floor·eh·suhns

- Electron
Pronunciation link: https://www.merriam-webster.com/dictionary/electron
IPA: /ɪˈlɛk.trɑn/
Phonetic Spelling: ih·lek·tron

- Ion
Pronunciation link: https://www.merriam-webster.com/dictionary/ion
IPA: /ˈaɪ.ɑn/
Phonetic Spelling: eye·on

- Microscope
Pronunciation link: https://www.merriam-webster.com/dictionary/microscope
IPA: /ˈmaɪ.krəˌskoʊp/
Phonetic Spelling: my·kruh·skohp

- Lamella
Pronunciation link: https://www.merriam-webster.com/dictionary/lamella
IPA: /ləˈmɛl.ə/
Phonetic Spelling: luh·mel·uh

- Autogrid
Pronunciation link: No confirmed link found
IPA: /ˈɔː.t̬oʊˌɡrɪd/
Phonetic Spelling: aw·toh·grid

- Cryo-FIB-SEM
Pronunciation link: No confirmed link found
IPA: /ˈkraɪ.oʊ fɪb sɛm/
Phonetic Spelling: kry·oh fib sem

- Correlative
Pronunciation link: https://www.merriam-webster.com/dictionary/correlative
IPA: /kəˈrɛl.ə.tɪv/
Phonetic Spelling: kuh·rel·uh·tiv

- Microscopy
Pronunciation link: https://www.merriam-webster.com/dictionary/microscopy
IPA: /maɪˈkrɑs.kə.pi/
Phonetic Spelling: my·kros·kuh·pee

- Fluorophore
Pronunciation link: https://www.merriam-webster.com/dictionary/fluorophore
IPA: /ˈflʊr.əˌfɔr/
Phonetic Spelling: floor·uh·for

- Z-stack
Pronunciation link: No confirmed link found
IPA: /ˈziː stæk/
Phonetic Spelling: zee·stak

- Micrometer
Pronunciation link: https://www.merriam-webster.com/dictionary/micrometer
IPA: /maɪˈkrɑm.ə.t̬ɚ/
Phonetic Spelling: my·krah·muh·ter

- Nanometer
Pronunciation link: https://www.merriam-webster.com/dictionary/nanometer
IPA: /ˈnæn.oʊˌmiː.t̬ɚ/
Phonetic Spelling: nan·oh·mee·ter

- Interferometric
Pronunciation link: https://www.merriam-webster.com/dictionary/interferometric
IPA: /ˌɪn.t̬ɚ.fɪr.oʊˈmɛt.rɪk/
Phonetic Spelling: in·ter·feer·oh·met·rik

- CryoTEM
Pronunciation link: No confirmed link found
IPA: /ˈkraɪ.oʊ tɛm/
Phonetic Spelling: kry·oh tem

- Tomography
Pronunciation link: https://www.merriam-webster.com/dictionary/tomography
IPA: /təˈmɑɡ.rə.fi/
Phonetic Spelling: tuh·mog·ruh·fee

- Tomogram
Pronunciation link: https://www.merriam-webster.com/dictionary/tomogram
IPA: /ˈtɑm.əˌɡræm/
Phonetic Spelling: tom·uh·gram

- Centriole
Pronunciation link: https://www.merriam-webster.com/dictionary/centriole
IPA: /ˈsɛn.triˌoʊl/
Phonetic Spelling: sen·tree·ohl

- Microtubule
Pronunciation link: https://www.merriam-webster.com/dictionary/microtubule
IPA: /ˌmaɪ.kroʊˈtuː.bjuːl/
Phonetic Spelling: my·kroh·too·byool

- Macrophage
Pronunciation link: https://www.merriam-webster.com/dictionary/macrophage
IPA: /ˈmæk.rəˌfeɪdʒ/
Phonetic Spelling: mak·ruh·fayj

- SiR-tubulin
Pronunciation link: No confirmed link found
IPA: /ˈɛs.aɪ.ɑr ˈtuː.bjʊ.lɪn/
Phonetic Spelling: ess·eye·ar too·byuh·lin

- MTOC
Pronunciation link: No confirmed link found
IPA: /ˌɛm.tiː.oʊˈsiː/
Phonetic Spelling: em·tee·oh·see

- Fluorophore
Pronunciation link: https://www.merriam-webster.com/dictionary/fluorophore
IPA: /ˈflʊr.əˌfɔr/
Phonetic Spelling: floor·uh·for

- Refractive
Pronunciation link: https://www.merriam-webster.com/dictionary/refractive
IPA: /rɪˈfræk.tɪv/
Phonetic Spelling: rih·frak·tiv

- Diffraction
Pronunciation link: https://www.merriam-webster.com/dictionary/diffraction
IPA: /dɪˈfræk.ʃən/
Phonetic Spelling: dih·frak·shuhn

- Registration
Pronunciation link: https://www.merriam-webster.com/dictionary/registration
IPA: /ˌrɛdʒ.ɪˈstreɪ.ʃən/
Phonetic Spelling: rej·ih·stray·shuhn

- Projective
Pronunciation link: https://www.merriam-webster.com/dictionary/projective
IPA: /prəˈdʒɛk.tɪv/
Phonetic Spelling: pruh·jek·tiv

- Transformation
Pronunciation link: https://www.merriam-webster.com/dictionary/transformation
IPA: /ˌtræns.fɚˈmeɪ.ʃən/
Phonetic Spelling: trans·fer·may·shuhn