# Submission ID #: 68238

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=20812638

# Title: Workflow Using a Cryogenic Coincident Fluorescence, Electron, and Ion Beam Microscope for Targeted Milling of Cells

# Authors and Affiliations:

# Jue Wang1,2, Anthony V. Sica2, Grant J. Jensen3, Peter D. Dahlberg2,4

# 1Division of Biology and Biological Engineering, California Institute of Technology

# 2SLAC National Accelerator Laboratory

# 3Department of Chemistry and Biochemistry, Brigham Young University

# 4Department of Photon Science and Structural Biology, Stanford University

# Corresponding Authors:

Peter D. Dahlberg 		pdahlb@slac.stanford.edu

# Email Addresses for All Authors:

Jue Wang 			jpwang@slac.stanford.edu

Anthony V. Sica 		asica1@slac.stanford.edu

Grant J. Jensen 		grant_jensen@byu.edu

Peter D. Dahlberg 		pdahlb@slac.stanford.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Yes  Yes

Authors: Please create scope videos of the shots labeled as SCOPE and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=20812638

Cryo-FM-FIB-SEM system

SCOPE: 2.3.2, 2.4.2, 2.5.1- 2.5.2, 2.6.1-2.6.2, 2.7.1-2.7.2, 2.8.1- 2.8.2, 2.9.1, 2.10.1, 2.11.1-2.11.2, 2.12.1- 2.12.3, 2.13.1, 2.14.1, 2.15.1-2.15.2, 2.16.1-2.16.2, 2.17.1-2.17.2, 2.18.1-2.18.2, 2.19.1-2.19.2

CryoTEM

SCOPE: 2.20.1- 2.20.2, 2.21.1, 2.22.1-2.22.3, 2.23.1

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=20812638

Current Protocol Length

Number of Steps:  23

Number of Shots:  41

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Jue Wang: My research develops optical microscopy methods to preserve rare cellular structures in cryogenic lamella during focused ion beam milling, enabling their detailed analysis by cryogenic electron tomography to uncover native biological organization.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.1

What are the most recent developments in your field of research?

  - Jue Wang: Optical microscopy has recently been integrated into cryoFIB-SEM systems, but rarely at the same focal position as the FIB. This necessitates registration-based guidance approaches to preserve fluorescently labelled structures in final lamella.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:2.19

What are the current experimental challenges?

  - Jue Wang: Accuracy of Fluorescence-based lamella targeting is limited by registration approaches, which suffer from registration error, diffraction-limited resolution, refractive index mismatch, and milling-induced motion.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What significant findings have you established in your field?

  - Peter Dahlberg: Simultaneous optical microscopy and FIB milling removes many of these sources of guidance error by enabling real-time feedback for precise milling that preserves fluorescently labeled structures for cryo-electron tomography within the final lamella.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What new scientific questions have your results paved the way for?

  - Jue Wang: Our research makes it possible to see cellular structures that were previously challenging to directly visualize by cryo-ET, shedding light on molecular mechanisms behind a wide spectrum of key macromolecular machines.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.4

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

- High-Precision Targeting for Cryo-Correlative Light and Electron Microscopy

Demonstrator: Jue Wang

  - To begin, manually cool the transfer station using liquid nitrogen once the system temperature nears cryogenic levels [1].

    - WIDE: Talent pouring liquid nitrogen into the transfer station while checking the temperature gauge.

  - Load one clipped Autogrid into the sample cassette designed for the cryogenic light microscopy system [1]. Ensure that the Autogrid is positioned between the glued coverslip and the spacer [2].

    - Talent placing a clipped Autogrid into the cassette.

    - Shot of the Autogrid correctly seated between the glued coverslip and spacer.

  - Tighten the ring screw to close the cassette [1], and insert it into the slot on the transfer station [2].

    - Talent tightening the ring screw on the cassette.

    - SCREEN: Talent inserting the closed cassette into the transfer station slot.

  - Attach the transfer station to the cryogenic light microscopy system [1]. Use the internal chamber camera to guide the sample cassette onto the sample stage [2].

    - Talent connecting the transfer station to the system interface.

    - SCREEN: Display the internal chamber camera feed guiding the sample cassette accurately onto the sample stage.

  - In the light microscopy software, move the sample stage from the Loading position to the predefined three-beam position [1]. Then, in the FIB-SEM (FIB-Sem) software, click on the top-left window and set the scanning electron microscope to a magnification of 100X (Hundred-Ex) [2].

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=20812638

    - SCREEN: Show the light microscopy software, selecting the option to move the stage from Loading to the 3-beam position.

    - SCREEN: Display the FIB-SEM software interface with the top-left window selected and set the magnification to 100 times.

  - In the Alignment tab of the light microscopy software, click on the “z” and select the “play” function to acquire a SEM atlas of the grid using previously set parameters [1]. Identify a square area with broken carbon film and double-click on it to move the sample stage to that location [2].

    - SCREEN: Navigate to the Alignment tab and select “z”. Under SEM setting, click “play” to display the atlas image of the grid.

    - SCREEN: Double click on a square with broken carbon film, prompting the sample stage to move.

  - Find the coincident FIB-SEM position by adjusting the axial stage position using “+z” or “-z.”

    - SCREEN: Return to the FIB-SEM software interface and activate the FIB and acquire an image. Adjust the magnification using the drop down manu on the top left corner.

    - SCREEN: Select “play” under FIB Setting in the light microscopy software to display the FIB image. Set the step size with the slide bar on the left. Then adjust the stage “z” position with set step size by selecting “+z” and “-z”. Repeat until the horizontal line present in the ion and electron images coincide with the same feature.

SCREEN: Select the SEM tab in the light microscopy software.

    - Use the “X” and “Y” buttons to align a common feature between the SEM and ion images.

  - Now turn on the fluorescence microscope [1].  Adjust the x, y, and z coordinates of the objective stage to match the SEM view [2].

    - Talent switching on the fluorescence microscope.

    - SCOPE/SCREEN: Select “FLM” tab. Set reflected light imaging conditions and select “play.” Use the “X,” “Y” and “Z” buttons with selected step size to align the reflected light image with the SEM image. Note there is an initial large axial offset of ~50 microns.

  - Mill a small pattern using the focused ion beam, such as a small rectangle or circle [1].

    - SCREEN: Show the FIB software interface milling a small circular or rectangular pattern on the carbon film.

  - Confirm that the milled pattern is visible in both the focused ion beam and fluorescent microscope views [1-TXT].

    - SCREEN: Side-by-side comparison of the milled pattern as seen in both the FIB and fluorescence microscope views. TXT: If necessary, repeat alignment step 2.8.2

  - In the light microscopy software, move to the Localization tab and activate the appropriate excitation and filtering conditions to set up the fluorescent channels [1]. Assign each fluorophore to a separate channel and include an additional channel for transmitted light or brightfield [2].

    - SCREEN: Navigate to the Localization tab and activate excitation and filtering settings for reflected light.

    - SCREEN: Assign multiple fluorophores to separate channels.

  - Adjust the LED (L-E-D) power and exposure time for each fluorescent channel to identify the target of interest [1-TXT]. Once the parameters are satisfactory, acquire a pre-milling fluorescent z-stack by stepping through the axial direction [2]. Then, mark the target of interest as a region of interest [3].

    - SCREEN: Show LED and exposure settings being adjusted for each fluorescent channel. TXT: Use 1.2 s and 30 mW when imaging SiR-tubulin

    - SCREEN: Fluorescent z-stack is being captured along the z-axis.

    - SCREEN: A region of interest is being marked on the fluorescent image.

  - Switch to the FIB-SEM software and draw a lamella at the region of interest [1].

    - SCREEN: Lamella outline being drawn on the FIB-SEM interface over the selected region.

  - Perform rough milling to remove the bulk material approximately 3 micrometers above and below the region of interest while keeping the fluorescence microscope on to monitor signal changes [1].

    - SCREEN: Talent initiating rough milling on the FIB-SEM system 
AND
SCREEN: Fluorescence signal changes are being seen.
Video Editor: Please play both shots side by side

  - Thin the ROI (R-O-I) to approximately 2 µM (micrometers) using the cleaning cross-section mode [1]. Open the Python-based interferometric toolkit from the terminal to monitor fluorescent signal during lamella thinning [2].

    - SCREEN: The cleaning cross section mode is being used and the ROI is being thinned to 2 µm.

    - SCREEN: Terminal command launching the Python toolkit and displaying real-time fluorescence plots.

  - For a target with an axial extent of over 300 nanometers, in the Python GUI (G-U-I), select the target in the fluorescence micrograph [1]. The software will automatically begin plotting brightness as a function of time with each new frame [2].

    - SCREEN: The region of interest within the Python GUI is being selected in the fluorescence micrograph.

    - SCREEN: Brightness plot as a function of time is being seen.

  - Monitor fluorescence intensity through the Python toolkit [1]. Mill the sample using cleaning cross-section mode to remove any remaining support film from below the region of interest [2-TXT].

    - SCREEN: Florescent intensity is being seen.

    - SCREEN: Sample is being milled from below the ROI. TXT: Fluorescence is enhanced when support film is removed

  - Observe for sudden drops in fluorescence intensity as cues to switch milling direction [1]. If milling from the bottom up and a sharp decrease is observed, stop and begin milling from the top down [2-TXT].

    - SCREEN: Highlight a steep decline in the fluorescence signal on the Python plot.

    - SCREEN: Bottom-up milling is being halted and reoriented to top-down milling in the FIB-SEM interface. TXT: If structure is removed too quickly, reduce milling current to slow down target removal

  - After completing top-down milling, finalize the lamella by milling from the bottom up in cleaning cross-section mode until the desired thickness is reached [1]. Then capture a post-milling fluorescence image of the lamella [2].

    - SCREEN: Milling process continuing in cleaning cross-section mode from the top or bottom up, with progress bar or depth scale indicating thinning to target thickness.

    - SCREEN: Fluorescence image acquisition showing the lamella post milling.

  - After transferring the sample to a CryoTEM (Cry-o-Tem) [1], collect an atlas overview of the grid at 182 X to cover the entire grid [2-TXT].

    - Shot of the sample in a CryoTEM.

    - SCREEN: Atlas overview of the entire grid captured at 182 X in the CryoTEM interface. TXT: Use Tomo5 software

  - Then acquire a montage series of images of the final lamella at 0 degree tilt at appropriate intermediate magnifications [1-TXT].

    - SCREEN: Montage imaging window displaying stitched images of the lamella at intermediate magnification and 0 degree tilt. TXT: Magnifications: 6500X to 11,000X

  - Load the post-milling multi-channel fluorescence image and the CryoTEM projection image into the custom projective transformation software for region of interest registration. Now select 8 to 10 pairs of reference points that are visible in both fluorescence and electron microscopy images to compute projective transformation [2]. Use the overlaid fluorescence and transmission electron microscopy image as a guide to select an imaging area within the camera frame size for high-magnification tilt-series collection [3].

    - SCREEN: Importing and thresholding fluorescence image, TEM atlas at low magnification, and an enlarged lamella atlas for identifying ROI into the software interface.

    - SCREEN: Select 6 to 10 pairs of reference points that are visible in both fluorescence and TEM atlas at low magnification to compute the projective transformation. Use the drop-down menu on the top to adjust the view. Move the image using the scroll bars on the left and the bottom. At this step, we recommend using features such as four edges of the lamella and other clearly distinguishable features located along the lamella edge. Other features, such as holes or broken areas on the EM grids, can also be used. Reference points are being marked across both image types, with transformation grid overlay updating dynamically.

    - SCREEN: Select again 6 to 10 pairs of reference points that are visible in both the low-magnification TEM atlas and the enlarged lamella atlas to correlate fluorescence signal with cellular features. We recommend using features, such as small ice crystals or large membrane structures visible at this magnification under TEM.

    - SCREEN: Display of overlaid image guiding the selection of a region within the camera's field of view.

  - Then select appropriate magnification and acquisition parameters for the tilt-series, depending on the biological objective [1].

    - SCREEN: User selecting a tilt-series profile in the acquisition interface, entering parameters.

# Results

- Results

  - Differentiated macrophage cells labeled with SiR (S-I-R) -tubulin displayed a single bright fluorescent punctum approximately 1 micrometer in diameter, corresponding to the MTOC (M-T-O-C) [1-TXT], along with fibril-like structures radiating towards the cell periphery, consistent with the microtubule network [2].

    - LAB MEDIA: Figure 1A (left image)  Video editor: Highlight the single bright red dot in each cell, marked by yellow  arrowheads 
TXT: MTOC : Microtubule Organizing Center

    - LAB MEDIA: Figure 1A (right image)  Video editor: Highlight the red fibril-like strands marked by yellow arrowheads

  - As the sample was milled to approximately 800 nanometers in thickness, the singular fluorescent MTOC punctum resolved into two distinct speckles, each approximately 700 nanometers in diameter [1].

    - LAB MEDIA: Figure 1B Video Editor: Please sequentially show the columns “pre-milling” to “800 nm” then highlight the 800 nm column images

  - Fluorescence intensity increased after removal of non-fluorescent bulk material and the absorptive support film [1], followed by a sharp decrease as the lamella was further thinned from 2 micrometers to below 200 nanometers [2].

    - LAB MEDIA: Figure 1 C. Video editor: Highlight the middle region of the timeline where the curve rises, marked by the label “Enhancement due to the removal of the absorptive support film”.

    - LAB MEDIA: Figure 1D. Video editor: Emphasize the downward trend in the red lines labeled 180 nm and the lighter red line labeled 500 nm.

  - Correlated fluorescence and electron microscopy imaging enabled precise targeting of structures during milling [1]. Reconstructed tomograms and segmented models reveal that the centrioles are composed of microtubule triplets [2].

    - LAB MEDIA: Figure 2A(a–c). Video editor: Highlight the areas pointed at by the yellow arrows

    - LAB MEDIA: Figure 2A (d). Video editor: Highlight the areas pointed at by the yellow arrows.

  - In cases where only one fluorescent speckle was retained in the final lamella, a single centriole was visualized in the corresponding tomogram [1].

    - LAB MEDIA: Figure 2B Video editor: Highlight the areas pointed at by the yellow arrows and the inset