# Submission ID #: 68273

# Scriptwriter Name: Deepthi Das

# Project Page Link: https://review.jove.com/files_upload.php?src=20823313

# Title: High-throughput Serial Tomography Using an Automated Sample-Changing Robot

# Authors and Affiliations: Shashidhara Marathe1, Malte Storm1,2, Kaz Wanelik1, Huw Shorthouse1, Nickie Griffin1, Anjali Goswami3, Gavin R. Broad3, Ben Price3, Oliver Kippax-Chui4,5, Christoph Rau1

# 1Diamond Light Source Ltd., Harwell Science and Innovation Campus

# 2Present address: Helmholtz-Zentrum Hereon GmbH, Max-Planck-Straße

# 3Natural History Museum, London

# 4Department of Earth Science and Engineering, Royal School of Mines, Imperial College

# 5Institute of Zoology, Zoological Society of London

# Corresponding Authors:

Shashidhara Marathe   	shashidhara.marathe@diamond.ac.uk

# Email Addresses for All Authors:

# a.goswami@nhm.ac.uk

b.price@nhm.ac.uk

christoph.rau@diamond.ac.uk

g.broad@nhm.ac.uk

huw.shorthouse@diamond.ac.uk

kaz.wanelik@diamond.ac.uk

malte.storm@hereon.de

nickie.griffin@diamond.ac.uk

o.kippax-chui23@imperial.ac.uk

shashidhara.marathe@diamond.ac.uk

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

Please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Yes

If Yes, how far apart are the locations?  

Filming will be in the same building ( I13-2 beamline of Diamond Light Source), however we will shoot in the different rooms experiment control room, experiment hutch and I13 chemical lab all are bout 20m from each other.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Yes

Current Protocol Length

Number of Steps:  15

Number of Shots:  48

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context. The answers have been revised for length, clarity, and consistency with journal style guidelines.

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Christoph Rau: This protocol describes automated sample exchange using a robot in I13-2, X-ray micro-tomography beamline of Diamond Light Source synchrotron facility.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Shashidhara Marathe: Manual sample exchange and alignment often exceed tomography scan time; this protocol automates them, improving beamtime efficiency and experimental throughput.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Christoph Rau: This protocol enables preparation of robot-compatible samples, allowing high-throughput acquisition of statistically significant datasets for data-intensive scientific studies.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the most important consideration or challenge when performing this protocol.

  - Shashidhara Marathe: Sample preparation and sample alignment are two critical steps for the automatic sample changer robot to work correctly and effectively.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Shashidhara Marathe: Future studies can develop larger grippers for bigger samples and improve designs by incorporating side-approach mechanisms.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

Videographer: Please capture all testimonial shots in a wide-angle format with sufficient headspace, as the final videos will be rendered in a 1:1 aspect ratio. Testimonial statements will be presented live by the authors, sharing their spontaneous perspectives.

- Testimonial statements will not appear in the video but may be featured in the journal’s promotional materials.

- Provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author delivering the testimonial.

- Please answer the testimonial question live during the shoot, speaking naturally and in your own words in complete sentences.

What problem in your field of research does your JoVE article address?

1.6    Prof. Christoph Rau, Principal Beamline Scientist, Diamond Light Source Ltd.: (authors will present their testimonial statements live)

1.6.1  INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

What is the key takeaway you hope viewers remember after watching your video article?

  - Prof. Christoph Rau, Principal Beamline Scientist, Diamond Light Source Ltd.​: (authors will   present their testimonial statements live) 

1.7.1  INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

# Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Sample Preparation for Robot Experiments

Demonstrators: Shashidhara Marathe, Christoph Rau

  - To begin, ensure all the paperwork and safety training needed for the experiment are complete [1]. Obtain insect samples from the Natural History Museum, mounted using a CryoPin on SPINE (spine) -compatible sample bases equipped with a Data Matrix barcode [2-TXT]. Document the Data Matrix code associated with each sample [3]. Mount the CryoCap on the CryoPin [4]. Verify that the mounted sample assembly is properly secured [5].

    - WIDE: Talent checking the permissions and safety training certificates.

    - CU: Shot of the insect samples mounted with a CryoPin on SPINE-compatible sample bases equipped with a Data Matrix barcode. TXT: SPINE - Structural Proteomics in Europe

    - SCREEN: Talent recording the Data Matrix code for a mounted sample.

    - Talent mounting the CryoCap on the CryoPin

    - Talent verifying that the assembly is properly positioned and secured.

  - Check whether the mounted sample fits inside one of the allowed size envelopes [1]. Ensure that the mounted configuration is centered on the sample base [2]. Confirm that the total height of the sample, mount, and base does not exceed 27 millimeters [3].

    - Talent checking the mounted sample against the allowed size envelopes for robot handling.

    - Talent centering the mounted configuration on the sample base.

    - Talent checking the total height of the sample, mount, and base.

  - When mounting samples on SPINE sample pins, position the sample on top of the pin with no part obscured by the pin shaft [1]. Mount the sample vertically and centrally on the pin base [2].

    - Talent positioning the sample on top of the sample pin.

    - Talent adjusting the sample to a vertical and central position on the pin base.

  - Fix the pin securely and vertically onto the sample base [1]. Use a securing method or reliable adhesive to prevent detachment [2].

    - Talent fixing the mounting element vertically onto the sample base.

    - Talent applying an adhesive or securing method to the mounting element.

  - Check that all samples are positioned at a consistent height on the sample pin [1]. Generate a samples.csv (samples dot c-s-v) file that links each sample name to its corresponding sample-base Data Matrix code and includes the parameters required for running tomography scans [2].

    - Talent comparing sample heights across multiple mounted samples.

    - SCREEN: Create the samples.csv file with the sample data and metadata.

- Aligning the X-ray Imaging System for Data Acquisition

Demonstrator: Click here to enter name of demonstrator(s)

  - Shift the full spectrum or filter the X-ray spectrum to achieve the required contrast [1]. Optimize the camera to sample distance based on the X-ray spectrum and the degree of phase contrast required for the sample [2]. Select the microscope magnification based on the sample size and the required field of view [3]. Select the corresponding microscope objective [4]. Align the camera horizontal position along the x-axis of the laboratory with respect to the X-ray beam [5].

    - SCREEN: The full spectrum being shifted or the X-ray spectrum being filtered to achieve the required contrast

    - Shot of the stages moving in the experimental hutch the distance between the camera and the sample is being adjusted.

    - SCREEN: The microscope magnification being selected based on the sample size and the required field of view.

    - SCREEN: The corresponding microscope objective is being selected on the linux workstation.

    - Shot of the camera being aligned horizontally relative to the X-ray beam.

  - Check the samples and sample mounting thoroughly at the start of the experimental session for damage or loose mounting and inform the users [1]. Mount the samples on the respective puck and puck slots identified in the samples.csvFs file [2]. Perform a dry run of the robot sample exchange procedure on a few samples [3].

    - Talent inspecting multiple mounted samples at the start of the session and noting the details of any damage.

    - Talent placing samples into the specified puck and puck slots.

    - Talent initiating a dry run of the robot sample exchange procedure.

  - Use a sample base with a sample pin for aligning the axis of rotation of the rotation stage to the center of the detector in the horizontal direction [1]. Move the top of the pin to the center of the X-ray camera [2]. Align the telescope camera, set at 90 degrees to the beam, to the center of the field of view [3]. Determine the coordinates of this position in telescope camera pixel coordinates [4]. Update the sample target position in the script that handles sample alignment [5].

    - Talent placing an empty sample pin on the rotation stage for alignment.

    - Talent moving the top of the pin to the center of the X-ray camera.

    - Talent aligning the telescope camera to the center of the field of view.

    - SCREEN: Show the telescope camera view while determining the pixel coordinates of the centered position.

    - SCREEN: Update the target position values in the sample alignment script.

  - Mount a representative sample [1]. Run the automatic sample alignment script to place the sample in the center of rotation on the telescope camera [2]. Check whether the sample region of interest is centered in the field of view of the X-ray data acquisition camera [3].

    - Talent mounting a representative sample onto the stage.

    - SCREEN: Run the automatic sample alignment script.

    - SCREEN: Show the X-ray data acquisition camera view with the sample region of interest centered in the field of view.

  - Repeat the previously demonstrated alignment procedure if the region of interest is not centered in the field of view [1]. Change the target position coordinates to refine the sample alignment [2]. Update the camera exposure time and the number of projections to acquire in the samples.csv file [3].

    - Representative shot of the talent repeating the sample alignment procedure.

    - SCREEN: Adjust the target position coordinates in the alignment script.

    - SCREEN: Edit the exposure time and number of projections in the samples.csv file.

- Tomography Data Acquisition and Reconstruction

Demonstrator: Shashidhara Marathe

  - Mount all samples on the correct puck and pin positions as instructed in the samples.csv file [1]. Use the Generic Data Acquisition or GDA (G-D-A) workbench for data collection [2]. Issue the command tomomatic.load (tomomatic dot load) in the GDA terminal to load the file [3]. Then, issue tomomatic.scan (tomomatic dot scan) to run the command [4].

    - Talent mounting all samples onto the designated puck and pin positions.

    - SCREEN: Show the GDA workbench interface ready for data collection.

    - SCREEN: Show the tomomatic object in the GDA workbench.

    - SCREEN: Enter the tomomatic.load (/dls/i13/data/year/… samples.csv) command in the terminal.

    - SCREEN: Enter the tomomatic.scan () command in the terminal.

  - Monitor the robot transferring the sample from the loading tray to the laser barcode reader [1]. Confirm that the barcode is read [2]. Ensure that the robot moves the sample to the tomography stage and mounts it correctly [3].

    - Talent monitoring the robot as it transfers the sample from the loading tray to the barcode reader.

    - Shot of the sample barcode being read by the barcode reader.

    - Talent monitoring the robot as it mounts the sample onto the tomography stage.

  - Monitor the stages moving to the scanning position [1]. Then, issue the tomomatic script to execute the autoCentrePin (auto-centre-pin) function to place the sample into the Center of Rotation and the TomoFlyscan (tomo-fly-scan) function for the tomography measurement to collect the images for image normalization [2].

    - Talent observing the stages move to the scanning position.

    - SCREEN: The tomomatic script is running to execute the autoCentrePin function and the TomoFlyscan function is being executed in continuation.

  - Collect projection images with the sample in the beam at the sample rotation angle intervals indicated in the samples.csv file [1]. Acquire 3001 projection images [2].

    - SCREEN: Show projection image acquisition with the sample in the beam.

    - SCREEN: Show the image acquisition progressing

  - Access the saved data by browsing to the appropriate file system and folder [1]. Use the Savu reconstruction pipeline for generating 2-dimensional slices from the projection images [2]. Access and visualize the reconstructed images using 3-dimensional rendering software such as AVIZO or ImageJ [3].

    - Talent browsing the file system to open the data folder.

    - SCREEN: Run the Savu reconstruction pipeline to generate 2-dimensional slices.

    - SCREEN: Open the reconstructed images in 3-dimensional rendering software and visualize the dataset.

# Results

- Results

  - Raw projection images of the insect were acquired using 1.25 times [1] and 4 times magnification during the tomography experiment [2].

    - LAB MEDIA: Figure 4 Video editor: Highlight the top row images to display the full insect

    - LAB MEDIA: Figure 4. Video editor: Highlight the bottom row images

  - Projection images were acquired at 0 degrees [1], 90 degrees [2], and 180 degrees [3] for both magnifications.

    - LAB MEDIA: Figure 4. Video editor: Highlight the label 0° and the panels A and D

    - LAB MEDIA: Figure 4. Video editor: Highlight the label 90° and the panels B and E

    - LAB MEDIA: Figure 4. Video editor: Highlight the label 180° and the panels C and F

  - Horizontal or transverse slices from the tomography reconstruction were well-centered within the field of view at both 1.25 times magnification [1] and 4 times magnification [2].

    - LAB MEDIA: Figure 5. Video editor: Highlight panel A labeled 1.25x

    - LAB MEDIA: Figure 5. Video editor: Highlight panel B labeled 4x