# Submission ID #: 68455

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20878823

# Title: MALDI-ToF MS Method for the Characterization of Synthetic Polymers with Varying Dispersity and End Groups

# Authors and Affiliations:

# Ashley V. Miles, Amman Nadeem, Mahi Ahmed, Brennan J. Curole, Scott M. Grayson

# Department of Chemistry, Tulane University

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# Scott M. Grayson 	sgrayson@tulane.edu

# Email Addresses for All Authors:

# Ashley V. Miles	amiles4@tulane.edu

# Amman Nadeem	anadeem@tulane.edu

# Mahi Ahmed		mahmad4@tulane.edu

# Brennan J. Curole 	bcurole1@tulane.edu

# Scott M. Grayson 	sgrayson@tulane.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  13

Number of Shots:  29 (18 SC)

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research questions will your laboratory focus on in the future?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

# Videographer: Please capture all testimonial shots in a wide-angle format with sufficient headspace, as the final videos will be rendered in a 1:1 aspect ratio. Testimonial statements will be presented live by the authors, sharing their spontaneous perspectives.

- Testimonial statements will not appear in the video but may be featured in the journal’s promotional materials.

- Provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author delivering the testimonial.

- Please answer the testimonial question live during the shoot, speaking naturally and in your own words in complete sentences.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Enter author name, Enter author title: (authors will present their testimonial statements live)

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Enter author name, Enter author title: (authors will present their testimonial statements live)

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Preparation of MALDI-ToF MS Target Plate

Demonstrator: Click here to enter name of demonstrator(s)

  - To begin, prepare the analyte acquisition solution by adding the cation, analyte, and matrix stock solutions into tube A [1] and vortex the tube to mix [2]. Then, add the cation, calibrant, and matrix stock solutions into tube B and mix thoroughly [3-TXT].

    - WIDE: Talent adding cation, analyte, and matrix stock solutions into tube A using micropipettes.

    - Talent vortexing the tube A.

    - Talent adding cation, calibrant, and matrix stock solutions into tube B and vortexing to mix. TXT: Cation:Analyte or Calibrant:Matrix stock = 1:5:20 µL  Video editor: This is a ratio

  - Now, pipette 0.5 microliters of the cation, analyte, and matrix sample mixture from tube A onto the MALDI-ToF MS (mall-dee-toff- M-S) target plate using the dried droplet method [1-TXT]. Then, pipette 0.5 microliters of the calibrant mix from tube B onto an adjacent spot on the target plate [2]. Record the exact row and column location of each spot for reference during data acquisition [3].

    - Talent pipetting the analyte sample mixture from tube A onto the MALDI-ToF MS target plate. TXT: MALDI-ToF MS: Matrix-Assisted Laser Desorption/Ionization Time-Of-Flight Mass Spectrometry

    - Talent pipetting the calibrant sample mixture from tube B onto an adjacent spot.

    - Talent writing down the row and column positions of each spot in a lab notebook.

  - Insert the MALDI-ToF MS target plate into the instrument [1] and dock the target plate using the instrument control software [2].

    - Talent placing the MALDI-ToF MS target plate into the instrument chamber.

    - Shot of docking the target plate.

- Selecting a Data Acquisition Method

Demonstrator: Click here to enter name of demonstrator(s)

  - Select and open a method file in the instrument control software [1]. Choose either the Reflector positive or Linear positive method based on the analyte’s mass range [2].

    - SCREEN: Show the Method file tab being opened.

    - SCREEN: selecting either Reflector positive or Linear positive method.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible:  https://review.jove.com/account/file-uploader?src=20878823

Please let me know if your institute’s system does not allow you to install the recording software.

  - Next, press the Start button on the instrument control software to begin MALDI-ToF MS data acquisition [1]. Allow the acquisition to complete automatically or press Stop when desired [2].

    - SCREEN: Show the Start button being clicked in the control software interface.

    - SCREEN: Show the acquisition process completing or Stop button being selected.

  - Then, click Add to transfer the MALDI-ToF MS single spectrum to the sum spectrum buffer [1]. Reposition the MALDI-ToF MS laser by viewing the camera feed and clicking anywhere within the same selected well to target a new position [2]. Repeat the MALDI-ToF MS acquisition to continue adding spectra until the results are satisfactory [3].

    - SCREEN: Show the Add button being clicked to transfer the spectrum to the sum buffer.

    - SCREEN: Show the camera feed of the well and mouse click repositioning the laser.

    - SCREEN: Show a good satisfactory specta.

- MALDI-ToF MS Calibration

Demonstrator: Click here to enter name of demonstrator(s)

  - Now, select the calibration sample on the MALDI-ToF MS target plate [1] and perform a new acquisition following the previously described data acquisition steps [1].

    - SCREEN: selecting the calibration sample spot on the target plate.

    - SCREEN: initiating acquisition.

  - Navigate to the calibration tab in the instrument control software and open the appropriate mass control list that corresponds to the selected calibrant [1].

    - SCREEN: Show the Calibration tab being accessed and a dropdown to select the relevant mass control list being opened.

  - Match the calibration peaks in the acquired MALDI-ToF mass spectrum with the reference masses listed in the mass control list [1].

    - SCREEN: cursor hovering over the acquired spectrum and the reference mass list, with visual matching of peaks.

  - In the mass control list, select the mass corresponding to the first calibration peak observed in the MALDI-ToF mass spectrum [1]. Set the reference peak by clicking to the left of the calibrant signal so that the peak is selected [2].

    - SCREEN: Show the mass control list being accessed and the first calibration mass being selected.

    - SCREEN: Show the calibrant signal in the mass spectrum being clicked to the left side to select the peak.

  - Then, click Apply to assign the calibrant signal as the reference mass [3]. Repeat the procedure for all additional calibration peaks by selecting the next mass of interest in the mass control list, aligning the corresponding calibrant peak, and applying the reference setting [4].

    - SCREEN: Show the Apply button being clicked to set the reference mass.

    - SCREEN: Show the other calibrated peaks.

  - For calibrated polymer analyte sample acquisition, locate and select the well containing the cation, analyte, and matrix mixture using the instrument control software [1]. Perform MALDI-ToF MS data acquisition using the calibrated method as demonstrated earlier [2]. Add an appropriate file name and save the calibrated analyte mass spectrum for further analysis [3].

    - SCREEN: Show the well containing the cation:analyte:matrix mixture being selected in the instrument software interface.

    - SCREEN: Show the Start of acquisition using the calibrated method, with steps such as laser targeting and spectrum acquisition briefly shown.

    - SCREEN: Show the file naming and Save dialog, and saving the mass spectrum file.

  - Finally, remove the MALDI-ToF MS target plate using the instrument control software [1] and clean the MALDI-ToF MS target plate thoroughly for reuse [2].

    - Shot of the target plate being removed.

    - Talent wiping the MALDI-ToF MS target plate with appropriate cleaning materials on a lab bench.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 120.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - The MALDI-ToF MS spectrum of TrisG1₆ (tris-G-1-6) showed a single peak at 505.239 Daltons, matching its theoretical mass and confirming complete deprotection [1].

    - LAB MEDIA: Figure 2B. Video editor: Highlight the peak labeled 505.239 in the spectrum.

  - PDMP polymers with chloride and proton end groups displayed a main peak at 859.426 Daltons [1-TXT], along with sodium-substituted [2] and cyclic variants [3].

    - LAB MEDIA: Figure 5. Video editor: Highlight the blue peak labeled 859.426. TXT: PDMP: Poly(2,2-dimethylpropanoate)

    - LAB MEDIA: Figure 5. Video editor: Highlight the orange peak labeled 881.413 to the right.

    - LAB MEDIA: Figure 5. Video editor: Highlight the green peak labeled 823.448 to the left.

  - The full PDMP distribution spanned 600 to 1600 Daltons with repeat units averaging 100.056 Daltons [1].

    - LAB MEDIA: Figure 4B. Video editor: Highlight the peaks from X axis range 600 to 1600.

  - l-PEB polymers exhibited alpha-hydroxy, alpha-propargyl, and mixed-end variants in three distributions [1].

    - LAB MEDIA: Figure 8B. Video editor: Sequentially highlight the three peaks labeled 1706.133, 1694.127, and 1700.139 and their coloured boxes above with green blue and red data.

  - Azide-functionalized l-PEB showed peaks for di-azido [1], azido-propargyl [2], and metastable ions [3].

    - LAB MEDIA: Figure 10. Video editor: Highlight the peak labeled 1994.296 in green.

    - LAB MEDIA: Figure 10. Video editor: Highlight the peak labeled 1839.184 in purple.

    - LAB MEDIA: Figure 11.

  - Comb polymers revealed oxidation states with 16 Dalton intervals [1] and K-plus adduct shifts confirming quadruple oxidation [2].

    - LAB MEDIA: Figure 14.

    - LAB MEDIA: Figure 15.

  - MeO-PEG (M-E-O-Peg) and unreacted alkynes were identified near 533 to 583 Daltons, validating reaction completeness [1].

    - LAB MEDIA: Figure 13. Video editor: Highlight the peaks at 533.309, 577.327, 539.263 and 583.306.

NOTE to Authors:

It's better if you upload each screen recording after naming each file according to the corresponding shot number (3-digit bullet, e.g, 2.1.1, 2.1.3, etc).

But if you prefer to record all the steps together and upload a single screen capture file, that is okay.

In this case, you need to indicate the corresponding timestamp against each shot.

For example if the action described in shot 2.1.2 is present in the file name xxx.mp4 and is in between timestamp 00:30 to 00:45 then write as following in script above (shown in red font below).

Example, shot 2.1.2 SCREEN: performing ........... action.   xxx.mp4 00:30-00:45.

2.2.2 SCREEN: clicking on.................button... xxx.mp4 01:00-01:25

Please note that every shot (3-digit step) should have the best corresponding video clip of only 20 to 25 seconds and not more than that, so that the voice narration can match the duration.

If a particular shot takes minutes to complete, we can choose timestamps at the beginning and end of the process so that the total time is again within the 25-second limit.

Example: 3.1.1 SCREEN: Clicking on the 'analyze and plot' button, analysis in progress and the graph being plotted.  xxx.mp4 00:00-00:10 and 02:30-02:40