# Submission ID #: 68515

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=20897123

# Title: A Multimodal Wide-Field Fourier-Transform Raman Microscope

# Authors and Affiliations:

# M. Riva1, B. Ardini1, A. Di Benedetto1, G. Valentini1,2, G. Cerullo1,2 , C. Manzoni2

# 1Dipartimento di Fisica, Politecnico di Milano

# 2Istituto di Fotonica e Nanotecnologie, IFN-CNR

# Corresponding Authors:

Cristian Manzoni 		(cristian.manzoni@cnr.it)

# Email Addresses for All Authors:

Martina Riva 		(martina.riva@polimi.it)

Benedetto Ardini		(benedetto.ardini@polimi.it)

Alessia Di Benedetto 	(alessia.dibenedetto@polimi.it)

Gianluca Valentini		(gianluca.valentini@polimi.it)

Giulio Cerullo		(giulio.cerullo@polimi.it)

Cristian Manzoni 		(cristian.manzoni@cnr.it)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done

3. Filming location: Will the filming need to take place in multiple locations?   No

Current Protocol Length

Number of Steps:  17

Number of Shots:  27

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

  - Martina Riva: We developed a multimodal hyperspectral microscope to acquire the Raman or photoluminescence spectrum from each pixel of an image which reduces the measurement time.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. 
AUTHOR’S NOTE: Please remove the symbol “P1” that appears at the top-left of the scene, above the drawing on the blackboard in file 0538.

What technologies are currently used to advance research in your field?

  - Giulio Cerullo : Conventional raster-scanning approach relies on dispersive spectrometers. Advanced techniques enhance the Raman signal either using light at resonance or with synchronized ultrashort laser pulses which drive molecular vibrations.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. 
Videographer’s Note : Use file 0540

What are the current experimental challenges?

  - Gianluca Valentini: The low Raman scattering cross-section poses two main challenges: (1) faint signal, which requires long acquisition times; (2) Raman signal may be overwhelmed by a stronger photoluminescence background.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:2.17
Videographer’s Note : Use file 0542

What advantage does your protocol offer compared to other techniques?

  - Cristian Manzoni: Fourier-transform spectroscopy enables parallel spectrum acquisition across all pixels, for faster measurement and a tunable sampling to access specific spectral information, such as isolating weak Raman peaks from strong luminescence background.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.1
Videographer’s Note : Use file 0543

What research questions will your laboratory focus on in the future?

  - Martina Riva:. The Fourier transform approach enables smart sampling and data analysis that leverage information in temporal traces. We aim at tailoring undersampling strategies to further shorten acquisition time.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:2.15

AUTHOR’S NOTE: Please remove the symbol “P1” that appears at the top-left of the scene, above the drawing on the blackboard in file 0539.

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

- Sample Preparation and Hyperspectral Raman Imaging

Demonstrator: Alessia Di Benedetto, Martina Riva

  - To begin, use a scraper to deposit 20 milligrams of each pigment powder on a precision balance, obtaining a one-to-one-to-one weight proportion [1]. Mix the powders with a mortar to remove clumps [2].

    - Talent placing pigment powders on the precision balance.
Videographer’s Note : Use file A7sIII_0545, A7sIII_0546, A7sIII_0547

    - Talent grinding powders to break up clumps.
Videographer’s Note : Use file A7sIII_0549

  - Pour the resulting mixture onto a microscope slide [1]. Then use the scraper tip to gently press it to obtain an almost uniform thickness of the layer [2].
Videographer’s Note : 2.2.1-2.3.2 merged. Use files A7sIII_0551, A7sIII_0552, A7sIII_0553

    - Talent pouring mixture onto the slide.

    - Talent using scraper tip to press mixture into uniform layer.

  - Apply nail polish on the edges of a microscope coverslip [1]. Place it on the mixture with the nail polish facing down and apply enough pressure to seal it [2-TXT].

    - Talent applying nail polish to coverslip edges.

    - Talent placing coverslip with nail-polish-side down and pressing to seal. TXT: Let it dry to harden the polish

  - Next, set the excitation wavelength and focus the excitation laser at the input of the large-core multimode fiber [1]. Put a test sample on the microscope stage and switch on the camera to view the laser illuminated spot [2].

    - Talent setting excitation wavelength and focusing laser into fiber.
Videographer’s Note : Use files A7sIII_0555, A7sIII_0556, A7sIII_0557

    - Talent putting a test sample on the stage and switching the camera on, showing the illumination spot.
Videographer’s Note : Use files A7sIII_0559, A7sIII_0560, 2_4_2

Added shot: SCREEN: 68515_screenshot_8.mp4	00:04-00:08, 00:14-00:20
Video editor: Please play 2.4.2 and added shot side by side in a split screen

  - Attach a middle section of the fiber to the vibrating membrane of a voice coil which will remove speckle [1]. Mechanically scramble the fiber by tightly bending it to merge all its spatial modes [2].

    - Talent attaching fiber section to voice-coil membrane.
Videographer’s Note : Use files A7sIII_0565

    - Talent tightly bending fiber to scramble modes.
Videographer’s Note : Use files A7sIII_0562

Added shot: SCREEN: 68515_screenshot_9.mp4 00:02-00:16
Video editor: Please play 2.5.2 and added shot side by side in a split screen

  - Insert a narrow bandpass filter at 532 nanometers with 2.0 nanometer bandwidth to clean the laser line and reject any unwanted pump spectral sidebands[1]. Use a dichroic mirror to reflect the laser light toward the sample and transmit the redshifted back-scattered radiation collected by the objective [2].

    - Talent inserting the narrow bandpass filter at 532 nm in the microscope. 
Videographer’s Note : Use files A7sIII_0568

    - Talent inserting the dichroic mirror to direct excitation and collect backscattered signal.
Videographer’s Note : Use files A7sIII_0569

  - Insert a long-pass filter at 532 nanometers to reject residual illumination light [1].

    - Talent inserting the longpass filter at 532 nm.
Videographer’s Note : Use files A7sIII_0567, A7sIII_0566

  - Measure the power on the sample plane with a power meter [1]. Adjust the pump beam to obtain an intensity at the sample which does not lead to damage [2].
Videographer’s Note : 2.8.1-2.8.2 merged. Use files A7sIII_0571, A7sIII_0572

    - Talent measuring sample-plane power with power meter.

Talent adjusting pump beam to safe intensity.

  - Now put the sample on the microscope stage and adjust the focus [1].

    - Talent placing slide on stage and focusing the microscope.
Videographer’s Note : use file A7sIII_0573

  - Switch on the driver of the motor that performs the wedge translation and press the Acquire option [1]. Then set the camera acquisition parameters like time frame and hardware binning to optimize the signal intensity [2-TXT].

    - SCREEN: 68515_screenshot_1.mp4  	00:00-00:16

    - SCREEN: 68515_screenshot_2.mp4 	00:00-00:13
TXT: Acquisition of photoluminescence.

  - Now set the step length, home position and number of steps [1-TXT].

    - SCREEN: 68515_screenshot_3.mp4  	00:00-00:16
TXT: Wait until the sample has reached a stationary condition

  - Add the name of the directory and the file name [1]. Move the cross point to the target position and click on Measure to start the acquisition of a monochrome image for each wedge position [2]. Once the measurement is finished, switch off the laser before data analysis [3].

    - SCREEN: 68515_screenshot_4.mp4  	00:00-00:14

    - SCREEN: 68515_screenshot_4.mp4  	 00:16-00:40

    - Talent switching off the laser.
Videographer’s Note : use file A7sIII_0576

  - Next, generate the spectral hypercube from the acquired dataset by loading the motor positions correction file specific to the stepper motor [1] and the frequency calibration file corresponding to the interferometer [2].

    - SCREEN: 68515_screenshot_5.mp4  	00:00-00:19

    - SCREEN: 68515_screenshot_5.mp4  	00:20-00:25

  - Set the wavelength for the spectra to be computed by Fourier transforming the interferograms of all pixels [1].

    - SCREEN: 68515_screenshot_5.mp4  	00:26-00:32

  - Apply the apodization function which offers a good trade-off between spectral broadening and artifact reduction, such as the Happ-Genzel window and select the pixel [1]. Then generate the spectral hypercube and save it in complex values [2].

    - SCREEN: 68515_screenshot_5.mp4  	00:33-00:46

    - SCREEN: 68515_screenshot_5.mp4  	00:47-00:53

  - Launch the analysis software and open the spectral hypercube in complex values [1]. Generate a false-colour RGB image [2] and obtain the average spectrum in selected areas to analyze the spectral hypercube [3].

    - SCREEN: 68515_screenshot_6.mp4  		00:00-00:08

    - SCREEN: 68515_screenshot_6.mp4  		00:09-00:19

    - SCREEN: 68515_screenshot_6.mp4  		00:20-00:32

  - For the Raman measurement on the same field of view, set the time frame and hardware binning [1]. Then set the set length, home position and number of steps [2]. Choose the file name and directory for saving and move the crosshair to the target position before pressing Measure [3].

    - Talent inserting the short pass filter at 600 nm.  
Videographer’s Note : use file A7sIII_0577

    - SCREEN: 68515_screenshot_7.mp4  	00:00-00:10	.

    - SCREEN: 68515_screenshot_7.mp4  	00:11-00:20

    - SCREEN: 68515_screenshot_7.mp4  	00:21-00:44

# Results

- Results

  - Three pigments are clearly distinguished within the field of view based on their Raman spectral signatures [1], with characteristic peaks for rutile at 454 and 616 inverse centimeters [2], anatase at 396, 514, and 641 inverse centimeters [3], and cadmium yellow at 301 and 605 inverse centimeters [4].

    - LAB MEDIA: Figure 4A. Video editor: Highlight the color-coded regions labeled 1, 2, and 3.

    - LAB MEDIA: Figure 4B. Video editor: Highlight the two peaks at 454 and 616 in the top cyan-colored graph labeled “Rutile (1)”.

    - LAB MEDIA: Figure 4B. Video editor: Highlight the three peaks at 396, 514, and 641 in the middle yellow-colored graph labeled “Anatase (2)”.

    - LAB MEDIA: Figure 4B. Video editor: Highlight the two peaks at 301 and 605 in the bottom magenta-colored graph labeled “Cadmium Yellow (3)”.

  - In the photoluminescence map, instead, cadmium yellow is the only pigment visibly distributed across the field of view [1], with strong emission dominating the entire image [2]. Indeed, the faint defect emission from rutile and anatase cannot be spectrally distinguished from the yellow powder [3].

    - LAB MEDIA: Figure 5A.

    - LAB MEDIA: Figure 5B. Video editor: Highlight the dominant magenta-colored spectrum rising across the full wavelength range labeled “Cd Yellow (3)”.

    - LAB MEDIA: Figure 5B. Video editor: Highlight the much smaller cyan and yellow peaks labeled “Rutile (1)” and “Anatase (2)” in the lower portion of the graph.

# Pronunciation List

- Multimodal
Pronunciation link: https://www.merriam-webster.com/dictionary/multimodal
IPA: /ˌmʌltiˈmoʊdəl/
Phonetic Spelling: mul-tee-moh-dəl

- Fourier-Transform
Pronunciation link: https://www.merriam-webster.com/dictionary/Fourier-transform
IPA: /ˈfʊəri.eɪ-ˌtrænsˌfɔːrm/ (American: /ˈfʊri.eɪ-ˌtrænsˌfɔrm/)
Phonetic Spelling: foor-ee-ay-trans-form

- Raman
Pronunciation link: https://www.merriam-webster.com/dictionary/Raman
IPA: /ˈrɑːmɑːn/ or /ˈreɪmən/ (American commonly /ˈreɪmən/)
Phonetic Spelling: ray-mən

- Microscope
Pronunciation link: https://www.merriam-webster.com/dictionary/microscope
IPA: /ˈmaɪkrəˌskoʊp/
Phonetic Spelling: my-kruh-skohp

- Hyperspectral
Pronunciation link: No confirmed link found
IPA: /ˌhaɪpərˈspɛktrəl/
Phonetic Spelling: hy-per-spek-trəl

- Photoluminescence
Pronunciation link: https://www.merriam-webster.com/dictionary/photoluminescence
IPA: /ˌfoʊtoʊˌluːməˈnɛsəns/
Phonetic Spelling: foh-toh-loo-muh-nes-uhns

- Interferogram
Pronunciation link: No confirmed link found
IPA: /ˌɪntərˈfɪrəˌɡræm/
Phonetic Spelling: in-ter-fi-ruh-gram

- Apodization
Pronunciation link: No confirmed link found
IPA: /əˌpoʊdəˈzeɪʃən/
Phonetic Spelling: uh-poh-duh-zay-shun

- Birefringent
Pronunciation link: https://www.merriam-webster.com/dictionary/birefringent
IPA: /ˌbaɪrəˈfrɪndʒənt/
Phonetic Spelling: by-ruh-frin-juhnt

- Dichroic
Pronunciation link: https://www.merriam-webster.com/dictionary/dichroic
IPA: /daɪˈkroʊɪk/
Phonetic Spelling: dy-kroh-ik

- Nanometer
Pronunciation link: https://www.merriam-webster.com/dictionary/nanometer
IPA: /ˈnænəˌmiːtər/
Phonetic Spelling: nan-oh-mee-ter

- Spectral
Pronunciation link: https://www.merriam-webster.com/dictionary/spectral
IPA: /ˈspɛktrəl/
Phonetic Spelling: spek-trəl

- Pixel
Pronunciation link: https://www.merriam-webster.com/dictionary/pixel
IPA: /ˈpɪksəl/
Phonetic Spelling: pik-səl

- Binning
Pronunciation link: No confirmed link found
IPA: /ˈbɪnɪŋ/
Phonetic Spelling: bin-ning

- Interferometer
Pronunciation link: https://www.merriam-webster.com/dictionary/interferometer
IPA: /ˌɪntərˌfɪrəˈmɛtər/
Phonetic Spelling: in-ter-fi-ruh-meh-ter