Submission ID #: 68532

Scriptwriter Name: Poornima G

Project page Link: https://review.jove.com/account/file-uploader?src=20902308

Title: Infusion of Targeted Radionuclide Therapy Using a Peristaltic Pump

Q: Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter the distance between locations.

Authors: To ensure that your script can be filmed in one day, the Protocol section is restricted to 35 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  12

Number of Shots:  30

# Protocol

NOTE TO THE AUTHORS: 
Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The one-digit numbers represent sections of the video. The text will appear onscreen.

- The two-digit numbers (e.g. 2.1., 2.2.) represent steps of your protocol. The text will be recorded by a professional voiceover talent.

- The three-digit numbers (e.g. 2.1.1., 2.2.2.) represent the shots that our videographer will capture at your lab.

- To ensure that your protocol can be filmed in one day, the protocol is restricted to 20 shots.

Please use this draft script to help you prepare for filming day.

- Filming should take no more than 10 minutes per step. If a step will take more than 10 minutes, prepare the product from that step in advance.

- Protocol

To begin, insert the infusion set spike into the inlet of the 250 milliliter sodium chloride 0.9 percent bag [1]. Apply gentle pressure to fill the infusion chamber three-quarters full [2].

WIDE: Talent inserting the spike of the infusion set into the inlet of the sodium chloride bag.

Talent squeezing the bag to fill the infusion chamber three-quarters full.

Declamp the infusion set and slightly unscrew the flow regulator [1]. Turn the perfusion set cassette upside down [2] and then back to its original position as soon as the first drops of saline enter the chamber [3].

Talent declamping the infusion set and loosening the flow regulator.

Talent flipping the cassette upside down.

Talent returning the cassette upright as saline enters the chamber.

Tap the cassette to remove air bubbles [1] and press the flow regulator to top the purge [2].

Talent tapping the cassette to dislodge air bubbles.

Talent pressing the flow regulator to complete the purge.

Now, switch on the pump with the on-off button [1] and lift the lever to open the pump compartment [2]. Hold the cassette by the flow regulator grip, position it in the compartment flap slots, and slide it all the way in [3].

Talent pressing the on/off button.

Talent lifting the pump lever.

Talent inserting the cassette into the pump compartment using the flow regulator grip.

Next, pull down the lever to close the compartment flap [1] and hold line A coming out of the cassette using the guide above the compartment [2].

Talent pulling down the pump lever to secure the cassette.

Talent aligning and securing line A using the guide on the pump.

Then, remove the stylet from the 19-gauge spinal needle and screw it onto the free luer lock adapter from the 2-way extension kit set [1]. Keep the needle in its protective packaging [2] and connect the spinal needle via this adapter to the extension kit line using the closed system transfer valve [3].

Talent removing the stylet from the spinal needle and attaching it to the luer lock adapter.

Talent holding the needle within its protective packaging.

Talent connecting the needle assembly to the extension line via the transfer valve.

Now, connect a sterile 10-milliliter saline syringe to the straight infusion connector on the extension line [1] and purge the tubing with approximately 5 milliliters of sodium chloride 0.9 percent [2].

Talent attaching a 10 milliliter syringe to the extension line.

Talent purging the tubing with saline solution.

Connect the luer lock end of the extension line to the line B valve [1] and keep the sterile saline syringe attached to the straight infusion connector [2]. Press and hold the corresponding button on the pump to begin back-priming of line B [3] until saline flows through the 19 gauge spinal needle [4].

Talent connecting the extension line to line B valve.

Talent leaving the syringe connected to the infusion connector.

Talent pressing the pump button.

Shot of saline flowing through the needle.

For infusion initiation, connect the end of line A, which has a luer lock connector, to the patient’s catheter valve [1]. Move the mobile shielded screen in line with the pump and patient to provide radiation protection [2] and start line A to begin the saline infusion to the patient [3].

Talent connecting line A’s luer lock to the catheter valve on the patient.

Talent adjusting the mobile shielded screen between the patient and pump.

Talent activating line A on the pump.

Place the shielded container holding the TRT vial onto the inclined vial stand on the tray of the shielded paravent [1]. Remove the cap from the container [2], and using forceps, apply an alcohol-soaked compress to disinfect the septum of the TRT vial [3]. Insert the venting filter into the septum, ensuring the needle does not touch the liquid [4].

Talent placing the shielded TRT container on the inclined stand.

Talent removing the cap.

Talent disinfecting the vial septum with forceps and alcohol compress.

Talent inserting the venting filter carefully into the vial septum.

Now, insert the 19-gauge spinal needle connected to line B into the septum of the TRT vial [1].

Talent inserting the spinal needle from line B into the TRT vial septum.

Finally, press the green button on the pump to start the infusion in simultaneous mode [1].

Talent pressing the green button to initiate simultaneous infusion.