# Submission ID #: 68612

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/files_upload.php?src=20926708

Title: Ultrasound Localization Microscopy for Super-Resolution Mapping of the Rodent Brain Microvasculature

# Authors and Affiliations:

Pia Virginia Pelaez1*, Sara Romanzi1*, Adrien Bertolo1, Mohamed Nouhoum1, Theo Meranger1, Nathalie Ialy-Radio2, Thomas Deffieux2, Sophie Pezet2, Jeremy Ferrier1

1Iconeus

2Physics for Medicine Paris, Inserm U1273, ESPCI Paris, PSL University

*These authors contributed equally

# Corresponding Authors:

Jeremy Ferrier				(jeremy.ferrier@iconeus.com)

# Email Addresses for All Authors:

Pia Virginia Pelaez		(pia.pelaez@iconeus.com)

Sara Romanzi			(sara.romanzi@iconeus.com)

Adrien Bertolo			(adrien.bertolo@iconeus.com)

Mohamed Nouhoum		(mohamed.nouhoum@iconeus.com)

Theo Meranger		(theo.meranger@iconeus.com)

Nathalie Ialy-Radio		(nathalie.ialy-radio@inserm.fr)

Thomas Deffieux		(thomas.deffieux@inserm.fr)

Sophie Pezet			(sophie.pezet@espci.fr)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

Current Protocol Length

Number of Steps: 31

Number of Shots: 55

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Sara Romanzi:  Our research aims to characterize the brain’s microvascular architecture and flow, while refining super-resolution ultrasound methods.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What are the most recent developments in your field of research?

  - Sara Romanzi:  Recent developments in our research field have optimized microbubble tracking and strengthened processing pipelines for high-resolution vascular mapping with ULM.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What significant findings have you established in your field?

  - Pia Pelaez This work establishes that ultrasound localization microscopy can reliably produce super-resolution maps of microvascular structure and flow in vivo.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

What research gap are you addressing with your protocol?

  - Pia Pelaez : This protocol fills the need for a standardized procedure that enables reliable ULM acquisitions in living rodents using a dedicated fUS platform.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

How will your findings advance research in your field?

  - Pia Pelaez : These findings establish a dependable protocol that allows researchers to capture microvascular dynamics with high precision and consistency.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

Ethics Title Card

This research has been approved by  by our local ethics committee (Comité d’éthique en matière d’expérimentation animale number 59, ‘Paris Centre et Sud’, project #2020-16)

# Protocol

- Animal Preparation and Tail Vein Catheterization

Demonstrator: Click here to enter name of demonstrator(s)

AUTHORS: Please provide the name of the demonstrator(s)

  - To begin, place an anesthetized mouse onto a stereotaxic frame with a heating pad set to 37 degrees Celsius [1-TXT].

    - WIDE: Talent placing the anesthetized mouse into the stereotaxic frame over the heating pad. TXT: Anesthesia: 2% isoflurane with air and O2

  - Using firm but non-damaging pressure, pinch the toe or tail to check for the absence of reflexes and confirm adequate anesthesia depth [1-TXT]. Apply eye ointment to both eyes to prevent cataract formation during the procedure [2].

    - Talent gently pinch the tail of the mouse. TXT: If reflex observed, slightly increase isoflurane and reassess

    - Talent gently applying eye ointment to the mouse's eyes using a sterile applicator.

  - Now, shave the head of the mouse using a trimmer [1]. Apply depilatory cream, let it sit for 1 to 1.5 minutes [2], then rinse with warm water and dry with gauze [2].

    - Talent trimming the fur on the mouse's head using an electric trimmer.

    - Talent applying depilatory cream to the shaved area.

    - Talent rinsing and drying the skin with gauze. TXT: Repeat as needed until all hair is removed; Avoid leaving cream on too long

  - Next, apply centrifuged ultrasound gel onto the head to ensure acoustic coupling [1].

    - Talent applying a layer of ultrasound gel evenly across the mouse’s scalp.

  - For tail vein catheterization, use a 27-gauge by half-inch winged infusion set connected to an 18-gauge by half-inch needle with 8 inches of tubing [1-TXT].

    - Talent holding a winged infusion set and attaching it to an 18-gauge needle using 8-inch tubing. TXT: Shorten the tubing, if possible, to minimize dead volume

  - Gently turn the animal’s tail 90 degrees to expose the lateral tail veins [1]. Then secure the tail in place with tape to maintain stability and prevent movement during catheterization [2].

    - Talent rotating the mouse’s tail sideways to a 90-degree position.

    - Talent taping the tail to the work surface to hold it in place securely.

  - Now, identify the tail vein. [1]. To improve vein visibility, warm the tail with lukewarm water to promote vasodilation [2].

    - Shot of the tail vein.

    - Talent applying a warm, damp cloth to the tail to facilitate vasodilation.

  - Using a finger, apply gentle pressure at the base of the tail and slide it toward the injection site against the direction of blood flow [1]. Insert the needle parallel to the vein with the bevel facing up [2].

    - Talent pressing at the tail base and sliding a finger toward the mid-tail region.

    - Shot of the needle being inserted at a shallow angle into the distended lateral tail vein with the bevel clearly positioned upward.

  - Observe the tubing for blood flashback to confirm correct needle placement [1].

    - Shot of blood visibly flowing back into the tubing, confirming successful vein entry.

  - On the imaging system, launch the functional ultrasound acquisition software and create a new experiment session [1]. Adjust the position of the ultrasound probe [2].

    - SCREEN: Show the software interface as the user clicks on New Session or Create Experiment within the IcoScan software.

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/files_upload.php?src=20926708

    - SCREEN: The panel probe position is being adjusted.

  - Now, start the Live View acquisition in the software and adjust the probe in real time using power doppler imaging to center the brain on the screen [1]. Optimize the imaging parameters, such as contrast and compression, to enhance visualization quality [2].

    - SCREEN: Show the Live View mode activated with real-time CBV imaging displayed, and talent adjusting the probe accordingly to center the brain.

    - SCREEN: Show the Contrast and Compression settings being adjusted to enhance image clarity and improve signal-to-noise.

  - Then, open the Angio3D (Angio-3-D) menu in the acquisition software [1]. Adjust the scanning parameters to fully cover the brain, then start the scan acquisition [2].

    - SCREEN: Show the user selecting the Angio3D menu.

    - SCREEN: Show the parameter fields being filled in and the Start Acquisition button being clicked to begin the scan

.

  - Keep the acquisition software open and launch the IcoStudio (Ico-Studio) software for analysis and visualization. Load the Angio3D scan for review [1].

    - SCREEN: Show the IcoStudio being launched from the desktop or applications menu. Load the angio scan.

  - Navigate to the Brain Registration panel and register the Angio3D scan to the Allen Mouse Common Coordinate Framework using either the fully automatic or manual mode [1]. Save the completed registration as a .bps (B-P-S) file [1].

    - SCREEN: Show the Brain Registration panel open in IcoStudio and the user selecting either Automatic Registration or Manual Registration

    - SCREEN: Show the Save As dialog box and the user typing a file name with the .bps extension and clicking Save.

  - Now, open the Brain Navigation menu and scroll down to the Atlas Manager panel [1]. Use the Parent and Child tree navigator to browse the Allen Mouse Brain Atlas and select the desired anatomical regions for overlay in the 3-view panel [2].

    - SCREEN: Show the Brain Navigation tab and Atlas Manager panel in IcoStudio.

    - SCREEN: Demonstrate selecting anatomical structures using the hierarchical tree and viewing the selected regions overlaid in the 3-view display.

  - Choose an imaging plane that overlaps with the selected regions of interest [1]. In the coronal view, manually place two markers to define the desired imaging slice [2].

    - SCREEN: Show the imaging plane being adjusted to intersect the targeted anatomical areas.

    - SCREEN: User clicking to set two markers within the coronal panel, defining the imaging slice.

  - Then, click on BPS Positioning to extract the motor coordinates for the selected imaging plane [1]. Review the image preview from the Angio3D scan and click Copy BPS Coordinates [2].

    - SCREEN: Show the BPS window displaying calculated coordinates along with a preview image of the slice.

    - SCREEN: User clicking on Copy BPS Coordinates to save the positional data.

  - Switch to the IcoScan software and open the Move Probe panel [1]. Click on Neuronavigation, paste the copied values, and allow the system to automatically align the probe with the selected imaging plane [2].

    - SCREEN: Show the user switching windows to IcoScan and opening the Probe Positioning panel.

    - SCREEN: Show the coordinates being pasted into the input field, and the probe adjusting its position automatically in response.

  - Then, start Live Imaging to confirm that the current imaging plane aligns with the previously selected plane [1].

    - SCREEN: Show the Live View screen displaying a real-time image matching the anatomical region previously selected.

- Ultrasound Localization Microscopy (ULM) Acquisition of the Brain Microvasculature

Demonstrator: Click here to enter name of demonstrator(s)

  - Using a sterile syringe, inject 5 milliliters of 0.9 percent sodium chloride solution into the vial through the septum [1]. Agitate the vial vigorously for 20 seconds to resuspend the microbubbles and ensure uniform dispersion [2].

    - Talent inserting a sterile syringe through the septum of a vial and slowly injecting 5 milliliters of sodium chloride solution.

    - Talent holding the vial and shaking it energetically for 20 seconds.

  - Predefine the recording parameters, including ultrasound sequence, frame rate, and total recording duration [1].

    - SCREEN: Show the software interface where the ultrasound sequence, frame rate, and recording time are entered in their respective fields.

  - Then, draw the desired volume of microbubbles for injection and begin the acquisition [1]. Once acquisition starts, inject 100 microliters for mice, using the appropriate dilution [2]. Confirm contrast enhancement on the live visualization panel [3]. Immediately flush the remaining microbubbles in the dead volume by injecting physiological fluid into the catheter [4].

    - Talent drawing microbubble solution into a syringe according to the species-specific volume.

    - Talent injecting the microbubbles into the catheter.

    - SCREEN: Live View showing real-time enhancement in the image due to microbubbles.

    - Talent injecting physiological fluid into the catheter to flush remaining microbubbles.

  - Wait for the acquisition to finish and save the recorded data [1-TXT].

    - SCREEN: Show the acquisition progress bar reaching completion and the Save button being clicked to store the data. TXT: Stop isoflurane and monitor the animal in a warm area until fully awake

- Image Reconstruction using the IcoLab Software

Demonstrator: Click here to enter name of demonstrator(s)

  - Open the IcoLab software and click on the Ultrasound Localization Microscopy tab located on the leftmost part of the screen to access the ultrasound localization microscopy interface [1]. Then, click on Compute ULM Maps to begin the processing workflow [2].

    - SCREEN: Show the IcoLab main interface with the ULM tab being selected on the left sidebar.

    - SCREEN: Show the user clicking on Compute ULM Maps to launch the workflow.

  - Select the source folder that contains the previously acquired data [1]. Then, select the specific scan file to process and adjust the processing time boundaries as needed [2].

    - SCREEN: Show the folder selection dialog with the path being chosen to locate the recorded scan data.

    - SCREEN: Display a list of available scan files and show the user selecting one, then editing the time boundaries using input fields.

  - Click on Next to move to the Reporting step. Set the output folder path for saving the processed files and configure the visualization parameters for the report images [2].

    - SCREEN: Show the Next button being clicked and the interface switching to the Reporting setup screen. Display the user browsing to set an output folder, then selecting visualization options.

  - Now click on RUN to begin processing the data [1]. Once completed, a .trk (T-R-K) file will be created containing localized microbubble coordinates, along with TIFF files representing density, velocity, and backscattered amplitude [2]. Load the .trk file into IcoStudio to begin exploring the processed data [3].

    - SCREEN: Show the RUN button being clicked.

    - SCREEN:  Show the output folder displaying a .trk file and several TIFF files being generated after processing.

    - SCREEN: Show IcoStudio interface with the .trk file being imported and the visualized microbubble tracks appearing.

  - Then, use the right-side panel in IcoStudio to adjust visualization parameters such as contrast, compression, and colormap for enhanced data representation [1].

    - SCREEN: Display the adjustment panel in IcoStudio with the user modifying the contrast, compression level, and selecting a new colormap for optimal visualization.

  - To set the integration time, adjust the double slider located at the bottom of the main view to define the start and end frames for the analysis [1]. The Density tab shows the distribution of microbubble occurrences. Choose between Absolute Density and Directional Density [2].

    - SCREEN: Show the user dragging the two ends of the double slider to select a specific frame range for integration.

    - SCREEN: Show the Density tab and the microbubble density map. Demonstrate toggling between Absolute Density and Directional Density options

  - After that, click on the Velocity tab to display microbubble speed in millimeters per second using Absolute, Directional, Axial, and Lateral modes [1]. Adjust the contrast, compression, and velocity thresholds to improve image clarity [2].

    - SCREEN: Show the velocity map appearing after selecting the Velocity tab, with display options for the four velocity modes.

    - SCREEN: User fine-tuning visualization parameters such as contrast sliders, compression levels, and threshold values.

  - Then explore the BSA tab to view microbubble reflectivity, which enhances contrast and reveals microvascular structures and out-of-plane motion [1]. Export images as high-resolution PNGs (P-N-Geez) or generate TIFF files for quantitative analysis with customizable pixel size [2].

    - SCREEN: Show the BSA tab being selected and a reflectivity-based image appearing with visible vascular patterns.

    - SCREEN: User exporting and saving images as PNG.

# Results

- Results

  - Microbubble density maps of a coronal rat brain slice acquired with ULM resolved both large vessels and fine capillary structures, revealing detailed vascular topology including bifurcations and branching patterns [1].

    - LAB MEDIA: Figure 5. Video editor: Highlight the leftmost full brain panel under “Density,” then zoom in on the bottom left inset

  - The velocity map of the same brain slice showed clear axial directional flow, with upward flow representing venous drainage and downward flow corresponding to arterial perfusion, consistent with known cortical vascular architecture [1].

    - LAB MEDIA: Figure 5. Video editor: Highlight the center full brain panel under “Velocities,” then zoom in on the bottom center inset showing vessels with both blue and orange

  - The backscattered amplitude map provided depth-dependent contrast and highlighted vessels based on their elevational position, useful for segmenting out-of-plane or low-amplitude vessels [1].

    - LAB MEDIA: Figure 5. Video editor: Highlight the rightmost full brain panel under “Backscattered Amplitude,” then zoom in on the bottom right inset

  - Vascular density maps acquired one week apart in the same mouse showed strong spatial overlap, confirming the reproducibility of coronal plane targeting using the Brain Positioning System [1].

    - LAB MEDIA: Figure 6A. Video editor: Highlight the “Comparison” image

  - Quantitative analysis of the left somatosensory cortex across sessions showed similar distributions in vessel radius [1], velocity [2], and flow rate, demonstrating measurement reproducibility [3].

    - LAB MEDIA: Figure 6B. Video editor: Highlight the box plots labeled “Radius” under “Week 1” and “Week 2” in the right panel.

    - LAB MEDIA: Figure 6B. Video editor: Highlight the box plots labeled “Velocity” under “Week 1” and “Week 2” in the right panel.

    - LAB MEDIA: Figure 6B. Video editor: Highlight the box plots labeled “Flowrate” under “Week 1” and “Week 2” in the right panel.

  - Similar reproducibility was observed in the left thalamus region, with consistent values for vessel radius [1], velocity [2], and flow rate between Week 1 and Week 2 sessions [3].

    - LAB MEDIA: Figure 6C. Video editor: Highlight the box plots labeled “Radius” under “Week 1” and “Week 2” in the right panel.

    - LAB MEDIA: Figure 6C. Video editor: Highlight the box plots labeled “Velocity” under “Week 1” and “Week 2” in the right panel.

    - LAB MEDIA: Figure 6C. Video editor: Highlight the box plots labeled “Flowrate” under “Week 1” and “Week 2” in the right panel.