# Submission ID #: 68722

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/account/file-uploader?src=20960258

# Title: An Open-Source Normothermic Perfusion System Designed for Research Scientists

# Authors and Affiliations:

Lindsay R. Friedman1*, Hannah E. Stepp1*, Jacob Lambdin1*, Allen Luna1*, Alexander Rossi1*, Stephanie C. Lux1*, Michael M. Wach1, Kenneth Luberice1, Samantha M. Ruff1, James McDonald1, Reed I. Ayabe1, Shahyan U. Rehman1, Carrie Ryan1, Stephanie Gregory1, Alyssa V. Eade1, Carolina Larrain1, Jack H. Victory1, Randy Pursley2, Taylor Noel2 , Max Freidlin2 , Shreya Gupta1, Tracey Pu1, Ashley Rainey1, Emily Verbus1, Martha Teke1, Areeba Saif1, Tahsin M. Khan1, Imani A. Alexander1, Surajit Sinha1, Sarfraz R. Akmal1, Ritish R. Maram3, Leila Sarvestani1, Stephen M. Hewitt4, James Doroshow5, Kirsten Remmert1, Emily Smith1,  Priyanka Desai1, Ning Miao6, Kevin Driscoll6, Kamille West7, Theo Heller8, Elliot Levy3, Jeremy Davis1, Andrew Blakely1, David E. Kleiner4, John Kakareka2 , Thomas Pohida2 , Jonathan M. Hernandez1,9

1Surgical Oncology Program, National Cancer Institute, National Institutes of Health

2Instrumentation Development and Engineering Application Solutions, National Institute of Biomedical Imaging and Bioengineering, National Institutes of Health

3Radiology and Imaging Sciences, Department of Interventional Radiology, National Institutes of Health

4Histopathology and Autopsy Pathology, Laboratory of Pathology, National Institutes of Health

5Division of Cancer Treatment and Diagnosis, National Cancer Institute, National Institutes of Health

6Department of Perioperative Medicine, National Institutes of Health

7Blood Services Section, Department of Transfusion Medicine, National Institutes of Health

8Translational Hepatology Section, Liver Disease Branch, National Institute of Diabetes and Digestive and Kidney Diseases, National Institute of Health

9Center for Immuno-Oncology, National Cancer Institute, National Institutes of Health

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# Jonathan M. Hernandez 		(jonathan.hernandez@nih.gov)

# Email Addresses for All Authors:

Lindsay R. Friedman			(lrfriedman.md@gmail.com)

Hannah E. Stepp			(hstepp15@gmail.com)

Jacob Lambdin			(jlambdin@email.gwu.edu)

Allen Luna				(allen.j.luna@vanderbilt.edu)

Alexander Rossi			(rossiaj12@gmail.com)

Stephanie C. Lux			(stephie.lux106@gmail.com)

Michael M. Wach			(michael.wach@moffitt.org)

Kenneth Luberice			(kenneth.luberice@gmail.com)

Samantha M. Ruff			(samantha.marilyn.ruff@gmail.com)

James McDonald			(JDMcDonald@utmck.edu)

Reed I. Ayabe				(reed.ayabe@gmail.com)

Shahyan U. Rehman			(shahyan.rehman@yale.edu)

Carrie Ryan				(carriery@buffalo.edu)

Stephanie Gregory			(sngregoryapp@gmail.com)

Alyssa V. Eade				(alyssa.eade@nih.gov)

Carolina Larrain			(carolina.larrain@nih.gov)

Jack H. Victory				(jhv00001@mix.wvu.edu)

Randy Pursley				(pursley@nih.gov)

Taylor Noel 				(taylor.noel@nih.gov)

Max Freidlin				(maxfreidlin@gatech.edu)

Shreya Gupta				(shgupta91@gmail.com)

Tracey Pu				(tracey.pu@nih.gov)

Ashley Rainey				(ashleyjrainey@gmail.com)

Emily Verbus				(verbuse@ccf.org)

Martha Teke				(martha.teke@utsouthwestern.edu)

Areeba Saif				(areeba.saif9@gmail.com)

Tahsin M. Khan			(tahsin.khan@mountsinai.org)

Imani A. Alexander			(imani.aalexander@gmail.com)

Surajit Sinha				(surajit.sinha@nih.gov)

Sarfraz R. Akmal			(sra112@rwjms.rutgers.edu)

Ritish R. Maram			(ritish.maram@nih.gov)

Leila Sarvestani			(leila.sarvestani@nih.gov)

Stephen M. Hewitt			(hewitts@mail.nih.gov)

James Doroshow			(doroshoj@mail.nih.gov)

Kirsten Remmert			(kirsten.remmert@nih.gov)

Emily Smith				(emily.smith@nih.gov)

Priyanka Desai				(Priyanka.desai@nih.gov)

Ning Miao				(nmiao@cc.nih.gov)

Kevin Driscoll				(kevin.driscoll@nih.gov)

Kamille West				(kamille.west@nih.gov)

Theo Heller				(theoh@intra.niddk.nih.gov)

Elliot Levy				(levyeb@cc.nih.gov)

Jeremy Davis				(jeremy.davis@nih.gov)

Andrew Blakely			(Andrew.Blakely@Nih.gov)

David E. Kleiner			(kleinerd@mail.nih.gov)

John Kakareka				(kakareka@exchange.nih.gov)

Thomas Pohida			(pohidat@exchange.nih.gov)

# Jonathan M. Hernandez 		(jonathan.hernandez@nih.gov)

## Author Questionnaire

1. We have marked your project as author-provided footage, meaning you film the video yourself and provide JoVE with the footage to edit. JoVE will not send the videographer. Please confirm that this is correct.

☐ Correct

☐ Incorrect

2. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

3. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen captured video files to your project page as soon as possible.

4. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: MM/DD/YYYY

Authors: Since we have protocol footage, this question is applicable for filming interview statements

DO NOT use this draft script for filming. Please wait until your script is finalized to begin the filming process.

When you are ready to submit your video files, please contact our Content Manager, Utkarsh Khare.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  27

Number of Shots:  55

# Introduction

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research questions will your laboratory focus on in the future?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

Ethics Title Card

This research is performed in accordance with the ethical standards of our institution’s Human Research Ethics Committee and adheres to all applicable guidelines and regulations for research involving human subjects and tissues

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that you/your videographer will capture.

- Machine Set-Up Normothermic Ex Vivo Perfusion System

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

NOTE: The script protocol section is drafted from available footage

  - To begin, place the drum funnel on the right side of the workstation with the spout extending past the front edge of the table [1]. Place a circular memory foam to fit the base of the drum with the diaphragm underneath the foam [2], leaving a 4-inch gap near the spout for outflow [3].

    - 1.1.4.MOV: 00:003-00:06

    - 1.1.4.MOV: 00:07-00:12

    - 1.1.4.MOV: 00:13-00:16

  - Then, place two water heating blankets on top of the foam insert with the connector tubing facing the top left corner and insert a three-sixteenth inch male Luer barb into each tube [1].

    - 1.1.5.MOV: 00:02-00:13

  - Take a folded sterile warmer drape, cut 2 millimeters off the corner of the fold to create an outflow hole, and place it on the drum [1].

    - 1.1.6.MOV: 00:02-00:09, 00:23-00:25

  - Next, obtain the 3D-printed outlet connector and insert the magnet into its groove [1]. Place this assembly beneath the drape through the outflow spout and align the cut hole with the hole in the connector and the magnet [2].

    - 1.1.7.MOV: 00:03-00:06

    - 1.1.7.MOV: 00:10-00:15, 00:20-00:25

  - Raise the excess drape to cover the drum funnel and secure it using rubber bands or elastic ties [1]. Fold the excess drape over the top to establish a sterile barrier [2]. Now, place the second water jacket on top of the folded drape [3]. Then, take a 30-inch piece of one-quarter inch tubing and connect the proximal end to the outlet at the base of the drum funnel spout and the distal end to the reservoir inflow barb [4].

    - 1.1.8.MOV: 00:07-00:14.

    - 1.1.8.MOV: 00:17-00:23

    - 1.1.8.MOV: 00:26-00:29

    - 1.1.8.MOV: 00:05-00:16.

  - Now, place the pumps on a mobile cart positioned nearest to the workstation edge [1]. Position the stand and two clamps directly behind the pumps so the clamps are elevated 6 inches above the pumps [2].

    - 1.2.1.MOV: 00:02-00:06

    - 1.2.1.MOV: 00:08-00:13.

  - Further, obtain two pump heads and secure each into one of the pumps [1]. Place two oxygenators into the clamps on the stand in a horizontal orientation [2]. Then, connect the arterial and venous one-eighth inch gas inflow tubing to the top right Luer connector of their respective oxygenators [3].

    - 1.2.2.MOV: 00:04-00:11

    - 1.2.3.MOV: 00:03-00:13

    - 1.3.3.MOV: 00:05-00:13

  - Now, remove all tubing from the hemodialysis kit except for the lines leading directly into the filter [1].

    - 1.4.1.MOV: 00:05-00:14

  - Place one large roller pump immediately to the right of the hemodialysis filter for blood inflow [1]. Then, place two smaller peristaltic pumps to the right of the blood roller pump [2]. Obtain two one-eighth-inch peristaltic tubings and insert them into the white pump clamp. Attach a female one-eighth-inch Luer barb to both ends of each tubing [3].

    - 1.4.2.MOV: 00:03-00:04

    - 1.4.2.MOV: 00:05-00:07.

    - 1.4.2.1.MOV: 00:05-00:14

  - To connect the dialysate bag to the filter inflow, attach one-eighth inch tubing to the Luer connector on the bag [1]. Attach a male one-eighth inch Luer to the tubing's end and connect it to the inflow peristaltic pump tubing [2]. Place the reservoir inside the reservoir holder with the exterior measurement markings and the yellow-capped connector facing outward [3].

    - 1.4.3.1.MOV: 00:03-00:10

    - 1.4.3.1.MOV: 00:10-00:14

    - 1.5.1.MOV: 00:04-00:10.

  - Next, cut a 36-inch piece of five-sixteenth-inch tubing, attach a Y connector to the distal end of this tubing, and connect the proximal end to the outflow spigot at the bottom of the reservoir [1].

    - 1.5.2.MOV: 00:08-00: 18

  - Cut two 4-inch pieces of five-sixteenth inch tubing and attach each to the end of the Y connector [1]. Connect the other ends of these tubes to the horizontal prong on each of the pump heads [2].

    - 1.5.3.MOV: 00:03-00:17

    - 1.5.3.MOV: 00:19-00:24

  - Then, cut two 4-inch pieces of one-quarter inch tubing and connect one to each vertical barb on the pump heads. Attach the other end to the left lower prong of each oxygenator [1].

    - 1.5.4.MOV: 00:03-00:20

  - Now, take a 30-inch segment of five-sixteenth inch tubing, cut the tubing in half, and place a five-sixteenth inch Y connector over the distal cut end [1]. Attach the proximal end to the outflow barb of the venous oxygenator. Reattach the other half to the top barb of the Y connector [2].

    - 1.6.1.MOV: 00:02-00:05, 00:12-00:15

    - 1.6.1.+.MOV: 00:19-00:21

  - For the side sampling circuit, cut a 2-inch piece of one-quarter inch tubing and, using a needle driver, stretch the proximal end [1]. Attach this stretched tubing to the remaining barb of the five-sixteenth inch Y connector [2].

    - 1.6.2.MOV: 00:02-00:07

    - 1.6.2.MOV: 00:12-00:16.

  - After connecting the cuvette shunt sensor, connect the horizontal part of a female T connector to a one-eighth inch tubing fitted with a male Luer. Attach a sampling clave to the perpendicular aspect of the T connector [1]. Finally, connect a one-eighth inch male Luer to the other horizontal side of the T connector [2].

    - 1.6.2.3. MOV: 00:03-00:13

    - 1.6.2.3. MOV: 00:14-00:20

  - Then, cut a 12-inch length of one-eighth inch tubing and connect the proximal end to the sampling clave [1]. Attach a male one-eighth inch Luer to the distal end and connect it to a second female T connector [2]. To one side of this T connector, attach a one-quarter-inch male Luer and use a 2-inch piece of one-quarter inch tubing to connect it to an inflow channel [3].

    - 1.6.2.3. MOV: 00:23-00:25

    - 1.6.2.3. MOV: 00:27-00:31

    - 1.6.2.3. MOV: 00:32-00:37

  - At the distal end of the five-sixteenth inch tubing, attach another five-sixteenth inch Y connector [1]. Obtain approximately 55 inches of one-quarter inch tubing and connect it to the top barb of the Y connector [2].

    - 1.6.3.MOV: 00:03-00:08

    - 1.6.3.MOV: 00:14-00:23

  - For the hemodialysis blood circuit, attach a one-quarter-inch male Luer to the distal end of the one-quarter inch tubing [1]. Thread this tubing through the large roller pump [2] and connect it to the blood inflow port at the top of the hemodialysis filter, identified by red-lined tubing [3].

    - 1.6.4.MOV: 00:02-00:11

    - 1.6.4.MOV: 00:12-00:19

    - 1.6.4.MOV: 00:26-00:31

  - Next, take a 24-inch piece of one-quarter inch tubing and attach a male Luer to the proximal end [1]. Connect this tubing to the blue-lined blood outflow port of the hemodialysis filter. Then, attach the distal end to the inflow barb of the reservoir [2].

    - 1.6.4.1.MOV: 00:01-00:09

    - 1.6.4.1.MOV: 00:10-00:18

  - Take a 36-inch piece of one-quarter inch tubing and stretch it to fit securely onto the barb of the five-sixteenth inch Y connector [1].

    - 1.6.5.MOV: 00:01-00:12.

  - Next, obtain three four-way stopcocks and connect them in a linear sequence [1]. Attach this assembly to the perpendicular port of a female T connector [2]. Insert this component 12 inches from the left side of the one-quarter inch tubing [3].

    - 1.6.6.MOV: 00:02-00:05

    - 1.6.6.MOV: 00:06-00:12

    - 1.6.6.MOV: 00:13-00:18

  - Then, attach external flow sensors to the one-quarter inch tubing immediately downstream of the drug infusion site [1]. Make a cut 2 inches distal to the flow sensors and insert pressure sensors with barbed ends into the one-quarter inch tubing. Place the distal portion of the tubing into the center of the drum funnel [3].

    - 1.6.7.MOV: 00:04-00:12

    - 1.6.7.part2MOV:00:07-00:20

  - Next, cut a 30-inch length of five-sixteenth inch tubing and connect its proximal end to the right outflow barb of the arterial oxygenator [1]. Cut this tubing at the midpoint and insert a five-sixteenth inch Y connector [2-TXT].

    - 1.7.1.MOV: 00:05-00:10

    - 1.7.1.MOV: 00:13-00:25   TXT: Create a duplicate side sampling circuit

  - Then, cut a 36-inch segment of one-quarter inch tubing and connect it to the barb of the five-sixteenth inch Y connector [2]. Next, cut a 30-inch segment of one-quarter inch tubing, attach a one-quarter inch male Luer to the proximal end, and connect it to the inferior vena cava outflow of the specimen [3].

    - 1.7.3.MOV: 00:01-00:10

    - 1.8.1.MOV: 00:02-00:09

  - Cut the tubing 6 inches from the reservoir and attach one-quarter-inch male Luer to both ends. Insert a female T connector with a sampling clave on the perpendicular port [1].

    - 1.8.2.MOV and 1.8.2.final.machine.jpeg

  - Introduce perfusate into the circuit following the previously described method. Rewarm and connect the specimen and begin perfusion [1-TXT].

    - 1.9.MOV: 00:11-00:25 TXT: Refer to the manuscript for specimen connection

# Results

Please review this section to make sure that it accurately reflects your findings.

- You/Your videographer does not have to record this section. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 181.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - In the first 24 hours of porcine liver perfusion, hepatic artery and portal vein flows averaged 107 milliliters per minute and 501 milliliters per minute [1], respectively, with stable inflow pressures of 80 and 7.1 millimeters of mercury [2].

    - LAB MEDIA: Figure 2A.

    - LAB MEDIA: Figure 2B.

  - Arterial, portal, and venous oxygen saturations in porcine livers averaged 95%, 85%, and 74% [1], with pH remaining near 7.2 [2].

    - LAB MEDIA: Figure 2C.

    - LAB MEDIA: Figure 2D.

  - Glucose levels decreased steadily [1], hematocrit and hemoglobin remained stable [2], and potassium dropped slightly [3]. Porcine livers produced 41 milliliters of bile [4] and showed moderate lactate elevation [5].

    - LAB MEDIA: Figure 2F.

    - LAB MEDIA: Figure 2E. Video editor: Show the top panel for hematocrit

    - LAB MEDIA: Figure 2E. Video editor: Show the bottom panels for hematocrit and potassium

    - LAB MEDIA: Figure 2H.

    - LAB MEDIA: Figure 2G.

  - In human perfusion, arterial and venous inflows averaged 113 and 317 milliliters per minute [1], and oxygen saturation reached 98%, 92%, and 86% in arterial, portal, and hepatic veins [2].

    - LAB MEDIA: Figure 3A.

    - LAB MEDIA: Figure 3C.

  - Human livers maintained physiologic pH [1], stable hematocrit and hemoglobin [2], and showed glucose decline and lactate increase [3], with bile production averaging 1.4 milliliters per hour [4].

    - LAB MEDIA: Figure 3D.

    - LAB MEDIA: Figure 3I.

    - LAB MEDIA: Figure 3F–G.

    - LAB MEDIA: Figure 3H.

  - Histology confirmed viable hepatic parenchyma and tumor morphology after perfusion [1].

    - LAB MEDIA: Figure 3J–L.

  - A mobile setup enabled CT (C-T) imaging of perfused livers, yielding radiographic quality comparable to in vivo scans [1]. The system also supported kidney and pancreas perfusion with preserved glomeruli and Islets of Langerhans [2].

    - LAB MEDIA: Figure 4B.

    - LAB MEDIA: Figure 4C–F.