# Submission ID #: 68767

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20973983

# Title: TD-DFT Guided Advanced E-Eye Sensing Technique for On-site Quantification of Fe, Cr, F, and As in the Environmental, Biological, and Food Samples

# Authors and Affiliations:

Shrikant Kashyap1*, Rajasekhar Ravula2*, Neha Majee1, Tapas K Mandal1

1Department of Chemical Engineering, and Center for Nanotechnology, Indian Institute of Technology Guwahati

# 2Department of Chemical Engineering, K.K. Wagh Institute of Engineering Education and Research

*These authors contribute equally

# Corresponding Authors:

Tapas K Mandal	tapasche@iitg.ac.in

# Email Addresses for All Authors:

Shrikant Kashyap 	shrikant2016@iitg.ac.in

Rajasekhar Ravula	rajravula@kkwagh.edu.in

# Neha Majee 		n.majee@iitg.ac.in

Tapas K Mandal	tapasche@iitg.ac.in

## Author Questionnaire

2. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No.

3. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, done.

Current Protocol Length

Number of Steps:  25

Number of Shots:  52  (19 SC)

# Introduction

Note: The script was filmed using the draft and their edits, the interview answers can be trimmed based on the answers here

  - Tapas Mandal:

The scope is to develop an indigenous portable device, named as E-Eye, for estimating of the toxicants in a multi-matrix sample.

What are the current experimental challenges?

  - Tapas Mandal:

The major experimental challenges with the POCT devices:

- Getting the selectivity of a targeted toxicant within the complex sample matrix, such as industrial effluents

- Detection of very low concentration

- The simultaneous detection of multiple analytes.

What significant findings have you established in your field?

1.5    Tapas Mandal:

The indigenously developed multiplexed 'E-Eye' device mimics the portable UV-Vis spectrophotometer. Here, E-Eye stands for “Optical sensor”.

This user-friendly Point-Of-Care-Testing (POCT) device comprises the followings:

- Working principle of a chemical-based colorimetric sensor 		with high selectivity

- Working principle of an optical sensor with high precision

- Fundamentals of both the principles;-  TD-DFT approach for chemical sensor

What new scientific questions have your results paved the way for?

1.9,  Tapas Mandal: Our results answer the following fundamental science questions:

- The pathway from computational exercise to translation technology for precise detection of the toxicants

- Sensing mechanism

- Selectivity

- Range of detection and

What research questions will your laboratory focus on in the future?

1.10,    Tapas Mandal: In the future, we will focus on computationally guided advanced POCT device design employing infrared spectroscopic principles. It will cover the following areas:

- Water quality management for the community services

- Detection of various noxious elements/compounds/biomarkers including the VOCs in:

- Agricultural applications

- Environmental applications

- Healthcare applications.

# Protocol

- Computing and Validation of the Reaction's Sensing Mechanism

Demonstrator: Shrikant Kashyap

  - To begin, select a lock-and-key reaction that specifically recognizes the targeted analyte based on its fundamental chemistry [1].

    - WIDE: Talent taking a seat at the computer station.

  - Launch the Gaussian 09 program to perform all the calculations using the Becke, 3-parameter, Lee-Yang-Parr (B3LYP) (B3-L-Y-P) hybrid method, and the Los Alamos National Laboratory 2 Double ζ (zee) basis set to obtain more accurate results in electronic computation involving the targeted analyte [1-TXT].

    - SCREEN: 2.2.1.revised.mp4  00:00:39 to 00:00:59. TXT: Analyte chosen here: Iron

Authors: We need only a 20-second clip for each shot showing the calculations

  - Activate the self-consistent reaction field using the polarizable continuum model to include solvent effects and set the solvent parameter to water [1].

    - SCREEN: 2.3.1.	SCREEN: 2.3.1 &2.4.1.mp4 00:20-00:40 .

  - In the Gaussian input, add SCRF=PCM and choose Water and confirm the setting and show the input containing geom=connectivity and bonded atom list in the connectivity block. Click Submit [1]. Compute the ground-state energy required to compute the optimized structure of the molecule [2].

    - SCREEN: 2.3.1 &2.4.1.mp4 00:30-00:55.

    - SCREEN: 2.4.2 00:10-00:20.

  - Obtain excited-state electronic transitions using the time-dependent self-consistent field method, requesting a number of states equal to 60 [1]. With the same method and solvent model, compute and export the predicted ultraviolet absorption spectrum [2].

    - SCREEN: 2.5.1 00:30-00:50.

    - SCREEN: 2.5.2. 00:20-00:40.

  - Perform a ground-state density functional theory geometry optimization using the B3LYP functional with the LanL2DZ (lan-L-2-D-zee) basis set to refine the molecular structure [1]. Export the optimized structures of 1,10-phenanthroline (1-ten-phen-anthroline) and the ferrous ion for subsequent studies [1-TXT].

    - SCREEN: 2.3.1 &2.4.1.mp4 00:10-00:25

    - SCREEN: 2.5.2. 00:12-00:18. TXT: Similarly, simulate the molecule structure of hexafluoroferrate (FeF6)

- Development of a Colorimetric Chemical Sensor

Demonstrator: Neha Majee

  - Transfer equal volumes of trisodium citrate dihydrate as masking agent, ascorbic acid as pH optimising agent, and 1,10-phenanthroline into a transparent glass culture tube [1-TXT]. Add 2 milliliters of stock solution to reach a total volume of 5 milliliters [2]

    - Talent pipetting equal volumes of each reagent. TXT: Add 1 mL of each reactant into a 5 mL tube

    - Talent adding 2 milliliters of the stock solution and gently mixing.

  - Place each sample and standard in a quartz cuvette with a 1-centimeter path length and keep all measurements at room temperature under standard laboratory conditions [1]. Measure iron in all standards and samples using UV–visible spectroscopy by scanning from 250 to 750 nanometers [2 AND 3].

    - Talent filling quartz cuvettes with the sample.

    - Placing the sample in the UV spectrophotometer. NOTE: 3.2.2 AND 3.2.3 MAY HABE BEEN COMBINED

    - Talent adjusting the settings in the instrument system.

  - Configure the instrument settings to match the specified conditions by selecting Scan speed as Medium [1]. Set Measurement type to Absorbance [2], adjust Slit width to 5 nanometers [3] and set the Detectors to Direct [4].

    - SCREEN: Open Scan settings and choose Medium under Scan speed. NOTE: All the screen captures in the section 3 are shot by the videographer, please use them

    - SCREEN: In Measurement mode, select Absorbance.

    - SCREEN: In Optics, set Slit width to 5.0 nanometers and confirm.

    - SCREEN: In Detector options, select Direct and apply the configuration.

  - Perform baseline correction using a blank that contains the same water matrix as the samples [1].

    - SCREEN: Select Baseline Correction, insert the matrix blank, and confirm Apply baseline.

  - After UV-absorbence is recorded, lambda max is observed at 510 nm which confirms the formation of the complex [1].

    - Cuvette is placed in the UV-Vis spectrophotometer and showing 510 nm as peak.

- Design and Fabrication of the E-Eye-Enable POCT Kit's Prototype

Demonstrator: Shrikant Kashyap

  - To begin, procure a LDR, a white LED, a LCD, a microcontroller or microprocessor board, and precision resistors [1].

    - Talent laying out all components on an electrostatic discharge-safe mat and checking them against a bill of materials.

  - Using the three-dimensional printed box, align the LED and the light-dependent resistor directly opposite each other across the cuvette slot [1].

    - Talent mounting the LED on one wall of the box and the light-dependent resistor on the opposite wall, facing each other.

  - Complete the circuit by installing a 3.3 kilo-ohm series resistor for the LED, routing clean wiring, and connecting the LCD model number 11497 [1].

    - Talent soldering the 3.3 kiloohm resistor to the LED lead and connecting power and ground.

  - Interface the sensor with the microcontroller across a voltage divider that converts light-dependent resistor resistance changes into a measurable voltage [1]. Using the Arduino Integrated Development Environment version 1.8.19 (1 point 8 point 19), set up data acquisition to read analog inputs from analog pins A0 to A3 [2].

    - Talent wiring the light-dependent resistor in a voltage divider configuration between the supply and an analog input, with the reference resistor to ground.

  - Install analog noise-suppression filters by adding a 10 kiloohm series or biasing resistor and a 1 microfarad capacitor to form a low-pass network at each analog input [1]. Place components close to the microcontroller pins and route short, twisted signal leads to improve stability [2].

    - Talent soldering the 10 kiloohm resistor and the 1 microfarad capacitor to the analog input node. NOTE: 4.5.1 and 4.5.2 were combined

    - Close-up of the filtered analog lines being secured near the microcontroller headers.

  - Print the printed circuit board when the prototype schematic and layout are finalized and electrically verified [1].

    - Shot of printing in progress.

  - Supply power by turning on the electronic switch and record the sensor response as resistance or converted voltage from each channel [1]. Observe the light path as the LED passes light through the reaction chamber and the light-dependent resistor senses transmitted intensity corresponding to ferroin formation [2]. Correlate the measured resistance with ferrous ion concentration to generate a calibration plot, and enter the resulting equation into the device algorithm for real-time reporting [3].

    - Talent flipping the main power switch and confirming the power indicator.

    - Close-up of the illuminated LED shining through a cuvette containing the reaction mixture toward the light-dependent resistor.

    - SCREEN: 4.7.4.

  - Review the complete circuit drawing to confirm final wiring and channel assignments, then print and assemble the finished device enclosure [1].

    - SCREEN: 4.8.1.

- Development of a Multiplexed Device

  - Create four dedicated slots in the three-dimensional box to hold four cuvettes in fixed positions [1]. Isolate each detection slot both optically with opaque physical barriers and electronically with separate LED and light-dependent resistor pairs and individual voltage dividers to prevent signal overlap [2].

    - Talent inserting opaque partitions between cuvette bays inside the enclosure.

    - Talent wiring four independent voltage divider circuits, one per channel, and confirming isolation with a multimeter.

  - Assign one slot to a single analyte and keep the assignment fixed [1]. Ensure the cuvette placed in each slot contains only the reagents specific to that analyte [2].

    - Talent labeling each slot with the analyte name and affixing matching labels to the corresponding cuvettes.

    - Talent inserting analyte-specific cuvettes into their designated slots.

  - Repeat the interfacing, programming, filtering, printing, powering, and calibration steps for the four-channel configuration, and connect all four LED–light-dependent resistor pairs in parallel at the board level while maintaining separate voltage dividers per channel [1]. Verify that code reads all four analog channels and that wiring supports future expansion to additional slots if needed [2].

    - Talent connecting four sensor channels to the board’s analog inputs and tying power and ground busses in parallel.

    - SCREEN: Update the Arduino sketch to read analog pins A0 to A3, compute concentration with the stored calibration for each channel, and upload the program.

  - Turn on the main switch to supply power so that all four LEDs illuminate simultaneously [1]. Allow the light-dependent resistors to sense each analyte in parallel and display the results for all channels at once [2].

    - Talent pressing the power switch and confirming all channel indicators illuminate.

    - Show the multi-channel readout with four concurrent measurements and corresponding calculated concentrations NOTE: 5.41. AND 5.4.2 WERE COMBINED

# Results

- Results

  - The multiplexed prototype accurately quantified total iron concentrations in peanut samples, with results closely matching the UV-Vis spectrophotometer [1] and remaining above the limit of detection in all cases [2].

    - LAB MEDIA: Figure 7A. Video editor: Highlight the green bars.

    - LAB MEDIA: Figure 7A. Video editor: Highlight the red horizontal line.

  - The chromium and fluoride levels measured in hair samples using the multiplexed prototype were comparable to UV-Visible measurements[1].

    - LAB MEDIA: Figure 7B and C. Video editor: Highlight the green bars.

  - Arsenic levels in Brahmaputra water samples increased progressively across sample numbers and were reliably detected by both the multiplexed prototype and UV-Vis system [1].

    - LAB MEDIA: Figure 7D. Video editor: Sequentially Highlight the green bars

  analyte
Pronunciation link: No confirmed link found
IPA: /ˈæn·əˌlaɪt/
Phonetic Spelling: AN-uh-lyt

  hybrid (as in “hybrid method”)
Pronunciation link: https://www.merriam-webster.com/dictionary/hybrid
IPA: /ˈhaɪ·brɪd/
Phonetic Spelling: HY-brid

  basis set
Pronunciation link: No confirmed link found (as a compound phrase; each word separately is common)
IPA: /ˈbeɪsɪs ˌsɛt/
Phonetic Spelling: BAY-sis set

  polarizable (as in “polarizable continuum model”)
Pronunciation link: No confirmed link found
IPA: /ˌpə·lər·əˈzaɪ·bəl/
Phonetic Spelling: puh-LER-uh-ZY-buhl

  continuum
Pronunciation link: https://www.merriam-webster.com/dictionary/continuum
IPA: /kənˈtɪn·ju·əm/
Phonetic Spelling: kun-TIN-you-um

  geometry (in computational-chemistry context: molecular geometry)
Pronunciation link: https://www.merriam-webster.com/dictionary/geometry
IPA: /dʒiˈɑ·mə·tri/
Phonetic Spelling: jee-AH-muh-tree

  optimization
Pronunciation link: https://www.merriam-webster.com/dictionary/optimization
IPA: /ˌɑp·təˌmɪˈzeɪ·ʃən/
Phonetic Spelling: op-tuh-mi-ZAY-shun

  ferrous
Pronunciation link: https://www.merriam-webster.com/dictionary/ferrous
IPA: /ˈfɛr·əs/
Phonetic Spelling: FERR-us

  phenanthroline (1,10-phenanthroline)
Pronunciation link: No confirmed link found
IPA (approximate): /ˌfɛn·ænˈθroʊˌliːn/
Phonetic Spelling: fen-an-THROH-leen

  hexafluoroferrate (FeF₆ complex)
Pronunciation link: No confirmed link found
IPA (approximate): /ˌhɛk·səˌflʊə·roʊˈfɛr·eɪt/
Phonetic Spelling: HEK-suh-fluh-roh-FERR-ate

  trisodium (as in “trisodium citrate dihydrate”)
Pronunciation link: No confirmed link found
IPA: /traɪˈsoʊ·di·əm/
Phonetic Spelling: try-SOH-dee-um

  citrate
Pronunciation link: https://www.merriam-webster.com/dictionary/citrate
IPA: /ˈsɪtrɪt/ or /ˈsɪ·treɪt/
Phonetic Spelling: SIH-trit or SIH-trayt

  dihydrate
Pronunciation link: https://www.merriam-webster.com/dictionary/dihydrate
IPA: /ˌdaɪˈhaɪ·dreɪt/
Phonetic Spelling: dye-HY-drate

  cuvette
Pronunciation link: https://www.merriam-webster.com/dictionary/cuvette
IPA: /kjuˈvɛt/
Phonetic Spelling: kew-VET

  nanometer (as in “nanometers” for spectral range)
Pronunciation link: https://www.merriam-webster.com/dictionary/nanometer
IPA: /ˈnæ·noʊˌmi·tər/
Phonetic Spelling: NAN-oh-mee-ter

  absorbance
Pronunciation link: https://www.merriam-webster.com/dictionary/absorbance
IPA: /əbˈzɔr·bəns/ or /əbˈzɔːr·bəns/
Phonetic Spelling: ub-ZOR-buhns

  resistor
Pronunciation link: https://www.merriam-webster.com/dictionary/resistor
IPA: /rɪˈzɪs·tər/
Phonetic Spelling: ri-ZIS-ter

  microcontroller
Pronunciation link: No confirmed link found
IPA (approximate): /ˌmaɪ·kroʊ·kənˈtroʊ·lər/
Phonetic Spelling: MY-kroh-kun-TROH-ler

  capacitor
Pronunciation link: https://www.merriam-webster.com/dictionary/capacitor
IPA: /kəˈpæs·ɪ·tər/
Phonetic Spelling: kuh-PAS-ih-ter

  multimeter
Pronunciation link: https://www.merriam-webster.com/dictionary/multimeter
IPA: /ˌmʌl·taɪˈmi·tər/
Phonetic Spelling: mul-tye-MEE-ter

  baseline (as in “baseline correction”)
Pronunciation link: https://www.merriam-webster.com/dictionary/baseline
IPA: /ˈbeɪsˌlaɪn/
Phonetic Spelling: BAY-s-line

  lambda (as in “λ-max” / “lambda max”)
Pronunciation link: https://www.merriam-webster.com/dictionary/lambda
IPA: /ˈlæm·də/
Phonetic Spelling: LAM-duh