# Submission ID #: 68788

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20980388

# Title: Updated Protocol for the Assembly and Use of the Minibioreactor Arrays (MBRAs)

# Authors and Affiliations:

# Jason D. Pizzini1, Firas S. Midani1,2, Robert A. Britton1,2

# 1Department of Molecular Virology and Microbiology, Baylor College of Medicine

# 2Alkek Center for Metagenomics & Microbiome Research, Baylor College of Medicine

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

# Jason D. Pizzini                              Jason.Pizzini@bcm.edu

# Email Addresses for All Authors:

# Firas S. Midani                              Firas.Midani@bcm.edu

# Robert A. Britton                          Robert.Britton@bcm.edu

# Jason D. Pizzini                              Jason.Pizzini@bcm.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps:  25

Number of Shots:  55

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer the 1st REQUIRED question and at least 2 other questions (1.2 – 1.10) below. Up to 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- If possible, each author should deliver no more than two statements.

- Answer in full sentences, in a style suitable for being spoken aloud.

- Limit the length of each statement to 30 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

What research questions will your laboratory focus on in the future?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 30 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

# Videographer: Please capture all testimonial shots in a wide-angle format with sufficient headspace, as the final videos will be rendered in a 1:1 aspect ratio. Testimonial statements will be presented live by the authors, sharing their spontaneous perspectives.

- Testimonial statements will not appear in the video but may be featured in the journal’s promotional materials.

- Provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author delivering the testimonial.

- Please answer the testimonial question live during the shoot, speaking naturally and in your own words in complete sentences.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Enter author name, Enter author title: (authors will present their testimonial statements live)

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Enter author name, Enter author title: (authors will present their testimonial statements live)

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Preparation for Minibioreactor Array (MBRA) Assembly

Demonstrator: Click here to enter name of demonstrator(s)

  - To begin, ensure that the Minibioreactor array strips are three-dimensional printed and contain 6 independent bioreactor chambers [1-TXT]. Using a one-quarter inch-28NF fraction tap with a T-handle tap wrench, thread the three one-quarter inch ports in each chamber to insert fittings [2].

    - WIDE: Talent showing a completed 3D-printed Minibioreactor array strip with 6 chambers. TXT: MBRA: Minibioreactor Array

    - Talent threading the ports using a one-quarter inch-28NF fraction tap with a T-handle tap wrench.

  - After washing the chamber with water, place a 10 by 3-millimeter magnetic stir bar into each chamber and add 1 milliliter of distilled water [1]. Then, position a rubber washer on top of each port of the bioreactor [2].

    - Talent placing a magnetic stir bar and pouring water into a chamber.

    - Talent placing a rubber washer onto a port.

  - For each chamber, screw in one waste straw-threaded male luer, one media straw-threaded male luer, and one empty threaded male luer into the ports [1].

    - Talent screwing the specified luers into the ports.

  - Now, insert 6 rubber septa onto three thirty-second-inch female luer barbs [1] and fold the upper sleeve of each septum down to cover the neck [2]. Attach these to the designated ports of each chamber [3]. Attach a one-eighth inch female luer barb to one end and a male luer lock connector to the opposite end of each length [4].

    - Talent placing rubber septa onto female luer barbs.

    - Talent folding the septum sleeve over the neck.

    - Talent attaching the septa to the designated ports. TXT: Cut C-flex tubing strips of desired lengths

    - Talent attaching a female luer barb and a male luer lock connector to each tubing length.

  - Then, insert a one-sixteenth-inch female luer barb into each end of the red two-stop E-lab tubing with a 1.14-millimeter inner diameter [1-TXT] and the orange two-stop E-lab tubing with a 0.89-millimeter inner diameter [2-TXT].

    - Talent inserting one-sixteenth inch female luer barb into  end of the red two-stop E-lab tubing. TXT: Briefly submerge the E-lab tubing’s ends in near-boiling water to soften them

    - Talent inserting into orange two-stop E-lab tubing with 0.89 millimeter inner diameter.  TXT: Repeat this process 6x for each MBRA strip

  - After connecting the prepared E-lab tubing to the C-flex tubing, ensure each of the 6 C-flex tubing lengths is connected to one red and one orange E-lab line via female luers [1].

    - Talent connecting orange E-lab tubing to the other end of the same C-flex tubing.

  - Attach a one-eighth inch female luer barb and a male luer lock connector to both ends of one 3-inch piece and the 12-inch piece of C-flex tubing [1]. Attach male luer lock connectors to both ends of the remaining pieces [2].

    - Talent attaching a female luer barb and a male luer lock connector to the 3 inch and 12 inch tubing.

    - Talent attaching male luer lock connectors to both ends of the remaining tubing.

  - Assemble the waste line tree according to the 3D diagram [1].

    - LAB MEDIA: Figure 3B.

  - Attach the exposed ends of the red two-stop E-lab tubing to the terminal male luer locks on the waste line tree in ascending order based on C-flex tubing length [1]. Then, connect the 3-inch C-flex tubing with the one-eighth-inch female luer barb and male luer lock connector to the top of the waste line tree [2].

    - Talent attaching red E-lab tubing to the waste line tree.

    - Talent securing the 3 inch C-flex tubing to the top of the assembly.

  - Assemble the feed line tree according to the 3D diagram [1].

    - LAB MEDIA: Figure 3B.

  - Bridge the exposed ends of the orange two-stop E-lab tubing to the terminal male luer locks on the feed line tree in ascending order based on C-flex tubing length [1] and attach the 12-inch C-flex tubing to the top of the feed line tree [2].

    - Talent attaching orange E-lab tubing to the feed line tree.

    - Talent securing the 12 inch C-flex tubing to the top of the assembly.

- MBRA Assembly Using the Prepared Components

Demonstrator: Click here to enter name of demonstrator(s)

  - Attach the variable-length C-flex tubing at the end of the feed line tree to the bioreactor [1], arranging them in ascending order from the shortest line on the left to the longest on the right [2].

    - Talent attaching the C-flex feed lines to the bioreactor ports.

    - Close-up of the feed lines arranged from shortest on the left to longest on the right.

  - Attach the variable-length C-flex tubing at the end of the waste line tree to the bioreactor strip in descending order [1], with the longest line on the left and the shortest on the right to accommodate pump placement [2]. Bundle all C-flex feed lines together on the left side of the strip and secure them with a twist tie [3-TXT].

    - Talent connecting waste lines to the bioreactor.

    - Close-up showing waste lines arranged from longest on the left to shortest on the right.

    - Talent bundling and tying the feed lines. TXT: Repeat the process for the waste lines on the right side

  - Form a loop with the orange two-stop E-lab tubing between the C-flex lines and secure the loop using autoclave tape [1] and repeat the process for the red two-stop E-lab tubing on the waste side of the bioreactor strip [2].

    - Talent looping and securing the orange tubing with autoclave tape.

    - Talent looping and securing the red tubing with autoclave tape.

  - Cover the female luer at the end of the waste and feed line trees with foil to prevent contamination [1]. Loosen the male threaded luers with septa on each bioreactor chamber to allow steam to escape during autoclaving [2].

    - Talent wrapping foil over the female luers.

    - Talent loosening threaded luers on each chamber.

  - After placing the assembly into an autoclave bin, stretch out the feed and waste line trees into separate bins adjacent to the one containing the MBRA strips [1-TXT].

    - Talent stretching feed and waste line trees into separate bins. TXT: Autoclave: 121 °C, ≥ 15 psi for 25 min; Once cooled, retighten the threaded male luers with the septa

- MBRA Connection and Operation

Demonstrator: Click here to enter name of demonstrator(s)

  - To attach the system to the pumps, remove the autoclave tape securing the E-lab tubing for both waste and feed lines [1] and untie the bundles of C-flex tubing [2]. Position the MBRA between the two pumps on top of the stir plate [3]. Clamp it down using the 3D-printed holders and align it with the marked stirring positions on the plate [4].

    - Talent peeling off autoclave tape from the E-lab tubing.

    - Talent untying bundled C-flex tubing.

    - Talent placing the MBRA between two pumps.

    - Talent clamping the MBRA into position on the stir plate.

  - Now, attach the feed line E-lab tubing to the peristaltic pump cartridges [1] and position the tubing stops into the cartridge slots [2]. Repeat the process for the waste line E-lab tubing on the pump located to the right of the stir plate [3]. Then, lock the peristaltic pump cartridges into the pump [4-TXT].

    - Talent attaching feed line E-lab tubing to pump cartridges.

    - Close-up of tubing stops seated in the cartridge slots.

    - Talent attaching waste line E-lab tubing to the right-hand pump.

    - Talent locking cartridges into the pump. TXT: Ensure the cartridges are seated fully

  - Arrange the C-flex tubing neatly using the 3D-printed tube holders [1], followed by the end of the waste line tree to the tubing connected to the waste bottles [2]. Next, attach the female luer on the feed line entry tubing to the male connector on the 12 inch tube from the media bottle cap [3].

    - Talent placing C-flex tubing into the 3D-printed tube holders.

    - Talent connecting waste line tree tubing to waste bottle tubing.

    - Talent connecting feed line entry tubing to media bottle cap tubing.

  - Turn on both pumps to start media flow [1] and ensure both pumps are set to clockwise rotation when waste is positioned to the right of the pumps [2].

    - Talent powering on both pumps.

    - Close-up of pump settings showing clockwise rotation.

  - Observe the droplet size and cadence in each bioreactor chamber [1]. If variability is seen, replace the orange two-stop E-lab tubing connected to the affected chamber to reduce flow rate variation [2]. Once the chambers are full, shut off both pumps and allow the bioreactors to sit for 24 to 48 hours to check for contamination before starting the experiment [3].

    - Close-up of media droplets falling into a chamber.

    - Talent replacing an orange E-lab tubing line.

    - Talent switching off the pumps.

- MBRA Disassembly and Refurbishment

Demonstrator: Click here to enter name of demonstrator(s)

  - Switch the media input to a 1 liter container of 10 percent bleach in deionized water [1] and increase the flow rate on both pumps to maximum to displace the contents of the bioreactor chambers with bleach [2].

    - Talent replacing the media input line with a container of bleach solution.

    - Shot of Pump control panel showing the flow rate being increased to maximum.

  - Once the chambers are clear of media, invert the MBRA to disinfect above the fill line for 5 minutes [1]. After 5 minutes, right the system and wait an additional 5 minutes for sterilization [2].

    - Talent inverting the MBRA assembly.

    - Talent placing the MBRA upright on the bench.

  - After replacing the bleach solution with deionized water, flush the system until 1 liter of water has passed through [1]. Then, disconnect the bioreactor E-lab tubing from the pumps and remove the MBRAs [2].

    - Talent observing the system flushing with water.

    - Talent detaching the E-lab tubing from the pumps and removing the MBRA assembly.

  - Finally, remove the used septa from the chambers [1]. Drain each chamber until only 1 milliliter of water remains [2] and refurbish the MBRA as required [3].

    - Talent removing septa from the chamber ports.

    - Talent tilting the MBRA to drain water, leaving a small volume behind.

    - TEXT ON PLAIN BACKGROUND:

- Autoclave up to 3 reuse cycles

- After 3rd reuse: Fully disassemble MBRA

Replace C-flex tubing and sterilize parts

Replace if damaged

- Reapply epoxy to waste and feed PTFE connections if brittle

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 84.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - A human fecal sample was prepared and grown in the MBRA  system. After four days of continuous flow [1], the microbial community in all nine bioreactors was dominated by 18 bacterial genera, each comprising at least 2% of relative abundance in any replicate [2].

    - LAB MEDIA: Figure 5A.

    - LAB MEDIA: Figure 5A. Video editor: Sequentially Highlight the colored stacked bars in the chart.

  - Twenty-two out of the 65 detected genera were present in all nine bioreactor replicates, demonstrating high reproducibility [1].

    - LAB MEDIA: Figure 5A.

  - Alpha diversity analysis showed minimal variation between replicates [1] in both the observed operational taxonomic units [2] and Shannon diversity index [3].

    - LAB MEDIA: Figure 5B.

    - LAB MEDIA: Figure 5B. Video editor: Highlight the top panel showing “Observed OTUs” .

    - LAB MEDIA: Figure 5B. Video editor: Highlight the bottom panel showing “Shannon Diversity” .