# Submission ID #: 68788

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=20980388

# Title: Updated Protocol for the Assembly and Use of the Minibioreactor Arrays (MBRAs)

# Authors and Affiliations:

# Jason D. Pizzini1, Firas S. Midani1,2, Robert A. Britton1,2

# 1Department of Molecular Virology and Microbiology, Baylor College of Medicine

# 2Alkek Center for Metagenomics & Microbiome Research, Baylor College of Medicine

# Corresponding Authors:

# Jason D. Pizzini                              Jason.Pizzini@bcm.edu

# Email Addresses for All Authors:

# Firas S. Midani                              Firas.Midani@bcm.edu

# Robert A. Britton                          Robert.Britton@bcm.edu

# Jason D. Pizzini                              Jason.Pizzini@bcm.edu

## Author Questionnaire

- Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  No

3. Filming location: Will the filming need to take place in multiple locations?   No

Current Protocol Length

Number of Steps:  25

Number of Shots:  55

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

REQUIRED: What is the scope of your research? What questions are you trying to answer?

  - Jason Pizzini: Our research examines how the gut microbiome can be engineered to protect against harmful pathogens while revealing the ecological rules and mechanisms behind colonization resistance.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 2.3.1

What technologies are currently used to advance research in your field?

  - Jason Pizzini: Next-generation sequencing, advanced bioinformatics, germ-free mice, and in vitro gut models are all technologies that are transforming how we study microbial communities and their role in health and disease.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 2.5.1

What are the current experimental challenges?

  - Jason Pizzini: Many in vitro gut models are challenging to operate at high throughput. Models that allow the functional interrogation of microbial communities at scale are needed.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 3.2.1

What advantage does your protocol offer compared to other techniques?

  - Jason Pizzini: The Minibioreactor array is a continuous flow culture system that aims to overcome the lack of throughput capacity in other systems, letting us scale up experiments while still capturing complex and reproducible microbial community behavior.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B-roll: 4.1.1

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

- Preparation for Minibioreactor Array (MBRA) Assembly

Demonstrator: Jason Pizzini

  - To begin, ensure that the Minibioreactor array strips are 3D printed and contain 6 independent bioreactor chambers [1-TXT]. Arrange all the components required for the assembly [2]. Using a one-quarter inch-28NF fraction tap with a T-handle tap wrench, thread the three one-quarter inch ports in each chamber to insert fittings [3].

    - WIDE: Talent showing a completed 3D-printed Minibioreactor array strip with 6 chambers. TXT: MBRA: Minibioreactor Array

    - LAB MEDIA: Table 1 Video editor: Please show the 3 pages of the table sequentially so that all the components are shown

    - Talent threading the ports using a one-quarter inch-28NF fraction tap with a T-handle tap wrench. Videographer: Please shoot the assembly steps in close-up so that the pieces being assembled are clear visible

  - After washing the chamber with water, place a 10 by 3-millimeter magnetic stir bar into each chamber and add 1 milliliter of distilled water [1]. Then, position a rubber washer on top of each port of the bioreactor [2].

    - Talent placing a magnetic stir bar and pouring water into a chamber.

    - Talent placing a rubber washer onto a port.

  - For each chamber, screw in one waste straw-threaded male luer, one media straw-threaded male luer, and one empty threaded male luer into the ports [1].

    - Talent screwing the specified luers into the ports.

  - Now, insert 6 rubber septa onto three thirty-second-inch female luer barbs [1] and fold the upper sleeve of each septum down to cover the neck [2]. Attach these to the designated ports of each chamber [3]. Attach a one-eighth inch female luer barb to one end and a male luer lock connector to the opposite end of each length [4].

    - Talent placing rubber septa onto female luer barbs.

    - Talent folding the septum sleeve over the neck.

    - Talent attaching the septa to the designated ports.

    - Talent attaching a female luer barb and a male luer lock connector to each tubing length. TXT: Cut C-flex tubing strips of desired lengths

  - Then, insert a one-sixteenth-inch female luer barb into each end of the red two-stop E-lab tubing with a 1.14-millimeter inner diameter [1-TXT] and the orange two-stop E-lab tubing with a 0.89-millimeter inner diameter [2-TXT].

    - Talent inserting one-sixteenth inch female luer barb into  end of the red two-stop E-lab tubing. TXT: Briefly submerge the E-lab tubing’s ends in near-boiling water to soften them

    - Talent inserting into orange two-stop E-lab tubing with 0.89 millimeter inner diameter.  TXT: Repeat this process 6x for each MBRA strip

  - Connect the prepared E-lab tubing to the C-flex tubing, and ensure each of the 6 C-flex tubing lengths is connected to one red and one orange E-lab line via female luers [1].

    - Talent connecting orange and red E-lab tubing the C-flex tubing.

  - Attach one-eighth inch female luer barb and a male luer lock connector to both ends of one 3-inch piece and the 12-inch piece of C-flex tubing [1-TXT]. Attach male luer lock connectors to both ends of the remaining pieces [2].

    - Talent attaching a female luer barb and a male luer lock connector to the 3 inch and 12 inch tubing. TXT: Cut C-flex tubing; Twenty-one 1” pieces; One 2” piece; Three 3” pieces; One 12” piece Authors: This text is added now

    - Talent attaching male luer lock connectors to both ends of the remaining tubing.

  - Assemble the waste line tree according to the 3D diagram [1].

    - LAB MEDIA: Figure 3B.

  - Attach the exposed ends of the red two-stop E-lab tubing to the terminal male luer locks on the waste line tree in ascending order based on C-flex tubing length [1]. Then, connect the 3-inch C-flex tubing with the one-eighth-inch female luer barb and male luer lock connector to the top of the waste line tree [2].

    - Talent attaching red E-lab tubing to the waste line tree.

    - Talent securing the 3 inch C-flex tubing to the top of the assembly.

  - Assemble the feed line tree according to the 3D diagram [1].

    - LAB MEDIA: Figure 3B.

  - Bridge the exposed ends of the orange two-stop E-lab tubing to the terminal male luer locks on the feed line tree in ascending order based on C-flex tubing length [1] and attach the 12-inch C-flex tubing to the top of the feed line tree [2].

    - Talent attaching orange E-lab tubing to the feed line tree.

    - Talent securing the 12 inch C-flex tubing to the top of the assembly.

- MBRA Assembly Using the Prepared Components

  - Attach the variable-length C-flex tubing at the end of the feed line tree to the bioreactor [1], arranging them in ascending order from the shortest line on the left to the longest on the right [2].

    - Talent attaching the C-flex feed lines to the bioreactor ports.

    - Close-up of the feed lines arranged from shortest on the left to longest on the right.

  - Attach the variable-length C-flex tubing at the end of the waste line tree to the bioreactor strip in descending order [1], with the longest line on the left and the shortest on the right to accommodate pump placement [2]. Bundle all C-flex feed lines together on the left side of the strip and secure them with a twist tie [3-TXT].

    - Talent connecting waste lines to the bioreactor.

    - Close-up showing waste lines arranged from longest on the left to shortest on the right.

    - Talent bundling and tying the feed lines. TXT: Repeat the process for the waste lines on the right side

  - Form a loop with the orange two-stop E-lab tubing between the C-flex lines and secure the loop using autoclave tape [1] and repeat the process for the red two-stop E-lab tubing on the waste side of the bioreactor strip [2].

    - Talent looping and securing the orange tubing with autoclave tape.

    - Talent looping and securing the red tubing with autoclave tape.

  - Cover the female luer at the end of the waste and feed line trees with foil to prevent contamination [1]. Loosen the male threaded luers with septa on each bioreactor chamber to allow steam to escape during autoclaving [2].

    - Talent wrapping foil over the female luers.

    - Talent loosening threaded luers on each chamber.

  - After placing the assembly into an autoclave bin, stretch out the feed and waste line trees into separate bins adjacent to the one containing the MBRA strips [1-TXT].

    - Talent stretching feed and waste line trees into separate bins. TXT: Autoclave: 121 °C, ≥ 15 psi for 25 min; Once cooled, retighten the threaded male luers with the septa

- MBRA Connection and Operation

  - To attach the system to the pumps, remove the autoclave tape securing the E-lab tubing for both waste and feed lines [1] and untie the bundles of C-flex tubing [2]. Position the MBRA between the two pumps on top of the stir plate [3]. Clamp it down using the 3D-printed holders and align it with the marked stirring positions on the plate [4].

    - Talent peeling off autoclave tape from the E-lab tubing.

    - Talent untying bundled C-flex tubing.

    - Talent placing the MBRA between two pumps.

    - Talent clamping the MBRA into position on the stir plate.

  - Now, attach the feed line E-lab tubing to the peristaltic pump cartridges [1] and position the tubing stops into the cartridge slots [2]. Repeat the process for the waste line E-lab tubing on the pump located to the right of the stir plate [3]. Then, lock the peristaltic pump cartridges into the pump [4-TXT].

    - Talent attaching feed line E-lab tubing to pump cartridges.

    - Close-up of tubing stops seated in the cartridge slots.

    - Talent attaching waste line E-lab tubing to the right-hand pump.

    - Talent locking cartridges into the pump. TXT: Ensure the cartridges are seated fully

  - Arrange the C-flex tubing neatly using the 3D-printed tube holders [1]. Connect the end of the waste line tree to the tubing connected to the waste bottles [2]. Next, attach the female luer on the feed line entry tubing to the male connector on the 12 inch tube from the media bottle cap [3].

    - Talent placing C-flex tubing into the 3D-printed tube holders.

    - Talent connecting waste line tree tubing to waste bottle tubing.

    - Talent connecting feed line entry tubing to media bottle cap tubing.

  - Turn on both pumps to start media flow [1] and ensure both pumps are set to clockwise rotation when waste is positioned to the right of the pumps [2].

    - Talent powering on both pumps.

    - Close-up of pump settings showing clockwise rotation.

  - Observe the droplet size and cadence in each bioreactor chamber [1]. If variability is seen, replace the orange two-stop E-lab tubing connected to the affected chamber to reduce flow rate variation [2]. Once the chambers are full, shut off both pumps and allow the bioreactors to sit for 24 to 48 hours to check for contamination before starting the experiment [3].

    - Close-up of media droplets falling into a chamber.

    - Talent replacing an orange E-lab tubing line.

    - Talent switching off the pumps.

- MBRA Disassembly and Refurbishment

  - Switch the media input to a 1 liter container of 10 percent bleach in deionized water [1] and increase the flow rate on both pumps to maximum to displace the contents of the bioreactor chambers with bleach [2].

    - Talent replacing the media input line with a container of bleach solution.

    - Shot of Pump control panel showing the flow rate being increased to maximum.

  - Once the chambers are clear of media, invert the MBRA to disinfect above the fill line for 5 minutes [1]. After 5 minutes, right the system and wait an additional 5 minutes for sterilization [2].

    - Talent inverting the MBRA assembly.

    - Talent placing the MBRA upright on the bench.

  - After the chambers have been cleared of media and have been sterilized for 10 minutes, flush the system until 1 liter of water has passed through [1]. Then, disconnect the bioreactor E-lab tubing from the pumps and remove the MBRAs [2].

    - Show the clean system and then show the system being flused with water.

    - Talent detaching the E-lab tubing from the pumps and removing the MBRA assembly.

  - Remove the used septa from the bioreactor [1] and drain each chamber until only 1 milliliter of water remains [2]. Replace the septa and the orange 2-stop E-lab tubing before autoclaving the assembled strip [3]. Finally, refurbish and repair the MBRA as required [4].

    - Talent removing septa from the chamber ports.

    - Talent tilting the MBRA to drain water, leaving a small volume behind.

    - Talent replacing the septa and the orange 2-stop E-lab tubing. Authors: This shot is added

    - TEXT ON PLAIN BACKGROUND:

- Autoclave up to 3 reuse cycles

- After 3rd reuse: Fully disassemble MBRA

Replace C-flex tubing and sterilize parts

Replace if damaged

- Reapply epoxy to waste and feed PTFE connections if brittle

# Results

- Results

  - A human fecal sample was prepared and grown in the MBRA system. After four days of continuous flow [1], the microbial community in all nine bioreactors was dominated by 18 bacterial genera, each comprising at least 2% of relative abundance in any replicate [2].

    - LAB MEDIA: Figure 5A.

    - LAB MEDIA: Figure 5A. Video editor: Sequentially Highlight the colored stacked bars in the chart.

  - Twenty-two out of the 65 detected genera were present in all nine bioreactor replicates, demonstrating high reproducibility [1].

    - LAB MEDIA: Figure 5A.

  - Alpha diversity analysis showed minimal variation between replicates [1] in both the observed operational taxonomic units [2] and Shannon diversity index [3].

    - LAB MEDIA: Figure 5B.

    - LAB MEDIA: Figure 5B. Video editor: Highlight the top panel showing “Observed OTUs” .

    - LAB MEDIA: Figure 5B. Video editor: Highlight the bottom panel showing “Shannon Diversity” .