# Submission ID #: 68911

# Scriptwriter Name: Pallavi Sharma

# Project Page Link: https://review.jove.com/account/file-uploader?src=21017903

# Title: High-Definition Transcranial Direct Current Stimulation During Sleep

# Authors and Affiliations:

# Seo Ho Song1,2, Shengzi Zeng1,2, Tony Cunningham1,2

# 1Center for Sleep and Cognition, Department of Psychiatry, Beth Israel Deaconess Medical Center

# 2Division of Sleep Medicine, Harvard Medical School

# Corresponding Authors:

# Tony Cunningham		(acunnin4@g.harvard.edu)

# Email Addresses for All Authors:

# Seo Ho Song			(ssong4@bidmc.harvard.edu)

# Shengzi Zeng			(szeng@bidmc.harvard.edu)

# Tony Cunningham		(acunnin4@g.harvard.edu)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  No

Videographer: Please film the computer screen for all shots labeled as SCREEN

3. Filming location: Will the filming need to take place in multiple locations?   No

No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Yes

.

Current Protocol Length

Number of Steps: 31

Number of Shots: 55

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

INTRODUCTION:

  - Seo Ho Song: We use neuromodulation during sleep to investigate the relationships between targeted brain regions and sleep physiology.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Seo Ho Song: High-definition transcranial direct current stimulation with simultaneous electroencephalography advances research on the regulation of and restoration from sleep.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

  - Seo Ho Song: HD-tDCS provides enhanced spatial precision and focal targeting versus conventional tDCS, minimizing off-target effects during sleep.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Seo Ho Song: This enables mechanistic studies of region-specific sleep physiology and development of targeted interventions for neuropsychiatric disorders.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Seo Ho Song: Future research will develop adaptive closed-loop systems and refine computational models for improved targeting.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

# Videographer: Please capture all testimonial shots in a wide-angle format with sufficient headspace, as the final videos will be rendered in a 1:1 aspect ratio. Testimonial statements will be presented live by the authors, sharing their spontaneous perspectives.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Tony Cunningham, Director of Center for Sleep and Cognition, BIDMC: (authors will present their testimonial statements live)

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Tony Cunningham, Director of Center for Sleep and Cognition, BIDMC: (authors will present their testimonial statements live)

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera.

Ethics Title Card

This research has been approved by the Institutional Review Board (IRB) at the Beth Israel Deaconess Medical Center

# Protocol

- EEG Cap and HD-tDCS Electrode Preparation

Demonstrators: Tony Cunningham and Shengzi Zeng

  - To begin, gather all necessary equipment and materials required for the procedure [1]. Ensure the conventional tDCS device has an adequate power supply and if required, change the batteries [2 and 3]. Turn on HD-tDCS (H-D-T-D-C-S) and EEG (E-E-G) devices and verify the power status according to the manufacturer’s specifications [4].

    - WIDE: Talent assembling all required equipment and materials on the workbench.

Added shots: Talent removes any used batteries at the back of the stimulator.

Talent inserts charged batteries at the back of the stimulator.

    - Talent powering on both devices and checking the power status.

  - Connect the cables of the silver or silver chloride sintered ring electrodes to the corresponding receivers on the 4 by 1 adapter output cable [1].

    - Talent connecting each electrode cable to its matching receiver port on the 4 by 1 adapter output cable. NOTE: This is moved before 5.1.1 and was slated 5.1.3 according to the author

  - Select an appropriately sized EEG cap based on the participant’s head circumference [1]. Use bilateral frontolateral positions F3 and F4 to target the prefrontal cortex [2].

    - Talent measuring head circumference using a tape and picking a correctly sized cap.

    - Talent marking the F3 and F4 positions on the EEG cap.

  - Next, input the electrode positions and stimulation parameters [1].

    - Show the software interface with fields for electrode positions and stimulation settings.

  - Insert the HD-tDCS electrodes or the electrode holders into the designated positions on the EEG cap [1] and reverse the EEG cap inside-out and place it on a foam head for better access [2].

    - Talent inserting each HD-tDCS electrode or holder into the pre-marked slots on the cap.

2.6.2, Talent flipping the cap inside-out and placing it securely over a foam head model
NOTE: Show 2.6.2 after 2.5.1, VO is moved here too

  - Then, apply the high-viscosity, chloride-free adhesive conductive paste to the cap [1]. Using a tongue depressor, cover each EEG electrode with paste, avoiding application on HD-tDCS stimulation electrodes [2].

    - Talent opening and preparing the adhesive conductive paste for use. NOTE: 2.6.1, 2.6.3 and 2.7.1 were shot together

    - Talent flipping the cap inside-out and placing it securely over a foam head model. NOTE: Show 2.6.2 after 2.5.1, VO is moved

    - Talent using a tongue depressor to apply paste on EEG electrodes, ensuring no paste is applied on the stimulation electrodes.

  - Shape the conductive paste into dome-like structures [1]. Apply a smaller amount of paste in the front of the scalp where there is less hair and a greater amount at the crown and back of the head where hair density is higher [2].

    - Talent shaping the applied paste into small domes on the electrodes.

    - Talent adjusting the paste quantity, applying less in the front and more at the crown and back of the head.

- Participant Preparation for HD-tDCS

Demonstrator: Seo Ho Song

  - Apply the facial and ground reference electrodes using gold cup electrodes filled with adhesive conductive paste [1]. Secure each electrode in place using medical-grade tape [2].

    - Talent placing gold cup electrodes with paste onto designated facial and ground sites.

    - Talent fixing electrodes in place with medical-grade tape to ensure secure contact.

NOTE: Show step 3.3 here, after 3.1.2

  - Now, clean each electrode site using an alcohol swab [1]. Exfoliate the site with mildly abrasive preparation gel using a cotton swab [2], then remove any residual fluid with a dry gauze square [3]. Target sites include the forehead and mastoids behind the ears for reference electrodes [4]. Outer canthi near the eyes to record eye movements, chin for electromyography electrodes, and clavicle for the ground electrode [5].

    - Talent cleaning electrode sites using an alcohol swab.

    - Talent exfoliating each site with preparation gel using a cotton swab.

    - Talent wiping the sites dry using a gauze square.

    - WIDE: Talent pointing to the forehead and mastoid where electrodes will be placed. NOTE: 3.2.4 and 3.2.5 were shot together

    - Talent pointing to outer canthi, chin, and clavicle.

NOTE: Move this 3.3 step after 3.1

  - Then, fill each electrode with adhesive conductive paste and apply it to the prepared site, ensuring that no paste is exposed on the back side of the electrode [1].

    - Talent filling a gold cup electrode with paste and carefully placing it on the cleaned site without exposing paste on the back.

  - For the HD-tDCS return electrode, position the stimulation electrode on the collar bone and secure it with tape [1-TXT]. Apply the saline solution as per manufacturer’s instructions [2].

    - Talent placing the stimulation electrode on the clavicle and taping it in place. TXT: Apply saline solution or gel as required

    - Talent applying saline solution or gel to the electrode as specified. NOTE: Not filmed, VO moved as on-screen text

  - Now, measure the distance from the nasion to the inion and mark the halfway point along this line [1].

    - Talent measuring from the nasion to the inion with a measuring tape and marking the halfway point.

  - Then, measure the distance between the left and right pre-auricular points and mark the halfway point on this line with a visible line [1].

    - Talent measuring the horizontal distance between the left and right pre-auricular points and marking the center point.

  - Identify the Cz location at the intersection of the nasion-inion and pre-auricular lines [1]. After that, lightly spray the participant’s hair with water using a spray bottle to assist in EEG cap placement [2].

    - Talent locating the intersection point of the two marked lines and indicating Cz.

    - Talent spraying water onto the participant's hair using a spray bottle.

  - Invert the prepared EEG cap and place it snugly on the participant’s head, aligning the Cz electrode with the marked Cz site [1]. Secure the cap in place using the chin straps for a comfortable and stable fit [2].

    - Talent flipping the cap right-side out and carefully placing it over the participant’s head, aligning it with the Cz marker.

    - Talent fastening the chin straps beneath the participant’s jaw to secure the cap.

  - To fill each electrode site with conductive gel, pull up each electrode slightly and insert the large-gauge, blunt-tip needle into its center [1]. Gently wiggle the needle to displace hair and dispense approximately 0.5 milliliters of conductive gel [2]. Push the electrode back down to establish firm contact between the gel, paste, and scalp [3-TXT].

    - Talent lifting the electrode gently and inserting the needle into its center. NOTE: 3.9.1, 2.9.2 and 3.9.3 were shot together

    - Talent wiggling the needle to clear hair and applying gel into the electrode.

    - Talent pressing the electrode down to ensure proper connection. TXT: Press each electrode firmly against the scalp to enhance contact

  - Position the silver or silver chloride sintered ring electrodes into each HD-tDCS electrode holder with the rough surface facing downward and the smooth, rounded surface facing upward [1]. Add more high-viscosity conductive gel over each inserted electrode to fully cover its surface [2]. Lock each electrode in place using the caps provided [1].

    - Talent placing the sintered ring electrodes into the holders on the EEG cap.

    - Talent applying a layer of conductive gel on top of each electrode in the holder.

    - Talent attaching and securing the cap over each electrode to hold it firmly in place.

- HD-EEG System Setup

  - Turn on the amplifier after connecting the required cords [1] and launch the EEG recording software on the connected computer [2]. NOTE: VO modified for the extra shot

    - Talent connects the cords between the cap (with the attached electrodes) to the headbox and connects the headbox to the amplifier. NOTE: The action was  modified

Added shot: Talent opens the EEG software.

  - Check the electrode impedances using the software. Set the impedance limit to 25 kiloohms and aim to achieve values below 10 kiloohms for all electrodes [1].

    - SCREEN: Show the impedance check window. The impedance limit is being set to 25 kiloohms.

Videographer: Please film the computer screen for all shots labeled as SCREEN

  - Ask the participant to open and close their eyes for approximately 10 seconds [1]. Observe the presence of alpha wave activity, typically more prominent in posterior channels than in anterior ones. Monitor the scalp for amplitude variations [2].

    - Talent instructing the participant to open and close their eyes.

    - SCREEN: Display of EEG signal showing clear alpha waves, with emphasis on posterior channels. EEG monitor showing amplitude differences across various scalp channels.

  - Ask the participant to look to the left, right, up, and down. Then, instruct them to clench their teeth [1] while observing the recordings from the eye and chin electrodes for movement signals [2].

    - Shot of the patient performing eye movements and clenching the teeth

    - Display of EEG traces showing muscle and eye movement artifacts at corresponding electrode channels.

  - Ask the participant to blink several times [1].

    - Participant blinking the eyes.

- HD-tDCS Stimulator Set-up

  - Connect the cables of the silver or silver chloride sintered ring electrodes to the corresponding receivers on the 4 by 1 adapter output cable [1]. Switch on the stimulator using its power button [2-TXT]. Once the stimulator is on, plug the electrode cable securely into the back of the device [3].

2.2.1,. Talent connecting each electrode cable to its matching receiver port on the 4 by 1 adapter output cable. Author’s NOTE: Show 2.2.1 here. In the footage, it was written as 5.1.3 on the clapperboard

    - Talent pressing the power switch on the stimulator. TXT: Ensure the stimulator is plugged in or has charged batteries

    - Talent inserting the electrode cable into the rear port of the powered-on stimulator.

  - Set the stimulation parameters according to the study design. Configure the intensity, select the waveform as direct current for transcranial direct current stimulation protocols, and set the duration and ramp times [1].

    - SCREEN: Show the stimulation settings interface, highlighting input fields for intensity, waveform (set to DC), duration, and ramp times.

  - Select either active stimulation or sham mode as directed by the study protocol [1]. If the HD-tDCS electrodes show high impedance, deliver a brief test stimulation under the SHAM setting to verify connectivity [2].

    - SCREEN: Show selection of Active or Sham mode from the stimulator settings menu.

    - SCREEN: Talent selecting SHAM mode and initiating a short test stimulation while monitoring impedance readings.

  - Start the recording [1] and mark the time in the EEG recording when the lights are turned off [2]. Monitor the EEG for signs of sleep onset and specific sleep stages to trigger stimulation according to the study protocol [3].

Added shot: Talent starts recording.   Author’s NOTE: This my have been slated as 5.4.1 and then the next 2 shots as 5.4.2 and 5.4.3

    - SCREEN: Show EEG recording software with a timestamped notation added indicating "Lights off."

    - SCREEN: Display of real-time EEG data showing talent observing the transition into defined sleep stages.

  - Mark the start of stimulation in the EEG recording with a clear notation [1].

    - SCREEN: EEG software showing the user entering a timestamped note labeled “Stimulation start.”

  - Once the final stimulation period ends, unplug the stimulation cord from the device [1]. Mark the time in the EEG recording when the lights are turned on [2].

    - Talent disconnecting the stimulation cord from the stimulator.

    - SCREEN: EEG software showing a timestamped note labeled “Lights on.”

  - After waking the participant from sleep, stop the EEG recording using the software interface [1].

    - SCREEN: Show the software screen where the Stop Recording button is clicked, and the recording session is saved.

  - Gently remove the silver or silver chloride ring electrodes from the participant’s scalp [1]. Clean the used electrodes with a damp paper towel and allow them to air dry before storing them [2].

    - Talent unplugging and lifting each electrode gently from the electrode holders.

    - Talent wiping each electrode clean with a moistened paper towel and setting them aside to dry.

  - Once the EEG cap is removed from the participant’s head, clean all electrode sites thoroughly to remove residual conductive gels [1].

    - Talent using cleaning wipes or a soft cloth to remove remaining gel from all electrode contact points on the scalp.

# Results

- Results

  - Bilateral frontolateral HD-tDCS produced a focal electric field distribution, highlighting the spatial precision achievable with a 4 by 1 ring electrode montage [1].

    - LAB MEDIA: Figure 4. Video editor: Show the color map across the three brain views, focusing on the regions with the highest intensity (red and yellow areas).

  - Clean EEG signals with minimal artifacts were observed during the ramp-up phase, confirming correct protocol implementation with optimal electrode preparation [1].

    - LAB MEDIA: Figure 5.

  tDCS
Acronym for transcranial direct-current stimulation.
Pronunciation link: No confirmed authoritative link found (acronyms typically follow letter-by-letter pronunciation).
IPA (AmE): /tiː-diː-siː-ɛs/
Phonetic: tee-dee-cee-ess

  HD-tDCS
Acronym for high-definition transcranial direct-current stimulation.
Pronunciation link: No confirmed authoritative link found (letter-by-letter).
IPA (AmE): /eɪtʃ-diː-tiː-diː-siː-ɛs/
Phonetic: eich-dee-tee-dee-cee-ess

  EEG
Acronym for electroencephalography.
Pronunciation link: No confirmed authoritative link found (letter-by-letter).
IPA (AmE): /iː-iː-dʒiː/
Phonetic: ee-ee-jee

  sintered
Pronunciation link: https://www.merriam-webster.com/dictionary/sintered (Merriam-Webster) merriam-webster.com
IPA (AmE): /ˈsɪn-tərd/
Phonetic: sin-turd

  frontolateral
Frontolateral is a compound describing the front and side (lateral).
Pronunciation link: No confirmed dictionary entry found — based on standard components (fronto- + lateral).
IPA (AmE, inferred): /ˌfrʌn.toʊˈlæt.ə.rəl/
Phonetic: fruhn-toh-LAT-uh-ruhl

  prefrontal
Pronunciation link: https://www.merriam-webster.com/dictionary/prefrontal merriam-webster.com
IPA (AmE): /priːˈfrʌn-təl/
Phonetic: pree-FRUN-tuhl

  nasion
Anatomical landmark between the eyes at the top of the nose.
Pronunciation link: No confirmed Merriam/Webster link found; standard medical pronunciation used.
IPA (AmE): /ˈneɪ-si-ən/
Phonetic: NAY-see-uhn

  inion
Anatomical landmark on the back of the skull.
Pronunciation link: No confirmed dictionary entry found.
IPA (AmE): /ˈɪn-i-ən/
Phonetic: IN-ee-uhn

  pre-auricular
Describes the area in front of the ear.
Pronunciation link: No direct Merriam-Webster entry found; based on medical usage.
IPA (AmE): /pri-ɔːrˈɪk.jə.lər/
Phonetic: pree-aw-RICK-yuh-luhr

  Cz
Label in EEG systems denoting electrode position (“C-Zee”).
Pronunciation link: No direct entry; spoken as two letters.
IPA (AmE): /siː-ziː/
Phonetic: see-zee

  impedance
Pronunciation link: https://www.merriam-webster.com/dictionary/impedance merriam-webster.com
IPA (AmE): /ɪmˈpiː.dəns/
Phonetic: im-PEE-duhns

  kiloohms
Unit of electrical resistance.
Pronunciation link: No specific Merriam-Webster entry for “kiloohms”; based on standard pronunciation.
IPA (AmE): /ˈkɪl.oʊ.oʊmz/
Phonetic: KIL-oh-ohmz