# Submission ID #: 69236

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21117028

# Title: A Methodological Protocol and Considerations for Transcranial Ultrasonic Stimulation in Exploratory Clinical Human Studies

# Authors and Affiliations:

# Ziping Huang1,2,*, Mengyue Chen3,*, Charalambos C. Charalambous1, Lei Zhu1, Ergi Spiro1,2, Shashank Shekhar1, Jody A. Feld1,4, Xiaoning Jiang1,3, Junjie Yao1,2, Wuwei Feng1,2

# 1Department of Neurology, Duke University School of Medicine

# 2Department of Biomedical Engineering, Duke University

# 3Department of Mechanical and Aerospace Engineering, North Carolina State University

# 4Department of Orthopedic Surgery, Duke University School of Medicine

# *These authors contributed equally

# Corresponding Authors:

Ziping Huang            ziping.huang@duke.edu

# WuweiFeng              wayne.feng@duke.edu

# Email Addresses for All Authors:

Ziping Huang       ziping.huang@duke.edu

Mengyue Chen      mchen23@ncsu.edu

Charalambos C. Charalambous    charalambos.charalambous@duke.edu

Lei Zhu                     l.zhu@duke.edu

Ergi Spiro                    ergi.spiro@duke.edu

Shashank Shekhar               shashank.shekhar@duke.edu

Jody A. Feld                     jody.feld@duke.edu

Xiaoning Jiang       xjiang5@ncsu.edu

Junjie Yao           junjie.yao@duke.edu

Wuwei Feng         wayne.feng@duke.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done.

Videographer : Please film the screen of the instrument for the following shots:
2.4.1, 2.5.1, 3.1.2, 3.2.1, 3.2.2, 3.3.2, 3.4.1, 3.5.1, 3.5.2

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Yes

Current Protocol Length

Number of Steps: 27

Number of Shots: 56

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

INTRODUCTION:

  - Ziping Huang: Our research encompasses the use of transcranial ultrasonic stimulation (TUS) in neurotypical and neurological adults.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Ziping Huang: TUS facilitates basic neuroscience inquiries and is increasingly investigated for neurological and psychiatric diseases like post-stroke complications, depression, and pain.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

  - Ziping Huang: We established TUS safety and effective intensity in stroke patients, showing motor-cortex TUS can enhance motor learning and cortical excitation.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Ziping Huang: There is no methodological article focusing on the use of TUS in neurologically diseased adults. This protocol addresses this gap.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

  - Ziping Huang: We hope this protocol can reduce the methodological variability and improve the reliability and reproducibility of future human ultrasound studies.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Testimonial Questions (OPTIONAL):

# Videographer: Please capture all testimonial shots in a wide-angle format with sufficient headspace, as the final videos will be rendered in a 1:1 aspect ratio. Testimonial statements will be presented live by the authors, sharing their spontaneous perspectives.

- Testimonial statements will not appear in the video but may be featured in the journal’s promotional materials.

- Provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author delivering the testimonial.

- Please answer the testimonial question live during the shoot, speaking naturally and in your own words in complete sentences.

How do you think publishing with JoVE will enhance the visibility and impact of your research?

  - Ziping Huang, Graduate research assistant: (authors will present their testimonial statements live)

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Can you share a specific success story or benefit you’ve experienced—or expect to experience—after using or publishing with JoVE? (This could include increased collaborations, citations, funding opportunities, streamlined lab procedures, reduced training time, cost savings in the lab, or improved lab productivity.)

  - Ziping Huang, Graduate research assistant: (authors will present their testimonial statements live)

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Ethics Title Card

This research has been approved by the Institutional Review Board (IRB) at Duke University

# Protocol

- Hydrophone-Based Characterization of Ultrasound Transducer Acoustic Output in a Water Tank Setup

Demonstrators: Yu Chu, Ziping Huang

  - To begin, prepare a water tank with degassed, deionized water to prevent potential damage to the hydrophone [1]. Fixate the ultrasound transducer onto the fixture on the water tank [2] and mount the hydrophone onto the 3-axis positioning system [3].

    - WIDE: Talent preparing the water tank with degassed, deionized water.

    - Talent securing the transducer onto the fixture on the water tank.

    - Talent mounting the hydrophone onto the 3-axis positioning system.

  - Connect the hydrophone to an oscilloscope [1] and then connect the needle tip along with the cable [2]. Switch on the hydrophone power only after it is fully immersed in water [3]. Drive the ultrasound transducer using a function generator [4-TXT].

    - Talent connecting the hydrophone cable to the oscilloscope.

Added shot: Talent connecting the hydrophone needle tip and the cable to the pre-amp

    - Talent immersing the hydrophone into the water and then switching on its power.

    - Talent driving the transducer using the function generator. TXT: Do this only after the transducer is fully immersed and hydrophone is on

  - Then position the hydrophone tip at the approximate focal region of the ultrasound transducer using the 3-axis positioning system [1]. Scan the hydrophone across the water tank at possible locations of primary pressure output [2].

    - Talent adjusting the hydrophone tip to the approximate focal region using the 3-axis positioning system.

    - Talent scanning the hydrophone across the tank.

  - Record the oscilloscope readings at each location to obtain the acoustic pressure distribution [1] and document the hydrophone location corresponding to the focal region [2].

    - Talent recording readings from the oscilloscope and noting the hydrophone location corresponding to the focal region.
Videographer: Please record the screen of the instrument for this step

    - Talent noting the hydrophone position corresponding to the focal region. NOTE: This was combined with Shot 2.4.1.

  - Assess the transmitting sensitivity of the ultrasound transducer at the focal point by applying different input voltages [1].

    - Talent applying varying input voltages and noting pressure outputs to calculate transmitting sensitivity.
Videographer: Please record the screen of the instrument for this step

- Configuration and Calibration of Transcranial Ultrasound Stimulation (TUS) System

Demonstrators: Yu Chu, Ziping Huang

  - Connect the transcranial ultrasound stimulation hardware and start oscilloscope recording [1]. Turn on the oscilloscope, set it to Run, and adjust both vertical and horizontal axes to the scale of interest [2].

    - Talent connecting the TUS hardware.

    - Talent showing the oscilloscope startup, setting to Run mode, and demonstrating adjusting axis scales.
Videographer: Please record the screen of the instrument for this step

  - Flip the MAIN POWER switch of the amplifier to the on position [1]. Dial the power output of the amplifier to zero [2]. Configure both function generators to the desired waveform, ensuring ramping is added to each pulse duration [3].

    - Talent flipping the amplifier’s MAIN POWER switch to on

    - Shot of the amplifier output being dialed to zero. 
Videographer: Please record the screen of the instrument for this step

    - Talent showing configuration of waveform with settings for ramping on.
Videographer: Please record the screen of the instrument for this step

  - Now press Output on the function generator [1]. Dial the amplifier output to 100 percent using the ADJUST wheel and monitor the Gain display [2]. Press the amplifier POWER button to start its output [3].

    - Talent pressing Output on the function generator.

    - Talent turning the ADJUST wheel and checking Gain on the amplifier display.
Videographer: Please record the screen of the instrument for this step

    - Talent pressing the amplifier POWER button.

  - Confirm the drive signal waveform on the oscilloscope and allow the drive system to run for 15 minutes to stabilize the output [1].

    - Talent checking the waveform displayed on the oscilloscope.
Videographer: Please record the screen of the instrument for this step

  - Check for ramping at the beginning and end of each pulse duration on the drive signal waveform [1]. Dial the amplifier output to zero [2].

    - Talent checking the oscilloscope waveform with clear ramping at start and end of pulse duration. 
Videographer: Please record the screen of the instrument for this step

    - Talent dialing amplifier output to zero. 
Videographer: Please record the screen of the instrument for this step

  - Fully immerse the transducer in water, positioning the exit plane approximately 1 inch below the surface facing upward [1]. Then dial the amplifier output back to 100 percent and observe whether the ramping is successful [2-TXT].

    - Talent immersing the transducer exit plane 1 inch below water surface, facing up.

    - Talent dialing amplifier output to 100 percent and observing the water surface.  TXT: Rippling indicates success; Bead ejection indicates failure

  - Now mount a neuronavigation tool tracker onto the transducer holder [1]. In the software, navigate to Window and Tool Calibrations, click New Calibration, enter a name and select the correct tool tracker [2].

    - Talent attaching neuronavigation tool tracker to transducer holder.

    - SCREEN: 69236_3.7.2.mp4		00:00-00:25

  - Hold the transducer with the tracker against the calibration block such that the reference indicator pin touches and is perpendicular to the center of the transducer’s exit plane [1]. Click Begin Calibration Countdown, hold steady for 5 seconds, and repeat if necessary until Success appears [2]. Then close both Edit Calibration and Tool Calibrations windows [3].

    - Talent pressing the transducer against the calibration block with correct alignment.

    - SCREEN: 69236_3.8.2-3.8.3.mp4 		00:00-00:12

    - SCREEN: 69236_3.8.2-3.8.3.mp4 		00:15-00:20

  - Press the amplifier POWER button to turn off its output when the participant is ready for transcranial ultrasound stimulation [1]. Disconnect the oscilloscope from the amplifier and connect the transducer to the amplifier output [2].

    - Talent pressing the amplifier POWER button.

    - Talent disconnecting oscilloscope from amplifier and connecting transducer to amplifier output.

- Neuronavigation-Guided Transcranial Ultrasound Stimulation (TUS) Administration and Data Export Procedure

Demonstrators: Ziping Huang, Lei Zhu

  - If using the simulation-based target, load the target into the neuronavigation software [1]. In the Sessions tab, open the New... (new) dropdown and select Online Session [2]. Under the Targets tab, click on the Trajectory X-mod (X-mod) target and press Add => (add) to add it for the session [3-TXT].

    - SCREEN: 69236_4.1.1-4.1.3.mp4 		00:00-00:10

    - SCREEN: 69236_4.1.1-4.1.3.mp4 		00:11-00:18

    - SCREEN: 69236_4.1.1-4.1.3.mp4 		00:19-00:33 
TXT: Follow registration and proceed to Perform

  - Bring up the Bullseye (Tool Centric) (Bulls-eye-Tool-centric) or Bullseye (Target Centric) (Bulls-eye-Target-centric) view in the Perform window [1]. Set the Driver to the correct TUS transducer and select the correct target in the Targets to sample list on the upper left [2].

    - SCREEN: 69236_4.2.1-4.2.2.mp4 		00:00-00:13

    - SCREEN: 69236_4.2.1-4.2.2.mp4 		00:16-00:30

  - Comb the participant's hair to expose the scalp at the TUS (T-U-S) transducer site and create an anterior-posterior parting line [1].  Apply degassed ultrasound gel along it [2-TXT].

    - Talent combing hair to expose scalp at the approximate transducer site and making an anterior posterior parting line,

    - Talent applying ultrasound gel along the parting line. TXT: Apply finger pressure from center out to burst air bubbles

  - Create 1 to 3 additional parting lines approximately 1.5 centimeters away on each side and repeat the same process [1]. Finally, return to the initial parting line and repeat the process four times [2].

    - Talent creating additional parting lines and repeating gel application and pressure steps.

    - Talent returning to initial parting line and repeating actions.

  - Now strap the TUS transducer onto the participant’s head and align it with the target using the Bullseye view [1]. Adjust the transducer to minimize distance, tilting, and rotational deviations to within operational limits [2-TXT].

    - Talent strapping the transducer to the participant’s head and checking alignment on Bullseye view.

    - Talent adjusting transducer for minimal deviations. TXT: Once aligned, tighten strap and check participant comfort

  - If safety monitoring is required, insert a thermocouple wire into the ultrasound gel, close to the scalp [1-TXT]. Tape the thermocouple wire on the head using paper tape to prevent movement [2].

    - Talent inserting thermocouple wire near the edge of the transducer into the ultrasound gel. TXT: Ensure wire is not centered to avoid interference with ultrasound transmission

    - Talent taping the thermocouple wire with paper tape to secure it.

  - Press the amplifier’s POWER button to start output and begin TUS [1]. Simultaneously, press Sample Now on the neuronavigation software to record initial transducer placement [2].

    - Talent pressing the amplifier POWER button.

    - SCREEN: 69236_4.7.2-4.8.1.mp4.		00:00-00:13

  - Continuously monitor the three alignment metrics in the Bullseye view [1-TXT].

    - SCREEN: 69236_4.7.2-4.8.1.mp4.  00:14-00:27 TXT: Adjust transducer position if metrics deviate from target

  - Monitor the scalp temperature using the thermocouple wire if installed [1-TXT].

    - Talent monitoring real-time thermocouple temperature. TXT: Intervene if temperature approaches 43.3 °C

    - Talent checking on participant comfort and preparing to intervene if necessary.

  - To finish TUS, press the amplifier POWER button to stop the output [1]. Remove the transducer and all head-mounted equipment from the participant [2].

    - Talent pressing the amplifier POWER button to stop stimulation.

    - Talent removing the transducer and any attached equipment from the participant’s head.

  - Inspect the scalp for any signs of redness and ask the participant about any discomfort or adverse effects [1]. If no issues are reported or after addressing concerns, gently wipe off the ultrasound gel from the scalp [2].

    - Talent visually inspecting the scalp and speaking with the participant.

    - Talent wiping off the ultrasound gel from the participant’s scalp.

  - Click Finish on the neuronavigation window and save the project by selecting File and pressing Save Project [1]. From the Sessions list, select the current session and click Review [2].

    - SCREEN: 69236_4.12.1.mp4	00:05-00:23

    - SCREEN: 69236_4.12.2-4.13.2.mp4		00:00-00:10

  - In the Sessions Review window, choose Brainsight text file (.txt) (Brain-sight-text-file) from the Export... (Export) dropdown menu. Ensure Orientation is selected and set Snap to: (Snap-to) to Nothing [1]. Configure other export settings as needed, then click Save to export the stimulation data [2].

    - SCREEN: 69236_4.12.2-4.13.2.mp4		00:12-00:32

    - SCREEN: 69236_4.12.2-4.13.2.mp4		00:33-00:42

# Results

- Results

  - A satisfactory simulation-based targeting result was achieved where the acoustic focus reached the intended cortical stimulation region [1]. A scenario in simulation-based targeting required additional iterations due to misalignment between the acoustic focus and the intended stimulation target [2]. Adjusting the Distance Tx (T-X) outplane to skin parameter to a more realistic value corrected the visualization scaling issue in the simulation interface [3].

    - LAB MEDIA: Figure 5A. Video editor: Highlight the center of the color-mapped acoustic field where the red focal point overlaps with the intended black-cross stimulation region.

    - LAB MEDIA: Figure 5B.

    - LAB MEDIA: Figure 5C. Video editor: Highlight the field labelled “Distance Tx outplane to skin (mm)”

  - A clean motor evoked potential was recorded, indicating good transcranial magnetic stimulation coil placement and orientation [1]. A waveform was observed that might not represent a motor evoked potential [2].

    - LAB MEDIA: Figure 6A. Video editor: Highlight the peak bounded by the “0.0373(2)” and the “-0.0307(5)” cursors

    - LAB MEDIA: Figure 6C. Video editor: Highlight the irregular signal fluctuations across the plot without a single sharp peak.

  - A successful MT1mV (M-T-One-Milli-Volt) determination was made at 52% maximum stimulator output, with five positive and five negative responses at that level [1]. A potentially inconclusive MT1mV determination was recorded at 87% maximum stimulator output due to consecutive rather than alternating positive and negative responses [2].

    - LAB MEDIA: Figure 6B. Video editor: Highlight the third column for the last ten 52% MSO rows where the values are five 1.0s and five 0.0s.

    - LAB MEDIA: Figure 6D. Video editor: Highlight the third column for the 87% MSO row where consecutive 1.0s and 0.0s appear instead of alternating responses.

  - A steady, even oscilloscope reading of the transducer driving signal was obtained, indicating proper equipment function [1].

    - LAB MEDIA: Figure 7C. Video editor: Highlight the consistent waveform on the oscilloscope screen.

  - Good alignment was achieved between the physical transducer placement and the pre-defined target using neuronavigation [1]. A misaligned transducer placement was observed, indicating the need for adjustment to match the target [2].

    - LAB MEDIA: Figure 8A. Video editor: Highlight the red and green overlay and the three near-zero alignment metrics

    - LAB MEDIA: Figure 8B. Video editor: Highlight the red transducer placement marker that is offset from the central green target area, and the three large alignment metrics.

- Degassed
Pronunciation link: https://www.merriam-webster.com/dictionary/degassed
IPA: /ˌdiːˈɡæst/
Phonetic Spelling: dee·gast

- Deionized
Pronunciation link: https://www.merriam-webster.com/dictionary/deionized
IPA: /diˈaɪəˌnaɪzd/
Phonetic Spelling: dee·eye·uh·nyzd

- Hydrophone
Pronunciation link: https://www.merriam-webster.com/dictionary/hydrophone
IPA: /ˈhaɪdrəˌfoʊn/
Phonetic Spelling: hy·druh·fohn

- Ultrasound
Pronunciation link: https://www.merriam-webster.com/dictionary/ultrasound
IPA: /ˈʌltrəˌsaʊnd/
Phonetic Spelling: uhl·truh·sownd

- Transducer
Pronunciation link: https://www.merriam-webster.com/dictionary/transducer
IPA: /trænzˈduːsər/
Phonetic Spelling: tranz·doo·ser

- Oscilloscope
Pronunciation link: https://www.merriam-webster.com/dictionary/oscilloscope
IPA: /əˈsɪləˌskoʊp/
Phonetic Spelling: uh·sil·uh·skohp

- Function Generator
Pronunciation link: https://www.merriam-webster.com/dictionary/function%20generator
IPA: /ˈfʌŋkʃən ˈdʒɛnəˌreɪtər/
Phonetic Spelling: funk·shuhn jen·uh·ray·ter

- Acoustic
Pronunciation link: https://www.merriam-webster.com/dictionary/acoustic
IPA: /əˈkuːstɪk/
Phonetic Spelling: uh·koos·tik

- Hemodynamics
Pronunciation link: https://www.merriam-webster.com/dictionary/hemodynamics
IPA: /ˌhiːmoʊdaɪˈnæmɪks/
Phonetic Spelling: hee·moh·dy·nam·iks

- Transcranial
Pronunciation link: https://www.merriam-webster.com/dictionary/transcranial
IPA: /trænzˈkreɪniəl/
Phonetic Spelling: tranz·kray·nee·uhl

- Neuronavigation
Pronunciation link: https://www.merriam-webster.com/dictionary/neuronavigation
IPA: /ˌnʊrəˌnævɪˈɡeɪʃən/
Phonetic Spelling: nyur·oh·nav·uh·gay·shuhn

- Thermocouple
Pronunciation link: https://www.merriam-webster.com/dictionary/thermocouple
IPA: /ˈθɝːmoʊˌkʌpəl/
Phonetic Spelling: thur·moh·kup·uhl

- Cortical
Pronunciation link: https://www.merriam-webster.com/dictionary/cortical
IPA: /ˈkɔːrtɪkəl/
Phonetic Spelling: kor·tih·kuhl

- Variability
Pronunciation link: https://www.merriam-webster.com/dictionary/variability
IPA: /ˌveriəˈbɪləti/
Phonetic Spelling: vair·ee·uh·bil·uh·tee

- Evoked
Pronunciation link: https://www.merriam-webster.com/dictionary/evoked
IPA: /ɪˈvoʊkt/
Phonetic Spelling: ih·vohkt

- Simulation
Pronunciation link: https://www.merriam-webster.com/dictionary/simulation
IPA: /ˌsɪmjəˈleɪʃən/
Phonetic Spelling: sim·yuh·lay·shuhn

- Alignment
Pronunciation link: https://www.merriam-webster.com/dictionary/alignment
IPA: /əˈlaɪnmənt/
Phonetic Spelling: uh·lyn·muhnt

- Calibration
Pronunciation link: https://www.merriam-webster.com/dictionary/calibration
IPA: /ˌkæləˈbreɪʃən/
Phonetic Spelling: kal·uh·bray·shuhn

- Spectral
Pronunciation link: https://www.merriam-webster.com/dictionary/spectral
IPA: /ˈspɛktrəl/
Phonetic Spelling: spek·truhl

- Amplifier
Pronunciation link: https://www.merriam-webster.com/dictionary/amplifier
IPA: /ˈæmpləˌfaɪər/
Phonetic Spelling: am·pluh·fy·er

Top of Form

Bottom of Form