# Submission ID #: 69236

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21117028

# Title: A Methodological Protocol and Considerations for Transcranial Ultrasonic Stimulation in Exploratory Clinical Human Studies

# Authors and Affiliations:

# Ziping Huang1,2,*, Mengyue Chen3,*, Charalambos C. Charalambous1, Lei Zhu1, Ergi Spiro1,2, Shashank Shekhar1, Jody A. Feld1,4, Xiaoning Jiang1,3, Junjie Yao1,2, Wuwei Feng1,2

# 1Department of Neurology, Duke University School of Medicine

# 2Department of Biomedical Engineering, Duke University

# 3Department of Mechanical and Aerospace Engineering, North Carolina State University

# 4Department of Orthopedic Surgery, Duke University School of Medicine

# *These authors contributed equally

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Ziping Huang                                                                      ziping.huang@duke.edu

# Wuwei Feng                                                                       wayne.feng@duke.edu

# Email Addresses for All Authors:

Ziping Huang                                                                      ziping.huang@duke.edu

Mengyue Chen                                                             mchen23@ncsu.edu

Charalambos C. Charalambous                                 charalambos.charalambous@duke.edu

Lei Zhu                                                                           l.zhu@duke.edu

Ergi Spiro                                                                       ergi.spiro@duke.edu

Shashank Shekhar                                                        shashank.shekhar@duke.edu

Jody A. Feld                                                                   jody.feld@duke.edu

Xiaoning Jiang                                                               xjiang5@ncsu.edu

Junjie Yao                                                                       junjie.yao@duke.edu

Wuwei Feng                                                                   wayne.feng@duke.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps: 27

Number of Shots: 57

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer up to 2 introduction and up to 3 conclusion questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Answer in full sentences, the questions will not be displayed in the video.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Enter author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What are the most recent developments in your field of research?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What technologies are currently used to advance research in your field?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What are the current experimental challenges?

  - Enter author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What significant findings have you established in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What research gap are you addressing with your protocol?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What advantage does your protocol offer compared to other techniques?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How will your findings advance research in your field?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What new scientific questions have your results paved the way for?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

What questions will future research focus on?

  - Enter author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

Ethics Title Card

This research has been approved by the Institutional Review Board (IRB) at Duke University

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Hydrophone-Based Characterization of Ultrasound Transducer Acoustic Output in a Water Tank Setup

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, fill a water tank with degassed, deionized water to prevent potential damage to the hydrophone [1]. Fixate the ultrasound transducer onto the fixture on the water tank [2] and mount the hydrophone onto the 3-axis positioning system [3].

    - WIDE: Talent filling the water tank with degassed, deionized water.

    - Talent securing the transducer onto the fixture on the water tank.

    - Talent mounting the hydrophone onto the 3-axis positioning system.

  - Connect the hydrophone to an oscilloscope [1] and switch on the hydrophone power only after it is fully immersed in water [2]. Drive the ultrasound transducer using a function generator and radio frequency power amplifier [3-TXT].

    - Talent connecting the hydrophone cable to the oscilloscope.

    - Talent immersing the hydrophone into the water and then switching on its power.

    - Talent driving the transducer using the function generator and radio frequency power amplifier. TXT: Do this only after the transducer is fully immersed and hydrophone is on

  - Then position the hydrophone tip at the approximate focal region of the ultrasound transducer using the 3-axis positioning system [1]. Scan the hydrophone across the water tank at possible locations of primary pressure output [2].

    - Talent adjusting the hydrophone tip to the approximate focal region using the 3-axis positioning system.

    - Talent scanning the hydrophone across the tank.

  - Record the oscilloscope readings at each location to obtain the acoustic pressure distribution [1] and document the hydrophone location corresponding to the focal region [2].

    - Talent recording readings from the oscilloscope.

    - Talent noting the hydrophone position corresponding to the focal region.

  - Assess the transmitting sensitivity of the ultrasound transducer at the focal point by applying different input voltages [1].

    - Talent applying varying input voltages and noting pressure outputs to calculate transmitting sensitivity.

- Configuration and Calibration of Transcranial Ultrasound Stimulation (TUS) System

Demonstrator: Click here to enter name of demonstrator(s)

  - Connect the transcranial ultrasound stimulation hardware and start oscilloscope recording [1]. Turn on the oscilloscope, set it to Run, and adjust both vertical and horizontal axes to the scale of interest [2].
Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=21117028

    - Talent connecting the TUS hardware.

    - SCREEN: Show oscilloscope startup, set to Run mode, and demonstrate adjusting axis scales.

  - Flip the MAIN POWER switch of the amplifier to the on position [1]. Turn the power output of the amplifier to zero [2]. Configure both function generators to the desired waveform, ensuring ramping is added to each pulse duration [3].

    - Talent flipping the amplifier’s MAIN POWER switch to on.

    - Talent dialing amplifier output to zero.

    - SCREEN: Show configuration of waveform with ramping on both function generators.

  - Now press Output on the top generator first, then on the bottom generator [1]. Dial the amplifier output to 100 percent using the ADJUST wheel and monitor the Gain display [2]. Press the amplifier POWER button to start its output [3].

    - Talent pressing Output on the top generator then pressing Output on the bottom generator

    - Talent turning the ADJUST wheel and checking Gain on the amplifier display.

    - Talent pressing the amplifier POWER button.

  - Confirm the drive signal waveform on the oscilloscope and allow the drive system to run for 15 minutes to stabilize the output [1].

    - SCREEN: Show the waveform displayed on the oscilloscope.

  - Check for ramping at the beginning and end of each pulse duration on the drive signal waveform [1]. Dial the amplifier output to zero [2].

    - SCREEN: Show the oscilloscope waveform with clear ramping at start and end of pulse duration.

    - Talent dialing amplifier output to zero.

  - Fully immerse the transducer in water, positioning the exit plane approximately 1 inch below the surface facing upward [1]. Then dial the amplifier output back to 100 percent and observe whether the ramping is successful [2-TXT].

    - Talent immersing the transducer exit plane 1 inch below water surface, facing up.

    - Talent dialing amplifier output to 100 percent and observing the water surface.  TXT: Rippling indicates success; Bead ejection indicates failure

  - Now mount a neuronavigation tool tracker onto the transducer holder [1]. In the software, navigate to Window and Tool Calibrations, click New Calibration, enter a name and select the correct tool tracker [2].

    - Talent attaching neuronavigation tool tracker to transducer holder.

    - SCREEN: Show Tool Calibrations window and enter calibration name and tracker selection.

  - Hold the transducer with the tracker against the calibration block such that the reference indicator pin touches and is perpendicular to the center of the transducer’s exit plane [1]. Click Begin Calibration Countdown, hold steady for 5 seconds, and repeat if necessary until Success appears [2]. Then close both Edit Calibration and Tool Calibrations windows [3].

    - Talent pressing the transducer against the calibration block with correct alignment.

    - SCREEN: Click Begin Calibration Countdown, show countdown and Success confirmation.

    - SCREEN: Closing Edit Calibration and Tool Calibrations windows.

  - Press the amplifier POWER button to turn off its output when the participant is ready for transcranial ultrasound stimulation [1]. Disconnect the oscilloscope from the amplifier  and connect the transducer to the amplifier output [2].

    - Talent pressing the amplifier POWER button.

    - Talent disconnecting oscilloscope from amplifier and connecting transducer to amplifier output.

- Neuronavigation-Guided Transcranial Ultrasound Stimulation (TUS) Administration and Data Export Procedure

Demonstrator: Click here to enter name of demonstrator(s)

  - If using the simulation-based target, load the target into the neuronavigation software [1]. In the Sessions tab, open the New... (new) dropdown and select Online Session [2]. Under the Targets tab, click on the Trajectory X-mod (X-mod) target and press Add => (add) to add it for the session [3-TXT].

    - SCREEN: Show neuronavigation software interface with simulation-based target being loaded.

    - SCREEN: Selecting Online Session from the New... dropdown menu.

    - SCREEN: Clicking Trajectory X-mod target and pressing Add =>. TXT: Follow registration and proceed to Perform

  - Bring up the Bullseye (Coil Centric) (Bulls-eye-Coil-centric) or Bullseye (Target Centric) (Bulls-eye-Target-centric) view in the Perform window [1]. Set the Driver to the correct TUS transducer and select the correct target in the Targets to sample list on the upper left [2].

    - SCREEN: Displaying Bullseye views in the Perform window.

    - SCREEN: Setting correct Driver and selecting target from Targets to sample list.

  - Comb the participant's hair to expose the scalp at the TUS (T-U-S) transducer site and create an anterior-posterior parting line [1].  Apply degassed ultrasound gel along it [2-TXT].

    - Talent combing hair to expose scalp at the approximate transducer site and making an anterior posterior parting line,

    - Talent applying ultrasound gel along the parting line. TXT: Apply finger pressure from center out to burst air bubbles

  - Create 1 to 3 additional parting lines approximately 1.5 centimeters away on each side, and repeat the same process [1]. Finally, return to the initial parting line and repeat the process four times [2].

    - Talent creating additional parting lines and repeating gel application and pressure steps.

    - Talent returning to initial parting line and repeating actions.

  - Now strap the TUS transducer onto the participant’s head and align it with the target using the Bullseye view [1]. Adjust the transducer to minimize distance, tilting, and rotational deviations to within operational limits [2-TXT].

    - Talent strapping the transducer to the participant’s head and checking alignment on Bullseye view.

    - Talent adjusting transducer for minimal deviations. TXT: Once aligned, tighten strap and check participant comfort

  - If safety monitoring is required, insert a thermocouple wire into the ultrasound gel, close to the scalp [1-TXT]. Tape the thermocouple wire on the head using paper tape to prevent movement [2].

    - Talent inserting thermocouple wire near the edge of the transducer into the ultrasound gel. TXT: Ensure wire is not centered to avoid interference with ultrasound transmission

    - Talent taping the thermocouple wire with paper tape to secure it.

  - Press the amplifier’s POWER button to start output and begin TUS [1]. Simultaneously, press Sample Now on the neuronavigation software to record initial transducer placement [2].

    - Talent pressing the amplifier POWER button.

    - SCREEN: Clicking Sample Now in neuronavigation software to record placement.

  - During TUS, press Sample Now in the software as needed to record transducer placement [1]. Continuously monitor the three alignment metrics in the Bullseye view [2-TXT] .

    - SCREEN: Clicking Sample Now at intervals during TUS.

    - SCREEN: The 3 alignment metrics are being seen.  TXT: Adjust transducer position if metrics deviate from target

  - Monitor the scalp temperature using the thermocouple wire if installed [1-TXT].

    - SCREEN: Showing real-time thermocouple temperature monitoring. TXT: Intervene if temperature approaches 43.3 °C

    - Talent checking on participant comfort and preparing to intervene if necessary.

  - To finish TUS, press the amplifier POWER button to stop the output [1]. Remove the transducer and all head-mounted equipment from the participant [2].

    - Talent pressing the amplifier POWER button to stop stimulation.

    - Talent removing the transducer and any attached equipment from the participant’s head.

  - Inspect the scalp for any signs of redness and ask the participant about any discomfort or adverse effects [1]. If no issues are reported or after addressing concerns, gently wipe off the ultrasound gel from the scalp [2].

    - Talent visually inspecting the scalp and speaking with the participant.

    - Talent wiping off the ultrasound gel from the participant’s scalp.

  - Click Finish on the neuronavigation window and save the project by selecting File  and pressing Save Project [1]. From the Sessions list, select the current session and click Review [2].

    - SCREEN: Navigating to File > Save Project.

    - SCREEN: Selecting current session and clicking Review.

  - In the export window, choose Brainsight text file (.txt) (Brain-sight-text-file) from the Export... (Export) dropdown menu. Ensure Orientation is selected and set Snap to: (Snap-to) to Nothing [1]. Configure other export settings as needed, then click Save to export the stimulation data [2].

    - SCREEN: Choosing Brainsight text file (.txt) and Orientation is being selected and Snap To: is being set to Nothing.

    - SCREEN: Other export settings are being configured then Save is clicked to complete data export.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 193.

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - A satisfactory simulation-based targeting result was achieved where the acoustic focus reached the intended cortical stimulation region [1]. A scenario in simulation-based targeting required additional iterations due to misalignment between the acoustic focus and the intended stimulation target [2]. Adjusting the Distance Tx (T-X) outplane to skin parameter to a more realistic value corrected the visualization scaling issue in the simulation interface [3].

    - LAB MEDIA: Figure 5A. Video editor: Highlight the center of the color-mapped acoustic field where the blue-green focal point overlaps with the intended stimulation region.

    - LAB MEDIA: Figure 5B.

    - LAB MEDIA: Figure 5C. Video editor: Highlight the field labelled “Distance Tx outplane to skin (mm)”

  - A clean motor evoked potential was recorded, indicating good transcranial magnetic stimulation coil placement and orientation [1]. A waveform was observed that might not represent a motor evoked potential [2].

    - LAB MEDIA: Figure 6A. Video editor: Highlight the peak labeled “0.079”

    - LAB MEDIA: Figure 6C. Video editor: Highlight the irregular signal fluctuations across the plot without a single sharp peak.

  - A successful MT1mV (M-T-One-Milli-Volt) determination was made at 52% maximum stimulator output, with five positive and five negative responses at that level [1]. A potentially inconclusive MT1mV determination was recorded at 87% maximum stimulator output due to consecutive rather than alternating positive and negative responses [2].

    - LAB MEDIA: Figure 6B. Video editor: Highlight the third column for the 52% MSO row where the five values are 1.0 and five are 0.0.

    - LAB MEDIA: Figure 6D. Video editor: Highlight the third column for the 87% MSO row where consecutive 1.0s and 0.0s appear instead of alternating responses.

  - A steady, even oscilloscope reading of the transducer driving signal was obtained, indicating proper equipment function [1].

    - LAB MEDIA: Figure 7C. Video editor: Highlight the consistent waveform on the oscilloscope screen.

  - Good alignment was achieved between the physical transducer placement and the pre-defined target using neuronavigation [1]. A misaligned transducer placement was observed, indicating the need for adjustment to match the target [2].

    - LAB MEDIA: Figure 8A. Video editor: Highlight the red and green overlay

    - LAB MEDIA: Figure 8B. Video editor: Highlight the green transducer placement marker that is offset from the central red target area.