# Submission ID #: 69288

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21132888

# Title: Clinical Imaging of Microwave Mammography

# Authors and Affiliations:

# Yoshihiko Kuwahara, Kimihito Fujii

# Department of Surgery, Aichi Medical University

# Corresponding Authors:

Yoshihiko Kuwahara                     kuwahara.yoshihiko.296@mail.aichi-med-u.ac.jp

# Email Addresses for All Authors:

Yoshihiko Kuwahara                     kuwahara.yoshihiko.296@mail.aichi-med-u.ac.jp

Kimihito Fujii                           kfujii@aich-med-u.ac.jp

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes, all done

3. Filming location: Will the filming need to take place in multiple locations?   Yes, 200 m apart

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

Current Protocol Length

Number of Steps: 18

Number of Shots: 35

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

AUTHORS: Please note that only 2 introduction and 3 conclusion statements may be presented

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Yoshihiko Kuwahara: We aim to develop a clinical imaging system for early breast cancer detection using microwave technology.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.1.1

What are the most recent developments in your field of research?

  - Yoshihiko Kuwahara: The use of AI is being actively considered as progress in image reconstruction algorithms and antennas has stalled.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What significant findings have you established in your field?

  - Yoshihiko Kuwahara: We demonstrated that microwave imaging can distinguish malignant from normal breast tissue in clinical tests.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:3.2

What research gap are you addressing with your protocol?

  - Yoshihiko Kuwahara: We address the lack of noninvasive, radiation-free breast cancer imaging techniques.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera. Suggested B.roll:2.8

What advantage does your protocol offer compared to other techniques?

  - Yoshihiko Kuwahara: Our method offers safe, low-cost imaging without radiation or contrast agents.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

Ethics Title Card

This research has been approved by the Institutional Review Board (IRB) at Aichi Medical University

# Protocol

- Microwave Imaging Setup and Image Reconstruction Using Analyzer Control and MATLAB/Excel Integration

Demonstrator: Yoshihiko Kuwahara and Kimihito Fujii

  - To begin, launch the measurement program Analyzer Control on the laptop [1].

    - WIDE: Talent powering on the laptop and launching the Analyzer Control program from the desktop.

  - Enter the start frequency as 1000 megahertz, the stop frequency as 6000 megahertz, the number of data points as 101, and the response waiting time as 200 milliseconds [1]. Use default values for all other parameters [2].

    - SCREEN: 69288_screenshot_1	00:11-00:30

    - SCREEN: 69288_screenshot_1	00:31-00:37

  - Press the DIO (D-I-O) button, then press the GPIB (G-P-I-B0 Confirmation button [1]. If all connections are correct, the CAL Wizard (cal-wizard) , CAL Load, Auto Measurement, and Manual Measurement buttons will become active [2].

    - SCREEN: 69288_screenshot_2	00:00-00:09

    - SCREEN: 69288_screenshot_2	00:10-00:15

  - Press CAL Load, select the appropriate file, and click Apply to load the calibration data into the vector network analyzer [1].

    - SCREEN: 69288_screenshot_3	00:03-00:25

  - Set the AUTO/MANUAL (Auto-or-Manual) switch on the control unit to AUTO [1].

    - Talent switching the AUTO/MANUAL toggle on the control unit from MANUAL back to AUTO.

  - Next, position the patient in a prone position on the bed [1]. Ask the patient to insert one breast into the imaging sensor, ensuring that no clothing obstructs the sensor [2].

    - Talent guiding the patient to lie face-down on the examination bed.

    - Shot of the patient placing one breast securely into the imaging sensor.

  - Turn on the aspirator to initiate suction [1].

    - Talent switching on the aspirator and observing the initiation of suction.

  - Adjust the vacuum pressure to 0.04 to 0.06 megapascal using the CTRL (control) knob [1]. If the pressure cannot be reduced, reposition the patient until proper suction is achieved [2]. Stop imaging if the aspirator pressure remains at 0 megapascal [3].

    - Talent rotating the CTRL knob and monitoring the pressure gauge until it reaches the desired range.

    - Talent adjusting the patient's body position and rechecking suction connection.

    - Talent observing a reading of 0 megapascal and discontinuing the imaging session.

  - Press the Auto Measurement button [1]. The measurement status window will appear [2]. Set the measurement range to 1 through 28 and enable Skip Reciprocity [3].

    - SCREEN: 69288_screenshot_4	00:00-00:07

    - SCREEN: 69288_screenshot_4	00:08
Video Editor: Please freeze frame here

    - SCREEN: 69288_screenshot_4	00:09-00:28

  - Now, press Start Measurement to begin data acquisition [1]. The transmission and reception combinations will be displayed during measurement [2]. After completion, results are saved in the VNA_Results (V-N-A-Results) folder in CSV (C-S-V) format [3].

    - SCREEN: 69288_screenshot_5	00:07-00:12

    - SCREEN: 69288_screenshot_5	00:13-00:23

    - SCREEN: 69288_screenshot_6	00:12-00:28

  - To create a patient spreadsheet for use in the image reconstruction program, open the data_read (data-Reed) program in MATLAB (mat-lab) and process the information in the measurement file [1].

    - SCREEN: 69288_screenshot_7	00:00-00:11

  - Set the file path and specify the measurement files, then assign an output data name [1]. Click the RUN button in the MATLAB editor to generate a MAT file for a spreadsheet [2].

    - SCREEN: 69288_screenshot_7	00:12-00:20

    - SCREEN: 69288_screenshot_7	00:25-00:33, 00:37-00:40

  - Next, open the initialize_for_clinic (initialize-for-clinic) program in MATLAB [1]. Enter the reconstruction frequency, path and names of the measurement files, breast designations, and patient worksheet path and name [2].

    - SCREEN: 69288_screenshot_8	00:00-00:08

    - SCREEN: 69288_screenshot_8	00:10-00:36

  - Click the RUN button to generate a patient worksheet for Excel VBA use [1]. A directory is automatically created to store the reconstruction results [2].

    - SCREEN: 69288_screenshot_9	00:03-00:11

    - SCREEN: 69288_screenshot_9	00:14-00:25

  - Open the image_reconstruction_# (image-reconstruction-number) MATLAB file [1]. Specify the patient worksheet and name, reconstruction frequency, and calibration file, then save the settings [2].

    - SCREEN: 69288_screenshot_10	00:00-00:07.

    - SCREEN: 69288_screenshot_10	00:10-00:36

  - Now , open the multi_person (multi-person) Excel image reconstruction program [1]. Enter the patient worksheet names created earlier [2].

    - SCREEN: 69288_screenshot_11	00:00-00:14

    - SCREEN: 69288_screenshot_11	00:15-00:36

  - Then open the VBA program from the Excel add-in [1]. Specify the worksheet processing range from 1 to 8 [2]. Confirm that worksheet directories are correctly set in the Data_read and Data_write standard modules [3].

    - SCREEN: 69288_screenshot_12	 00:00-00:10

    - SCREEN: 69288_screenshot_12	00:11-00:17

    - SCREEN: 69288_screenshot_12	00:18-00:33

  - Select getPDTdata (get-P-D-T-Data) and click the Play button to begin the image reconstruction process [1]. After processing, verify that the reconstructed 3D tomographic images are saved in the correct directory alongside the patient worksheets [2].

    - SCREEN: 69288_screenshot_13	00:04-00:26

    - SCREEN: 69288_screenshot_14	00:04-00:12,00:15-00:26

# Results

- Results

  - High-permittivity and high-conductivity regions were observed in the left breast containing cancer [1], whereas no such regions appeared in the right, cancer-free breast of Patient 1 [2].

    - LAB MEDIA: Figure 10	Video Editor:Please highlight images A, C and D

    - LAB MEDIA: Figure 10	Video Editor:Please highlight images B, E and F

  - In Patient 2, contrast-enhanced MRI clearly detected the cancer in the left breast [1], but it was difficult to identify on X-ray mammography [2]. High-permittivity and high-conductivity areas were also observed in the left breast using the prototype [3], but no significant regions were detected in the cancer-free right breast [4].

    - LAB MEDIA: Figure 11	Video Editor:Please highlight images C to E

    - LAB MEDIA: Figure 11	Video Editor:Please highlight images A And B

    - LAB MEDIA: Figure 11. Video Editor:Please highlight images F and G

    - LAB MEDIA: Figure 11. Video Editor:Please highlight images H and I

  - In Patient 3, the cancer was visible near the nipple on MRI images before chemotherapy [1], but was absent after treatment [2].  Images obtained with the prototype also confirmed the disappearance of the cancer after chemotherapy [3].

    - LAB MEDIA: Figure 12	Video Editor:Please highlight image A

    - LAB MEDIA: Figure 12	Video Editor:Please highlight image B

    - LAB MEDIA: Figure 12	Video Editor:Please highlight images C to F