# Submission ID #: 69305

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=21138073

# Title: Porcine Cardiac Arrest Model Using an Implantable Defibrillator

# Authors and Affiliations:

# Benjamin Stage Storm1–3, Knut Tore Lappegård2,3, Charlotte Björk Ingul1,4, Erik Waage Nielsen1,2,3,5, Bent Aksel Nilsen1,3, Anette Nyjordet1,6, Ole-Jakob How1,2, Bjørn Faldaas1,4

# 1Nord University

# 2UiT The Arctic University of Norway

# 3Nordland Hospital Trust

# 4NTNU The Norwegian University of Science and Technology

# 5Department of Pain Medicine and Research, Oslo University Hospital and University of Oslo

# 6Helgeland Hospital Trust

# Corresponding Authors:

# Benjamin Stage Storm		benjamin.s.storm@nord.no

# Email Addresses for All Authors:

# Knut Tore Lappegård			knut.tore.lappegard@nordlandssykehuset.no

# Charlotte Björk Ingul			charlotte.b.ingul@ntnu.no

# Erik Waage Nielsen			erik.waage.nielsen@nordlandssykehuset.no

# Bent Aksel Nilsen			bent.aksel.nilsen@nordlandssykehuset.no

# Anette Nyjordet			anette.nyjordet@helgelandssykehuset.no

# Ole-Jakob How			ole-jakob.how@uit.no

# Bjørn Faldaas				bjorn.faldaas@nord.no

# Benjamin Stage Storm		benjamin.s.storm@nord.no

## Author Questionnaire

1. 1. We have marked your project as author-provided footage, meaning you film the video yourself and provide JoVE with the footage to edit. JoVE will not send the videographer. Please confirm that this is correct.

☐ Correct

☐ Incorrect

2. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar? No

3. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

Authors: Please create screen capture videos of the shots labeled as SCREEN, write a screenshot summary, and upload the files to your project page as soon as possible

Author response: We will produce the screen capture videos in conjunction with the experiment on February the 4th. We will test the screen capture shootings in conjunction with experiments on February the 2nd and 3rd

4. Proposed filming date: To help JoVE process and publish your video in a timely manner, please indicate the proposed date that your group will film here: 02/04/2026

When you are ready to submit your video files, please contact our Content Manager, Utkarsh Khare.

Current Protocol Length

Number of Steps: 25

Number of Shots: 56 (23 SC)

# Introduction

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone statements with sufficient context.

INTRODUCTION:

What is the scope of your research? What questions are you trying to answer?

  - Erik Nielsen: We aimed to establish a reproducible ventricular fibrillation-cardioversion model that enhances consistency and reduces animal use in resuscitation research.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

What technologies are currently used to advance research in your field?

  - Charlotte Ingul: We use an implantable cardioverter-defibrillator to reliably induce ventricular fibrillation, perform cardioversion and record intracardiac electrocardiogram.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

CONCLUSION:

What advantage does your protocol offer compared to other techniques?

  - Knut Lappegård: Compared to other methods, our approach achieved a 92% circulation restoration success with artefact-free rhythm monitoring during cardiopulmonary resuscitation.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

What are the current experimental challenges?

  - Knut Lappegård: The main challenge is the need for specialized equipment and skilled operators for ICD implantation and programming.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

How will your findings advance research in your field?

  - Knut Lappegård: Our model increases resuscitation success from ~50% to 90%, enabling repeated ventricular fibrillation-CPR-cardioversion cycles per animal and significantly reduces animal use.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Videographer: Obtain headshots for all authors available at the filming location.

Ethics Title Card

This research has been approved by the Norwegian Animal Research Authority

# Protocol

- Arterial Catheter and Central Venous Catheter (CVC) Placement

Demonstrators: Benjamin Storm, Bent Nielsen

  - To begin, turn the anesthetized pig into a supine position [1-TXT] and fixate the limbs using a gauze bandage [2].

    - WIDE: Talent repositioning the pig into a supine position on the procedure table. TXT: Anesthesia: Pentobarbital (200 - 300 mg IV) + O2 (6 L/min)

    - Talent securing the pig’s limbs with gauze bandages.

  - Disinfect the neck, jaw, groin, and belly using chlorhexidine at a concentration of 5 milligrams per milliliter in ethanol [1] and cover the disinfected areas with sterile drapes [2].

    - Talent applying chlorhexidine solution to the neck, jaw, groin, and belly.

    - Talent placing sterile drapes over the disinfected regions.

  - Stretch a hind leg to extend the iliac artery [1]. Using ultrasound, locate the iliac artery [2] and insert a 5 French Pulse Index Continuous Cardiac Output arterial catheter using the Seldinger technique [3].

    - Talent extending and holding the hind leg to expose the iliac region.

    - Talent placing the ultrasound probe to identify the iliac artery.

    - Talent inserting the arterial catheter using the Seldinger technique under ultrasound guidance.

  - Connect the arterial catheter to a prefilled pressure transducer and flush set [1]. Zero the transducer and confirm a good arterial pressure tracing [2], then release the hind leg [3-TXT].

    - Talent attaching the arterial catheter to the pressure transducer and flush system.

    - SCREEN: Monitor showing a stable arterial pressure waveform.

    - Talent releasing and repositioning the hind leg. TXT: Perform bladder catheterization

Authors: Please create screen capture videos of the shots labeled as SCREEN, write a screenshot summary, and upload the files to your project page as soon as possible : https://review.jove.com/account/file-uploader?src=21138073

If the VGA/HDMI works, that would be good (please follow the instructions pasted at the end of the script).

Also, if this doesn’t work, then no problem. You can film the computer screen/monitors with the camera such that all the text and images are clearly visible and provide us with the footage.

  - Tilt the pig into the Trendelenburg position [1] and increase ventilator positive end-expiratory pressure to 10 centimeters of water to distend the neck veins [2].

    - Talent adjusting the table to place the pig in the Trendelenburg position.

    - Shot of the positive end-expiratory pressure being increased to 10 centimeter H2O.

  - Using ultrasound, locate the left external jugular vein and insert a three-lumen central venous catheter using the Seldinger technique [1]. Verify venous backflow on all lumens and transfer sedation and fluid infusions to the central venous catheter [2]. Then, connect a pressure transducer to the third lumen and confirm a central venous pressure tracing [3].

    - Talent inserting the three-lumen central venous catheter using the Seldinger technique.

    - Talent reconnecting sedation and fluid infusion lines to the central venous catheter.

    - Talent attaching the pressure transducer to the third lumen of the catheter.

- Vascular Introducer and Implantable Cardioverter-Defibrillator (ICD) Electrode Placement

Demonstrators: Benjamin Storm, Bent Nielsen, Knut Lappegård

  - After locating the right external jugular vein, insert a 7 French vascular introducer using the Seldinger technique [1]. Confirm venous backflow and flush the introducer with saline [2].

    - Talent inserting the vascular introducer using the Seldinger technique.

    - Talent flushing the introducer with saline.

  - Decrease ventilator positive end-expiratory pressure to 5 centimeters of water [1] and level the operating table [2-TXT].

    - SCREEN: Ventilator interface showing positive end-expiratory pressure adjusted to 5 centimeter H2O.

    - Talent adjusting the operating table to a level position. TXT: Confirm introducer tip placement by transesophageal echocardiography or fluoroscopy

  - Now, insert a single coil implantable cardioverter-defibrillator electrode through the vascular introducer and advance it into the apex of the right ventricle [1]. Connect the implantable cardioverter-defibrillator lead to a pacing system analyzer using the dedicated connector lead [2-TXT].

    - Talent advancing the implantable cardioverter-defibrillator electrode through the introducer.

    - Talent connecting the lead to the pacing system analyzer. TXT: Intracardiac R-waves must be >6 mV

  - Then, expose the helix of the implantable cardioverter-defibrillator lead [1] and fasten the tip to the endocardium of the right ventricle at the apex or distal septum [2-TXT].

    - Talent extending the helix mechanism of the lead.

    - SCREEN: Imaging view showing the lead tip being fixed to the right ventricular endocardium. TXT: Confirm lead position: Pacing threshold: <2.5 V; 0.4 ms; Impedance: 500-1500 Ω for pacing and 50-100 Ω for shock coil

  - Now, using diathermy, make a 4-centimeter horizontal skin incision below the left clavicle [1]. Using forceps and the fingers, create a small subcutaneous pocket large enough to contain the implantable cardioverter-defibrillator [2].

    - Talent performing a horizontal skin incision with diathermy below the left clavicle.

    - Talent creating a subcutaneous pocket using forceps and fingers.

  - Connect the lead to the implantable cardioverter-defibrillator and insert the device into the subcutaneous pocket [1]. Verify stable right ventricular lead position and secure attachment of the retractable screw to the myocardium [2].

    - Talent connecting the lead to the implantable cardioverter-defibrillator header and placing the device into the subcutaneous pocket.

    - SCREEN: Pacing system analyzer confirming stable electrical parameters.

- Pacemaker/Defibrillator Programming

Demonstrator: Knut Lappegård

  - Switch the pacing system analyzer from pacing system analyzer mode to programmer mode [1]. Using the supplied antenna, connect the implantable cardioverter-defibrillator to the programmer and confirm wireless telemetry [2-TXT].

    - WIDE: Talent navigating the pacing system analyzer interface.

    - SCREEN: Mode selection on the pacing system analyzer switching to Programmer mode. TXT: Repeat measurements of intracardiac R-wave, pacing threshold, and impedances

  - Then, program the pacing mode and the tachyarrhythmia therapy parameters to the desired values [1]. Now, on the programmer screen, select the Test icon and open the Fibber and NIPS folder [2]. Press the Ventricular Fibber icon and select direct current Fibber eleven at 7.5 volts for 2 seconds to enable ventricular fibrillation induction [3].

    - SCREEN: Tachyarrhythmia therapy programming panel showing all the settings.

    - SCREEN: Programmer interface highlighting the Test icon and opening the Fibber and NIPS folder.

    - SCREEN: Selecting the Ventricular Fibber icon and selecting direct current Fibber eleven settings.

  - When ventricular fibrillation is induced, select a therapy delivery option from the lower left corner of the screen [1]. Use Automatic delivery to allow the device to deliver therapy according to programmed parameters [2-TXT].

    - SCREEN: Therapy delivery options displayed on the programmer.

    - SCREEN: Automatic therapy option selected. TXT: Choose Timed or Manual delivery as required

  - Induce ventricular tachycardia by selecting the Ventricular NIPS icon [1]. Choose the Burst icon and set the desired cycle length, such as 300 milliseconds [2], then press Hold to Apply Burst for up to 20 seconds [3].

    - SCREEN: Selection of the Ventricular NIPS icon.

    - SCREEN: Burst pacing settings with cycle length input.

    - SCREEN: Hold to Apply Burst activated for ventricular tachycardia induction.

  - If ventricular tachycardia or ventricular fibrillation is induced, deliver therapy according to the programmed settings in the Tachy folder [1]. Select the Parameters icon to review settings and ensure restoration of sinus rhythm [2].

    - SCREEN: Tachy folder displaying programmed therapy parameters.

    - SCREEN: Parameters icon selected on the device interface and show the rhythm display confirming restoration of sinus rhythm.

- Establishing Non-Invasive Continuous Carotid Doppler Monitoring

Demonstrators: Charlotte Ingul and Bjørn Faldaas

  - Monitor carotid flow using the intra-arterial pressure tracing from the left carotid artery [1] and supplement this, if available, with a non-invasive ultrasound-guided flow velocity measurement in the right carotid artery using a hands-free continuous Doppler [2].

    - Talent observing the monitor displaying the intra-arterial pressure tracing from the left carotid artery.

    - Talent positioning the hands-free Doppler system for continuous monitoring of the right carotid artery.

  - Attach the Doppler probe perpendicular to the left carotid artery [1] and apply generous amounts of ultrasound gel to ensure proper skin contact [2].

    - Talent positioning the Doppler probe perpendicular to the left carotid artery.

    - Close-up shot of ultrasound gel being applied to the skin under the probe.

  - Check the ultrasound monitor to confirm a good carotid flow trace [1]. Next, fixate the probe thoroughly using adhesive tape [2]. Monitor the cardiac rhythm using the intracardiac electrocardiogram obtained from the pacemaker and the superficial electrocardiogram tracings displayed on the monitor [3].

    - SCREEN: Ultrasound monitor showing a clear carotid flow waveform.

    - Talent securing the Doppler probe in place with adhesive tape.

    - Talent observing the pacing system and monitor simultaneously.

  - Confirm the return of spontaneous circulation using the invasive arterial pressure from the left carotid artery on the monitor [1] and the non-invasive ultrasound-guided flow velocity measurement in the right carotid artery [2].

    - SCREEN: monitor displaying invasive arterial pressure tracing.

    - SCREEN: Doppler display showing carotid flow velocity signal.

  - After inducing ventricular fibrillation or ventricular tachycardia, observe cessation of carotid flow on the Doppler trace [1-TXT]. Perform manual cardiopulmonary resuscitation, mechanical compressions, or other interventions as specified in the study protocol [2].

    - SCREEN: Doppler display showing loss of carotid flow signal. TXT: Maintain the no-flow period according to the study protocol

    - Talent performing cardiopulmonary resuscitation or initiating protocol-defined interventions.

  - When required, convert the pig to sinus rhythm by delivering a manual direct current shock through the implantable defibrillator [1]. Adjust the programmed energy level based on the size of the animal [2].

    - SCREEN: Programmer interface showing shock energy selection.

    - SCREEN: Delivery of a manual defibrillation shock through the device.

  - Terminate ventricular fibrillation within 120 seconds to ensure a high conversion rate [1]. If ventricular tachycardia is present, achieve conversion using anti-tachycardia pacing or synchronized cardioversion [2].

    - SCREEN: Therapy delivery timing displayed on the programmer.

    - SCREEN: Anti-tachycardia pacing or synchronized shock being delivered.

  - Finally, allow the animal to rest [1] and ensure normalization of physiological parameters, including blood pressure, end-tidal carbon dioxide, and pH [2].

    - Talent monitoring the resting animal post-intervention.

    - SCREEN: monitor showing stabilization of physiological parameters.

# Results

- Results

  - Using the ventricular fibrillation fibber mode consistently induced ventricular fibrillation or ventricular tachycardia in all 11 animals and cardioverted all animals to sinus rhythm with return of spontaneous circulation [1].

    - LAB MEDIA: Table 1.

  - The success rate for achieving return of spontaneous circulation after induced ventricular fibrillation was 92 percent across all sequences in all animals [1].

    - LAB MEDIA: Table 1. Video editor: Highlight the “Arrhythmia conversion rate” column.

  - On average, 18 arrhythmias, including ventricular fibrillation and ventricular tachycardia, were induced per animal using an implantable cardioverter-defibrillator [1].

    - LAB MEDIA: Table 1. Video editor: Sequentially highlight the second and third columns under the label “pacemaker induced arythmia” consisting “ventricular fibrillation” and “ventricular tachycardia”

NOTE to the Authors:

You have two options for filming and submitting the SCREEN footage:

Option 1: Uploading Separate Clips (Highly Recommended)

Please upload a separate video file for each SCREEN shot (tagged as SCREEN), naming the file according to the shot number (e.g., 2.4.3, 1.5.1, 3.2.1, 3.1.3, etc.).

- Each clip must be a maximum of 20 to 25 seconds to match the voice narration.

- If an action takes longer, provide only the best, most representative 20-25 second segment.

Option 2: Uploading a Single Continuous Video

If you prefer to record all steps together and upload a single video file, that is acceptable. In this case, you must write the precise timestamp for each shot within the script.

For example, if the action described in shot 2.1.2 is present in the file xxx.mp4 and occurs between timestamps 00:30 and 00:45 , then write the following in the script after each shot description or as a comment bubble.

(write the file name and timestamp after each shot): [File name] [Start Time]-[End Time]

- Example 	2.1.2  SCREEN: performing ...........action. xxx.mp4 00:30-00:45

    - SCREEN: cutting the…..........action. xxx.mp4 01:10-01:20

- The duration between the start and end times for each shot must be limited to 20 to 25 seconds.

- If a shot is longer than 25 seconds, choose two short segments (one at the beginning and one at the end) that collectively do not exceed 25 seconds.

- Example: 3.1.1 SCREEN: inserting a needle... xxx.mp4 00:00-00:10 and 02:30-02:40