# Submission ID #: 69531

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21207003

# Title: Whole-Cell Super-Resolution Imaging via DNA-PAINT on a Spinning Disk Confocal with Optical Photon Reassignment

# Authors and Affiliations:

# Cecilia Zaza1, Miruna Tanase1,2, Olivia P. L. Dalby1,2, Sabrina Simoncelli1,2

# 1London Centre for Nanotechnology, University College London

# 2Department of Chemistry, University College London

☒   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Cecilia Zaza			(c.zaza@ucl.ac.uk)

Sabrina Simoncelli		(s.simoncelli@ucl.ac.uk)

# Email Addresses for All Authors:

Cecilia Zaza			(c.zaza@ucl.ac.uk)

Miruna Tanase		(miruna.tanase.18@ucl.ac.uk)

Olivia P. L. Dalby		(olivia.dalby.21@ucl.ac.uk)

Sabrina Simoncelli		(s.simoncelli@ucl.ac.uk)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

If Yes, can you record movies/images using your own microscope camera?

N/A

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

The microscope (CSU-W1 SoRa, Nikon) has a camera port, there is no need to use the scope kit.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

N/A

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Yes

If Yes, how far apart are the locations? 2 min walk, inside UCL campus

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps: 22

Number of Shots: 40

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Answer up to 2 introduction and up to 3 conclusion questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Answer in full sentences, the questions will not be displayed in the video.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

INTRODUCTION:

What are the most recent developments in your field of research?

  - Sabrina Simoncelli: DNA-PAINT now achieves extremely high resolution, with multiplexing and quantitative capabilities, surpassing other SMLM techniques such as PALM and STORM.

What are the current experimental challenges?

  - Sabrina Simoncelli: High-resolution DNA-PAINT mainly works under TIRF or HILO illumination, restricting imaging depth to just a few microns.

CONCLUSION:

What research gap are you addressing with your protocol?

  - Miruna Tanase: We overcame depth limitations by performing DNA-PAINT on a spinning-disk confocal microscope with optical photon reassignment.

What advantage does your protocol offer compared to other techniques?

  - Cecilia Zaza: Our protocol achieves deep-cell and tissue DNA-PAINT imaging without sacrificing resolution, using straightforward, commercially available optics.

How will your findings advance research in your field?

  - Olivia Dalby: Our findings make DNA-PAINT accessible to many laboratories, enabling nanoscale imaging in cells and tissues at depth.

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Deep Whole-Cell Single-Molecule Localization Microscopy Using Spinning Disk Confocal Systems

Demonstrator: Cecilia Zaza

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

AUTHORS: Please note that all pronunciation guides are given in red, italics. Kindly go through the same and change where necessary.

  - To begin, use a spinning disk confocal system equipped with an optical photon reassignment unit and an sCMOS (S-C-M-O-S) camera [1].

    - WIDE: Talent showing the spinning disk confocal microscope setup with camera and lasers visible.

  - Start the system 30 minutes before imaging by turning on the microscope, camera, and lasers [1], then start the computer [2]. Wait for the NIS (N-I-S) -Elements software to fully load [3].
Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=21207003

    - Talent switching on the microscope hardware components.

    - Talent powering on the computer.

    - SCREEN: NIS-Elements software loading on the computer screen.

  - Now select the appropriate imaging mode in NIS-Elements by choosing either standard spinning disk confocal or SoRA (So-Rah) super-resolution before setting acquisition parameters [1].

    - SCREEN: Talent selecting the imaging mode within the NIS-Elements interface.

  - Choose the 60X Apo (A-poh) numerical aperture 1.49 oil immersion objective to maximize detection of single-molecule events [1].

    - Talent rotating the objective turret to select the 60X oil objective.

  - Then select the appropriate magnification based on the objective, camera settings, and sample [1]. Adjust camera binning accordingly to match the selected magnification [2-TXT].

    - SCREEN: Talent selecting magnification settings in the software.

    - SCREEN: Talent adjusting camera binning parameters.TXT: Binning: Effective pixel sizes :108 nm for 1X (no binning) , 4X (4X binning); 78 nm for 2.8X (2X binning).

  - Set the exposure time for DNA-PAINT (D-N-A-Paint) acquisition and synchronize it with the spinning disk rotation speed  [1]. Press the Sync button to automatically match compatible rotation speed [2].

    - SCREEN: Talent entering the exposure time parameter.

    - Talent pressing the Sync button on the spinning disk control pad.

  - Now, configure the laser wavelength, laser power, and filter wheel according to the imager dye spectrum [1].

    - SCREEN: Talent selecting laser wavelength and filter settings in the software.

  - Adjust the laser power to 100 percent using the laser control pad for DNA-PAINT imaging [1].

    - SCREEN: Talent adjusting the laser power slider to maximum.

  - To record the laser power after the objective, place a calibrated laser power meter sensor on the microscope stage near the sample plane [1] and record the measured power [2].

    - Talent positioning the laser power meter sensor on the stage.

    - Shot of the power meter display showing the recorded value.

  - Next, apply a drop of immersion oil to the clean 60X Apo numerical aperture 1.49 oil objective [1]. Secure the prepared immunostained cell sample onto the microscope stage using a stage adapter to minimize drift [2].

    - Talent dispensing immersion oil onto the objective lens.

    - Talent placing and securing the sample on the stage.

  - Switch to brightfield mode and focus on the glass plane using nanoparticles as a reference [1]. Then activate the Perfect Focus System to maintain focus during acquisition [2].

    - SCREEN: Brightfield image used to focus on the glass plane.

    - SCREEN: Talent activating the Perfect Focus System.

  - With Perfect Focus System active, temporarily change the widefield excitation to the green fluorescent protein channel to identify the optimal focal plane [1-TXT]. Use low laser power during this step [2].

    - SCREEN: Green fluorescent protein signal used to identify the focal plane. TXT: For U2OS mEGFP-Nup96 cells/ HeLa Kyoto mEGFP-Nup107 cells, use the 488 GFP signal

    - SCREEN: Laser power set to low percentage.

    - Return to the previously defined SoRA (So-Rah) excitation settings [1].

  - Define the experiment name and saving path [2]. Tick the Time box to set the interval, duration and number of loops [3]. Set the interval to No Delay and enter the number of loops corresponding to total number of frames [4].

    - SCREEN: Talent clicking on SoRa Spinning Disk Confocal imaging mode.

    - SCREEN: Talent entering the experiment name and saving directory.

    - SCREEN: Talent clicking on the Time box.

    - SCREEN: Interval is being set to No Delay and number of loops is being entered.

  - Make sure the PFS box at the bottom is selected, then press Run [1]. Acquire 17,000 to 20,000 frames for the first protein using a 300-millisecond integration time [2].

    - SCREEN: The PFS box is being seen and Run is being pressed.

    - SCREEN: Ongoing time-lapse acquisition displaying frame count progression.

- Single-Molecule Localization and Reconstruction Workflow Using Picasso Software

Demonstrator: Miruna Tanase

  - For image analysis, load the deconvoluted image stack into the Picasso Localize module [1]. Select Analyze and then Parameters from the menu bar and input the camera-specific parameters according to the manufacturer specifications [2].

    - SCREEN: Talent opening the image stack in the Picasso Localize interface.

    - SCREEN: Analyze > Parameters is being selected then camera settings are being entered.
AND
TEXT ON PLAIN BACKGROUND:
For the sCMOS camera
Parameter Setting:
EM gain: 1
Baseline: 100
Sensitivity: 0.23 (conversion factor, electrons/counts)
Quantum efficiency: 0.95 (depending on the used excitation wavelength)

    - Pixelsize (nm): 108 (depending on the imaging configuration)

  - Next, set the box side length to 7 to optimize single-molecule localization precision [1]. Then select the appropriate pixel size corresponding to the imaging configuration [2].

    - SCREEN: Talent entering the box side length value.

  - Choose a suitable minimum net gradient value and enable Preview to visualize detected localizations [1]. Adjust the gradient threshold to minimize background detections [2].

    - SCREEN: Preview of detected localizations is being seen.

    - SCREEN: Adjustment of gradient threshold slider is being done.

  - Now, select Analyze and then Localize to start spot identification and fitting across all frames [1].

    - SCREEN: Menu> Analyse> Localize is being selected.

  - Load the resulting .hdf5 (dot-H-D-F-Five)  into the Picasso Render module to visualize the reconstructed localization map [1].

    - SCREEN: Reconstructed localization map displayed in Picasso Render.

  - Choose Postprocess and then Undrift by RCC [1].  Enter an initial segmentation value starting at 1000 to obtain a xy (X-Y) drift plot [2]. Repeat with lower segmentation values to improve drift correction [3].

    - SCREEN: Postprocess> Undrift by RCC is being selected,

    - SCREEN: Initial segmentation value of 1000 is being entered and xy drift plot is seen.

    - SCREEN: Drift plots corrected and updated plot is seen.

  - Click on File and Save Localization to save the drift-corrected localization file before closing the Render module [1].

    - SCREEN: Talent saving the corrected localization file.

  - Now, load the drift-corrected file into the Picasso Filter module [1] and apply thresholds based on localization precision, point spread function widths, photon counts, and other quality metrics [2].

    - SCREEN: The drift corrected file is being loaded into the Picasso Filter module.

    - SCREEN: Filtering interface showing applied localization quality thresholds.

  - Click on File and Save Localization to save the filtered localization file before closing the Filter module [1].

    - SCREEN: Talent saving the filtered localization file.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 186

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - DNA-PAINT imaging of nucleoporin 96 tagged with monomeric enhanced green fluorescent protein in U2OS (U-Two-O-S) cells produced a clear super-resolved nuclear pore complex pattern across the field of view [1]. Zoomed-in views revealed paired nucleoporin 96 subunits arranged with the expected 8-fold symmetry of nuclear pores [2].

    - LAB MEDIA: Figure 1C.

    - LAB MEDIA: Figure 1D. Video editor: Highlight the paired bright spots indicated by arrows in individual nuclear pores.

  - Measurement of the Euclidean distance between aligned nucleoporin 96 pairs showed a peak-to-peak separation of 13 plus or minus 2 nanometers [1].

    - LAB MEDIA: Figure 1E. Video editor: Show the cross-sectional histogram with the distance marked between the two peaks.

  - Nucleoporin 96, mitochondria, and microtubules were simultaneously visualized in U2OS cells with preserved single-molecule localization precision across all targets [1].

    - LAB MEDIA: Figure 2B. Video editor: Show the multicolor reconstruction with nuclear pore complexes in blue, mitochondria in red, and microtubules in green.

  - Large field-of-view imaging of microtubules in HeLa cells demonstrated that localization precision varied with magnification and field size, with reduced precision observed at lower magnification [1].

    - LAB MEDIA: Figure 3A.

  - Higher magnifications produced improved and more uniform localization precision compared with 1X magnification [1].

    - LAB MEDIA: Figure 3D. Video editor: Please highlight the dots corresponding to 1x, 2.8x and 4x sequentially

  - Whole-cell imaging was demonstrated by DNA PAINT acquisition of the microtubule network across multiple axial planes spanning the full cell depth [1]. Localization precision gradually decreased with increasing imaging depth [2].

    - LAB MEDIA: Figure 4A (left images). Video editor: Sequentially show images from z= 0µm to z = 9 µm

    - LAB MEDIA: Figure 4A (right plot).

  - Double-walled microtubule structures were resolved at multiple axial positions with peak-to-peak distances between 30 nanometers and 40 nanometers [1].

    - LAB MEDIA: Figure 4C. Video editor: Sequentially show images 1 – 3 with their corresponding histogram plots