TITLE:

Development and Standardization of an Ex Vivo Micromethod for Intracellular Quantification of Vincristine in Primary ALL Cells by LC-MS/MS

AUTHORS AND AFFILIATIONS:

Adrielli Caroline Soares1, Guilherme Reis-de-Oliveira2, Samara de Sousa Mariano1*, Juliana Ronchi Corrêa1*, Amilcar Cardoso de Azevedo3, Fabio Gozzo2, José Andrés Yunes1

1Leukemia Biology Laboratory, Boldrini Children's Center, Campinas, Brazil
2Mass Spectrometry Laboratory, Boldrini Children's Center, Campinas, Brazil

3Boldrini Children's Center, Campinas, São Paulo, Brazil

E-mail addresses of the co-authors:

Guilherme Reis-de-Oliveira		(guioliveirareis@gmail.com)

Samara de Sousa Mariano		(samarasouzama@gmail.com)

Juliana Ronchi Corrêa			(juliana.ronchi@gmail.com)

Amilcar Cardoso de Azevedo		(amicar.cardoso@boldrini.org.br)

Fabio Gozzo				(fgozzo@gmail.com)

José Andrés Yunes			(andres@boldrini.org.br)

Corresponding author:

Adrielli Caroline Soares		(dricaroline8@gmail.com)

SUMMARY:

This protocol aimed to establish and optimize an LC-MS/MS method for the relative quantification of vincristine accumulation in primary pediatric ALL cells.

ABSTRACT:

Quantifying intracellular vincristine in primary pediatric acute lymphoblastic leukemia (ALL) cells is essential for understanding sample-specific differences in drug uptake and for supporting experimental studies using scarce patient material. Here, we present a streamlined LC–MS/MS micromethod specifically optimized for primary lymphoblasts, addressing key challenges such as limited cell availability, small cell size, and the need for reproducible recovery after multiple washing steps. Time-course assays ranging from 1 to 5 h demonstrated that a 3 h incubation period yields the highest and most consistent intracellular vincristine levels, with lower variability across replicates. Method optimization also established that using 2.5 × 10⁶ cells per condition improves analytical robustness compared with 1 × 10⁶ cells, which showed greater dispersion in repeated experiments. Additional critical refinements included the use of cold methanol to enhance protein precipitation efficiency and the incorporation of an internal standard to monitor and minimize procedural variation. The method enables sensitive and precise quantification of intracellular vincristine across relevant concentration ranges, although the lowest tested dose (0.01 µM) was not quantifiable after the 3 h incubation. Overall, this micromethod provides a practical, reproducible, and scalable approach for measuring intracellular vincristine in primary ALL samples. Its design supports comparative analyses, validation of ex vivo models, and future methodological applications aimed at integrating intracellular drug quantification into broader pharmacological assessments.

INTRODUCTION:

Acute lymphoblastic leukemia (ALL) accounts for approximately 25% of all childhood cancers and remains the most common pediatric malignancy1,2. Although advances in therapy have markedly improved survival rates—now exceeding 90%—the biological complexity of ALL continues to drive efforts to develop tools that better characterize the disease at the cellular level3.

Vinca alkaloids, particularly vincristine (VCR), are key components of ALL treatment protocols4. Because VCR exerts its cytotoxic effects intracellularly, understanding its uptake and accumulation in leukemic cells is fundamental for studies involving drug response, pharmacodynamics, and cellular behavior in ex vivo systems. However, most available approaches for assessing intracellular drug levels rely on indirect measurements or semi-quantitative methods, limiting reproducibility and sensitivity5-7.

In parallel, many investigations still depend heavily on leukemia cell lines, which lack the biological heterogeneity and clinical relevance of patient-derived primary cells. Primary samples often present challenges such as limited cell numbers, variable sample quality, and the need for highly sensitive analytical techniques capable of detecting low-abundance intracellular compounds8.

Liquid chromatography–tandem mass spectrometry (LC-MS/MS) provides the specificity and sensitivity required for precise quantification of chemotherapeutic agents, yet standardized protocols for measuring intracellular VCR directly in primary leukemic cells remain scarce. Establishing robust micromethods compatible with limited patient material may facilitate a broad range of applications, including pharmacological profiling, evaluation of drug–cell interactions, and comparative studies across disease subtypes or therapeutic conditions.

Here in this protocol, we describe the development and optimization of a micromethod for ex vivo quantification of intracellular VCR in primary leukemic cells using LC-MS/MS. To obtain sufficient material for method development, patient samples were transplanted into immunodeficient NSG mice to generate patient-derived xenograft (PDX) models9. After in vivo expansion, human leukemic cells were isolated and cryopreserved10. After, they were thawed, incubated with VCR, washed, lysed, and analyzed by LC-MS/MS to determine intracellular drug levels.

This method offers several advantages: it requires a small number of cells (2.5 million), is sensitive and rapid, and is compatible with clinically relevant primary samples. Because it directly measures intracellular drug content, it provides a reliable quantitative tool for researchers investigating drug uptake, cellular responses to chemotherapeutics, or other pharmacological questions in hematologic malignancies.

Overall, this protocol expands the analytical toolkit available for studying pediatric ALL and can be readily adapted for other drugs or sample types, supporting future investigations into leukemia biology, drug behavior, and precision pharmacology.

PROTOCOL:

The leukemia samples used in this study were obtained from bio-banked vials of patients treated at Centro Infantil Boldrini, who allowed the use of their cells through an informed consent form. The projet was approved by the Institutional Ethics Committee (CAAE 34601120.7.0000.5376). he protocols for the generation of patient-derived xenografts (PDX) in animals were carried out in accordance with the regulations and ethical guidelines of the Animal Ethics Committee (Comissão de Ética no Uso de Animais) of Centro Infantil Boldrini (CEUA/Boldrini 0011-2020).

- Sample preparation and drug incubation (Day 1)

  - Preparation of stock and working solutions

    - Weigh 2.5 mg of Vincristine sulfate (molecular weight: 923 g/mol) and dissolve it in 2.0 mL of dimethyl sulfoxide (DMSO) in a 2.0 mL microtube to prepare a 1.354 mM stock solution. Vortex until fully dissolved.

CAUTION: DMSO is an irritant; wear gloves and handle it in a chemical fume hood.

    - Prepare the 100 µM working solution by mixing 147.71 µL of the VCR stock solution with 1,852.28 µL of DMSO in a 2.0 mL microtube. Vortex thoroughly.

    - Distribute 500 µL of the working solution into four 0.6 mL microtubes and store at -20 °C for up to several months. Avoid thawing the aliquots more than three times.

    - Prepare a 100 µM vinblastine (VINB) working solution by mixing 181.81 µL of a 1.1 mM stock solution (Vinblastine sulfate 1 mg/mL, molecular weight: 909.07 g/mol) with 1,818.19 µL of sterile ultrapure water in a 2.0 mL microtube. Vortex until homogeneous.

    - Aliquot the VINB working solution the same way as done for VCR.

    - Prepare methanol + 0.1 µM VINB solution by mixing 49.95 mL of methanol with 50 µL of the VINB working solution in a 50 mL tube. Vortex until homogeneous. Store it at -20 °C up to 3 months.

CAUTION: Methanol is toxic and flammable; handle with appropriate PPE in a fume hood.

    - Prepare a water + 0.1% formic acid solution by mixing 49.95 mL of ultrapure water with 50 µL of formic acid in a 50 mL tube. Vortex well. Store at room temperature (RT).

CAUTION: Formic acid is corrosive; handle with gloves and eye protection.

  - Thawing and washing primary cells

    - Thaw cryopreserved cell samples by gently swirling the vial in a beaker containing 15 mL of 37 °C distilled water until fully thawed.

NOTE: Information regarding the cell collection and freezing protocols can be found in previous publications9,10.

    - Transfer the contents of the vial to a 50 mL conical tube containing 10–15 mL of sterile 1x PBS with a sterile disposable Pasteur pipette. Centrifuge at 300 × g for 5 min at RT.

    - Discard the supernatant by inverting the tube.

    - Resuspend the pellet in 10 mL of AIM V medium with a disposable Pasteur pipette. Count viable cells and measure average cell size using an automated counter.

NOTE: Using cells that exhibit at least 80% viability relative to the total number of cells collected is recommended.

    - Calculate the volume needed to collect 2.5 × 10⁶ viable cells using a simple proportion based on the measured concentration. For example, if the concentration is 5 × 10⁶ cells/mL, the required volume is 0.5 mL.

NOTE: Perform each treatment in triplicate for data reproducibility. Save the medium size of the cells provided in the automated cell counter; use it to calculate the volume of the cells.

  - Drug treatment setup

    - Calculate the volume of VCR working solution required to reach a final concentration of 1 µM in a total volume of 2.0 mL. Vortex the stock solution before pipetting.

    - Add AIM V medium, then 2.5 × 10⁶ cells, and finally the VCR solution to a 2.0 mL microtube. Vortex briefly and spin down lightly (5 s) to collect any liquid from the tube cap.

    - Incubate the open tubes at 37 °C, 5% CO₂ for 3 h.

    - After incubation, close the lids and centrifuge the samples at 300 × g for 5 min at RT.

    - Carefully discard the supernatant using a pipette, then add 1 mL of PBS at RT to the tube.

    - Repeat the centrifugation at 300 × g for 5 min at RT, supernatant removal, and PBS wash two additional times, for a total of three washes.

    - Carefully remove all supernatant using a pipette, resuspend the pellet in 100 µL of PBS, and vortex until the pellet is fully resuspended.

    - Add 400 µL of cold methanol + 0.1 µM vinblastine (-20 °C) directly to the sample, avoiding contact with the tube walls.

    - Store the samples overnight at -80 °C.

NOTE: Samples can be stored at -80 °C for a few weeks (safe stop).

- Sample processing for LC-MS/MS (Day 2)

  - Sample extraction

    - Pre-cool the centrifuge to 4 °C before removing samples from the -80 °C freezer.

    - Immediately after taking the samples from the freezer, centrifuge them at 13,000 × g for 20 min at 4 °C.

    - Transfer 300 µL of supernatant to a 1.5 mL microtube.

    - Dry the samples completely using a vacuum concentrator at 40–45 °C for 1 h and 40 min.

NOTE: Dried pellets can be stored at -20 °C for a few weeks (safe stop).

    - Resuspend each pellet in 400 µL of water + 0.1% formic acid (previously prepared). Vortex gently until the pellet is fully dissolved.

    - Centrifuge the samples again at 13,000 × g for 20 min at 4 °C.

    - Transfer 300 µL of supernatant to a 2.0 mL screw-top vial for LC-MS/MS analysis.

- LC-MS/MS setup and data acquisition

  - Mobile phase preparation and column conditioning

    - Prepare Solvent A by mixing ultrapure water with 0.1% formic acid.

    - Prepare Solvent B by mixing acetonitrile with 0.1% formic acid.

CAUTION: Acetonitrile is toxic and flammable; handle in a chemical fume hood.

    - Sonicate each solvent in an ultrasonic bath for 5 min. Connect the bottles to the pump module and purge the system at high flow to remove contaminants.

    - Install a reversed-phase BEH C18 column (1.7 μm, 2.1 mm × 50 mm).

NOTE: To perform the analysis, this study used an Acquity H-Class chromatographic system coupled to a Xevo TQS-Micro triple-quadrupole mass spectrometer.

    - Use a binary gradient composed of Solvents A and B, both previously prepared.

    - Run the chromatographic gradient from 5% to 95% of Solvent B over the first 2 min, followed by 2 min of Solvent A for column washing and re-equilibration.

    - Set the flow rate to 0.2 mL/min throughout the 4-min run. Maintain the column temperature at 40 °C and inject 10 µL of each sample.

  - LC-MS/MS method setup

    - Open the LC-MS/MS software. Click File > New > MS Method to start a new method.

    - Go to the Function tab and set the function type to multiple reaction monitoring (MRM). Enter ion transitions and parameters listed in Table 1 for VCR and VINB (internal standard).

    - Set the Acquisition Time to 4 min. Click Apply to save.

NOTE: Do not exceed 4 min to ensure column re-equilibration between injections.

    - Configure UHPLC parameters under Inlet Method > Edit LC Method and save as a .mth file (Table 2).

    - Set source parameters under the Tune tab according to Table 3. Save as a .tun file.

    - Save the complete method: File > Save As, and Export as a .mth file.

  - Creating a sample list and starting the run

    - In the software's home screen, open Sample List and create a list with the sample names, injection volumes, vial positions, and the method to be used.

    - Place a vial with ultrapure water in the autosampler tray and designate it as a blank. Run the blank between every set of triplicate injections to avoid carryover.

NOTE: Before starting the run, ensure the column pressure (ΔP) is stable and less than 10 psi.

    - Select the desired rows in the Sample List and click Run > OK to begin data acquisition.

- Data analysis

  - Data visualization and export

    - After the run completes, view chromatograms by selecting a sample and clicking Chromatogram in the top menu bar.

    - On the home screen, click Browse and navigate to the folder where the ".raw" files are saved. Select the files and click Process > Process All to generate output tables.

    - Click Report > Export Table and save the processed data as a .csv or .xlsx file.

NOTE: Export raw peak areas for both the analyte and internal standard.

  - Data organization and ratio calculation

    - Open the exported file and create the following columns:

      - Sample: Add the sample or vial name.

      - Run: Indicate which replicate the value corresponds to.

      - VCR Area: Input the area under the peak for vincristine.

      - VINB Area: Input the area under the peak for vinblastine.

      - VCR/VINB Ratio: Normalize each VCR value by dividing it by its corresponding VINB value.

    - Add columns for standard deviation values of VCR area, VINB area, and VCR/VINB ratio for each triplicate set.

NOTE: Normalization using the internal standard allows correction for variation in extraction or injection volume.

  - Plotting results

    - Open the preferred graphing software and select the Grouped Table format.

    - Choose Enter or import data into a new table, and under Options, select Enter 3 replicate values in side-by-side subcolumns.

    - Plot only the normalized VCR/VINB ratios for each sample, omitting raw area values.

REPRESENTATIVE RESULTS:

Representative chromatographic results obtained after intracellular accumulation assays using primary ALL PDX cells incubated with VCR and VINB are shown in Figure 1. Samples were prepared using 2.5 × 10⁶ cells per condition. Under the chromatographic conditions described, VCR was consistently retained at 2.85 min with minimal matrix effect, and peak shape was symmetrical and well-resolved. These profiles illustrate the expected chromatographic behavior of VCR and VINB when the protocol is executed correctly.

Prior to applying this protocol to primary leukemia cells, preliminary assays were performed to determine the method's linearity, limit of detection (LOD), optimal cell number, and incubation time. These steps are not detailed in the present protocol, but representative results from the calibration curve are shown below as part of the method validation.

To assess the sensitivity and linearity of the method, calibration curves for VCR were constructed under two experimental conditions. In AIM V media, concentrations ranging from 0.001 µM to 10 µM showed excellent linearity after log₁₀-transformation (r² = 0.9986). When 2 × 10⁶ RS4;11 cells were added to the medium, the linear range was evaluated from 0.01 to 1 µM, also displaying a strong correlation (r² = 0.9967) (Figure 2). These results confirm that the method maintains accuracy and sensitivity both in cell-free media and in the presence of cells, supporting its applicability to intracellular quantification.

Figure 3 shows the incubation test to determine the time of incubation of the cells (lineage and PDX cells) with VCR, which showed an equilibrium between influx and efflux. We incubated 1 x 106 of lineage RS 4;11 and 2,5 x 106 primary cells with VCR from 1 h to 5 h, in triplicate, and then analyzed the samples. Analysis of the VCR/VINB ratio in RS4;11 cells revealed a plateau in intracellular accumulation between 2 h and 3 h, followed by a slight increase at 4–5 h. In contrast, PDX-derived primary cells showed a peak at 2 h, with a progressive decline thereafter. These differences likely reflect intrinsic biological properties: immortalized RS4;11 cells maintain higher metabolic stability and tolerate longer drug exposure, whereas primary cells exhibit reduced proliferative capacity and are more susceptible to prolonged incubation. Although the 2-h time point yielded the highest uptake in primary cells, it also showed substantial variability across replicates. Therefore, 3 h was selected as the standard incubation period, representing a balanced condition with adequate accumulation, improved reproducibility, and preserved cell viability across models.

Figure 4 shows the results of intracellular accumulation of VCR in six different PDX patient samples. Quantification was based on the ratio of the peak area of VCR to that of VINB, used as an internal standard. To account for differences in cell size, this ratio was further normalized by the average cell volume of each patient sample. A positive correlation was observed between the unnormalized VCR/VINB ratios and cell volume (Spearman, one-tailed; p = 0.0075), indicating that cell size may influence apparent drug accumulation. Therefore, the additional normalization by cell volume was applied to minimize this confounding factor. These data illustrate inter-individual variability in VCR accumulation, which may reflect differences in ABC transporter activity, membrane composition, or intracellular compartmentalization across patient-derived samples.

The fragmentation patterns of VCR and VINB are presented in Figure 5. For VCR, the base peak corresponds to the precursor ion at m/z 413.38, from which the quantification ion (m/z 392.47) and the confirmation ion (m/z 383.38) are generated using 40 V of cone energy and 16 V of collision energy (CE). For VINB, because the energies required to obtain the product ions differ, the spectra are shown separately. In both spectra, the precursor ion at m/z 406.51 is observed; however, the quantification ion (m/z 376.39) is obtained using 40 V of cone energy and 24 V of collision energy, whereas the confirmation ion (m/z 271.70) is generated using 2 V of cone energy and 18 V of CE.

An example of a poorly defined chromatographic peak is presented in Figure 6, corresponding to VCR at a concentration of 0.001 µM. Under these conditions, the peak showed reduced intensity and broadening, likely due to limitations in the quantification of the method after incubation in cells at low concentrations. This highlights the importance of using appropriate concentration ranges and reinforces the lower limit of quantification established by the method validation.

FIGURE AND TABLE LEGENDS:

Figure 1: Representative chromatograms of vincristine (VCR) and vinblastine (VINB) after intracellular accumulation assay in primary ALL PDX cells. Chromatographic profiles were obtained after incubation with 1 µM VCR for 3 h at 37 °C using 2.5 × 10⁶ cells per condition. VCR was consistently detected at 2.85 min with minimal matrix interference and adequate peak symmetry. VINB was used as an internal standard and was added in a solution with methanol after washing the cells. These peaks illustrate expected results when the extraction and chromatographic procedures are properly executed.

Figure 2: Calibration curve for VCR quantification by LC-MS/MS. (A) Standard curve generated in AIM V media using VCR concentrations from 0.001 to 10 µM. (B) Standard curve generated in AIM V media containing 2 × 10⁶ RS4;11 cells, using VCR concentrations from 0.01 to 1 µM. In both cases, data were log₁₀-transformed to account for exponential spacing, and the linear regression equation with correlation coefficient (r² = 0.9986 and r² = 0.9967, respectively) is shown. These curves illustrate the method's dynamic range, sensitivity, and detection limits under the tested conditions.

Figure 3: Intracellular accumulation of VCR over time in RS4;11 and PDX-derived primary ALL cells. Time-course analysis (1–5 h) of the VCR/VINB ratio shows a 2–3 h plateau in RS4;11 cells, followed by a modest increase at later time points. Primary PDX-derived cells display a peak at 2 h, with decreasing accumulation at 4–5 h. The 3-h incubation was selected as the standard condition due to its balance between intracellular uptake, reproducibility, and cell viability across both models.

Figure 4: Intracellular accumulation of VCR in six different PDX patient samples. Data represent the ratio of VCR to VINB peak areas normalized by average cell volume for each patient. Values highlight inter-individual variability in vincristine accumulation across primary samples. Each sample was analyzed in triplicate, and error bars represent standard deviation.

Figure 5: Fragmentation patterns of VCR and VINB. (A) MS/MS fragmentation spectrum of VCR showing the precursor ion (base peak, m/z 413.38) and the product ions used for quantification (m/z 392.47) and confirmation (m/z 383.38), obtained with 40 V cone energy and 16 V collision energy. (B) MS/MS spectrum of VINB showing the precursor ion (m/z 406.51) and the quantification product ion (m/z 376.39), acquired with 40 V cone energy and 24 V collision energy. (C) MS/MS spectrum of VINB showing the precursor ion (m/z 406.51) and the confirmation product ion (m/z 271.70), acquired with 2 V cone energy and 18 V collision energy.

Figure 6: Chromatographic profile of VCR at low concentration in lineage cells (RS4;11), illustrating poor peak definition. At 0.01 µM VCR in 1 × 10⁶ RS4;11 cells after 3 h of incubation, the chromatographic peak is broad, with low intensity and signal-to-noise ratio, illustrating the lower limit of the method's detection capability. These results emphasize the importance of working within validated concentration ranges for reliable quantification.

Table 1: LC-MS/MS acquisition parameters. Mass spectrometry parameters for VCR (analyte) and VINB (internal standard), including precursor and product ion transitions, polarity mode, cone voltage, collision energy, and retention times used for MRM acquisition.

Table 2: UHPLC parameters. Chromatographic conditions used for LC separation of vincristine and vinblastine, including column temperature, flow rate, gradient mode, mobile phases, and injection volume.

Table 3: Ion source parameters. Ionization source settings applied during LC-MS/MS analysis, including capillary voltage, source and desolvation temperatures, and gas flow rates.

DISCUSSION:

Accurate assessment of intracellular drug levels in primary leukemic cells is essential for understanding drug handling at the cellular level and for supporting the development of functional assays that complement molecular profiling. Existing analytical approaches often require large numbers of cells, involve lengthy processing steps, or rely on indirect readouts that do not reflect the true intracellular availability of chemotherapeutic agents5,6. These limitations highlight the need for rapid, sensitive, and scalable methods capable of quantifying intracellular drug accumulation directly in patient-derived samples.

This protocol describes a micromethod for the relative quantification of intracellular VCR in primary pediatric acute lymphoblastic leukemia (ALL) cells using LC-MS/MS. he method was developed to provide a direct and sensitive assessment of intracellular drug content, addressing the limitations of traditional analytical approaches that often depend on bulk cell material, semi-quantitative detection strategies, or indirect surrogate markers. By enabling the measurement of VCR within a small number of primary cells, this micromethod offers a practical and reproducible framework for investigating drug uptake, intracellular handling, and pharmacological behavior in ex vivo leukemia models.

A critical aspect of this protocol is the number of cells used per condition. Based on optimization experiments, we standardized the use of 2.5 × 10⁶ primary ALL cells, as these cells are relatively small and the protocol involves several washing steps that may lead to cell loss11. Using 1 × 10⁶ cells resulted in higher variability and less consistent intracellular VCR measurements, making 2.5 × 10⁶ a more reliable input. Equally important is the use of samples with ≥80% viability, which helps ensure consistent uptake and reduces artifacts related to compromised cell membranes. Another key step is the addition of cold methanol (–20 °C), which improves protein precipitation efficiency and stabilizes the analytes during extraction12. The inclusion of vinblastine as an internal standard is also essential, as it minimizes analytical variation across samples and enables accurate normalization throughout the workflow13.

Time of incubation is one of the key parameters that may require adjustment depending on the cellular model. In the optimization experiments using RS4;11 cells and PDX-derived primary ALL cells, intracellular VCR accumulation showed distinct temporal profiles: RS4;11 cells exhibited a plateau between 2–3 h, whereas primary cells displayed a peak at 2 h followed by decreased accumulation at later time points. However, the 2-h condition showed substantial variability in primary samples, while the 3-h condition provided more consistent measurements across replicates. For this reason, 3 h was selected as the standard incubation time. Laboratories implementing this protocol may need to adjust the time course depending on cell type, viability, or metabolic stability, and we recommend performing a preliminary time-course assay (e.g., 1–5 h) to identify the optimal incubation period for each model. Furthermore, researchers working with larger or slower-growing cells may consider increasing the cell input or adjusting the washing steps to avoid excessive cell loss.

This micromethod presents some limitations that should be considered when adapting it to different experimental contexts. First, extremely low extracellular VCR concentrations may fall below the quantificable range after incubation; for example, 0.01 µM VCR was not quantifiable in our experiments using 1 × 10⁶ RS4;11 cells after 3 h of incubation, indicating that detection sensitivity is constrained by both drug concentration and cell number. Additionally, because the method relies on relative quantification, variability in cell size or volume can influence raw VCR/VINB ratios, making volume normalization essential when comparing different models. The protocol also requires a minimum of 2.5 × 10⁶ small lymphoblast-like cells to ensure robust signal after multiple washing steps, which may pose challenges when working with limited patient samples.

The significance of this micromethod lies in its ability to provide a direct and sensitive measurement of intracellular vincristine in primary leukemic cells using a small number of cells and a rapid analytical workflow. Existing approaches for assessing intracellular drug levels in hematological samples often require larger cell inputs, extensive extraction procedures, or longer chromatographic runs, limiting their applicability in studies involving scarce patient-derived material. By combining a short 4-min LC-MS/MS acquisition with efficient protein precipitation and a robust internal standard, this protocol enables reproducible quantification while maintaining compatibility with clinical or preclinical samples. The method therefore fills an important technical gap for researchers who aim to evaluate intracellular drug exposure in primary cells using a streamlined and scalable assay.

This micromethod can be readily extended to additional applications beyond vincristine quantification. With appropriate optimization of MRM transitions and extraction parameters, the workflow can be adapted to measure other chemotherapeutic agents or small molecules of interest in hematologic models. Its low sample requirement makes it suitable for pharmacokinetic–pharmacodynamic studies in ex vivo cultures, for testing drug uptake profiles in patient-derived xenografts, or for evaluating intracellular drug levels in limited clinical samples. The protocol may also support the development of multiplexed assays for simultaneous quantification of multiple compounds, provided that chromatographic separation and ionization conditions are optimized. As quantitative mass spectrometry continues to advance, this workflow provides a practical foundation for integrating intracellular drug measurements into broader analytical or translational research pipelines.

ACKNOWLEDGMENTS:

We thank the Centro Infantil Boldrini and all the patients who donated their samples for the study.  Adrielli C. Soares received a fellowship from the Coordenação de Aperfeiçoamento de Pessoal de Nível Superior—Brasil (CAPES) and from the São Paulo Research Foundation (FAPESP, 2023/18389-3). José A. Yunes received a productivity fellowship from the National Counsel of Technological and Scientific Development (CNPq, 308399/2021‐8). This work was supported by research funding from the Brazilian Ministry of Health by the PRONON program (Programa Nacional de Apoio à Atenção Oncológica, NUP 25000.211174/2019‐45).

DISCLOSURES:

The authors declare no conflicts of interest.

REFERENCES:

- Inaba, H., Pui, C.H. Advances in the diagnosis and treatment of pediatric acute lymphoblastic leukemia. J Clin Med. 10 (9), 1926 (2021).

- Pui, C.H. et al. Childhood acute lymphoblastic leukemia: progress through collaboration. J Clin Oncol. 33 (27), 2938–2948 (2015).

- Gaudichon, J. et al. Mechanisms of extramedullary relapse in acute lymphoblastic leukemia: reconciling biological concepts and clinical issues. Blood Rev. 36, 40–56 (2019).

- Dhyani, P. et al. Anticancer potential of alkaloids: a key emphasis on colchicine, vinblastine, vincristine, vindesine, vinorelbine and vincamine. Cancer Cell Int. 22, 206 (2022).

- Mateus, A., Matsson, P., Artursson, P. Rapid measurement of intracellular unbound drug concentrations. Mol Pharmaceutics. 10 (6), 2467–2478 (2013).

- Chien, H. C. et al. Rapid method to determine intracellular drug concentrations in cellular uptake assays: application to metformin in organic cation transporter 1-transfected human embryonic kidney 293 cells. Drug Metab Dispos. 44 (3), 356–364 (2016).

- Calcagno, A. M., Ambudkar, S.V. Analysis of expression of drug resistance-linked ABC transporters in cancer cells by quantitative RT-PCR. Methods Mol Biol. 637, 121–132 (2010).

- Richter, M. et al. From donor to the lab: a fascinating journey of primary cell lines. Front Cell Dev Biol. 9, 648272 (2021).

- Borgmann, A. et al. Childhood ALL blasts retain phenotypic and genotypic characteristics upon long-term serial passage in NOD/SCID mice. Pediatr Hematol Oncol. 17 (8), 635–650 (2000).

- Rosati, M. M., Corrêa, J. R., de Sousa Mariano, S., Cardoso, A., Andrés Yunes, J. Murine model of leukemia relapse to induction chemotherapy for acute lymphoblastic leukemia. J Vis Exp. 224, e68797 (2025).

- Dams, R., Huestis, M. A., Lambert, W. E., Murphy, C. M. Matrix effect in bioanalysis of illicit drugs with LC–MS/MS: influence of ionization type, sample preparation, and biofluid. J Am Soc Mass Spectrom. 14 (11), 1290–1294 (2003).

- Yang, Y., Cruickshank, C., Armstrong, M., Mahaffey, S., Reisdorph, R., Reisdorph, N. New sample preparation approach for mass spectrometry-based profiling of plasma results in improved coverage of metabolome. J Chromatogr A. 1300, 217–226 (2013).

- Zhang, J., Dasgupta, A., Ayala, R., Bonapace, C., Kassim, S., Faustino, P.J. Internal standard variability: root cause investigation, parallelism for evaluating trackability and practical considerations. Bioanalysis. 16 (15), 771–776 (2024).