# Submission ID #: 69640

# Scriptwriter Name: Poornima G

# Project Page Link: https://review.jove.com/account/file-uploader?src=21240248

# Title: Construction Methods for Hybrid Polycarbonate-Silicon Microfluidic Devices

# Authors and Affiliations:

# Nicholas C. Higgins1, Jacqueline C. Kading2, David G. Blauvelt3,4, Shuvo Roy1*

# 1Department of Bioengineering and Therapeutic Sciences, University of California, San Francisco

# 2Department of Pediatrics, Division of Neonatology, University of California, San Francisco

# 3Department of Pediatrics, Nemours Children’s Health

# 4Department of Pediatrics, Sidney Kimmel Medical College at Thomas Jefferson University

# Corresponding Authors:

# Shuvo Roy					shuvo.roy@ucsf.edu

# Email Addresses for All Authors:

# Nicholas C. Higgins				nicholas.higgins@ucsf.edu

# Jacqueline C. Kading				jacqueline.kading@ucsf.edu

# David G. Blauvelt				david.blauvelt@nemours.org

# Shuvo Roy					shuvo.roy@ucsf.edu

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  No

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. No

Current Protocol Length

Number of Steps:  21

Number of Shots:  51

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context

Answers are slightly edited in accordance with journal guidelines.

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Jackie Kading: Our research focuses on bonding hybrid microfluidic devices for prototyping in academic lab settings. We aim to answer the question of what are low-cost yet optimal bonding strategies for a hybrid microfluidic channel.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Jackie Kading: Bonding thermoplastics to rigid substrates in hybrid microfluidics is challenging, and conventional high-temperature, high-pressure methods are unsuitable; this protocol provides a practical, research-friendly bonding approach.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Jackie Kading: This protocol allows researchers to study different low-cost bonding modalities suitable for bonding hybrid microfluidic devices in a laboratory setting.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Describe the most important consideration or challenge when performing this protocol.

  - Jackie Kading: The most important consideration when performing this protocol is ensuring precision in both the design fabrication and alignment of the layers for a leak-tight interface and adequate bond surface area.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Jackie Kading: Future studies can compare sealing surface areas across bonding methods and evaluate this protocol’s applicability to hybrid microfluidic systems incorporating brittle or diverse substrate materials.

    - INTERVIEW: Named talent says the statement above in an interview-style shot, looking slightly off-camera

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

Videographer: Please always record an audio slate while filming each shot mentioned in the shotlist.

- Solvent Bonding of Polycarbonate Microfluidic Device with Integrated Silicon Chip

Demonstrators: Jackie Kading, Charlotte Li

  - To begin, obtain the polycarbonate housing to achieve [1] a flat bonding surface [2]. After cleaning the components, integrate the silicon chip using epoxy, and allow the assembly to cure [2-TXT].

    - WIDE: Talent placing the polycarbonate housing on the work surface.

    - Close-up of showing the polycarbonate housing.

    - Talent placing the silicon chip into the housing using epoxy. TXT: Seal any residual gaps with fresh epoxy

  - For solvent bonding, laser-cut the adhesive tape to match the shape of the channel walls. [1]

    - Talent laser-cutting the adhesive tape to match the shape of the channel walls.

  - To apply adhesive tape, remove 1 liner layer from the tape [1]. Then, align and apply the tape to the channel wall surfaces using dowel pins for alignment [2].

    - Talent removing 1 liner layer from the tape.

    - Talent aligning and applying the tape to the channel wall surfaces with dowel pins.

  - Coat the fluidic channel with polyvinyl alcohol [1]. Add enough 1% polyvinyl alcohol solution to cover the entire fluidic channel area [2] and allow the polyvinyl alcohol solution to dry completely at room temperature [3].

    - Talent adding polyvinyl alcohol to the fluidic channel.

    - Talent showing the fully covered fluidic channel area.

    - Talent leaving the polyvinyl alcohol-coated channel aside to dry at room temperature.

  - To remove the adhesive mask once the polyvinyl alcohol layer has fully cured [1], carefully peel off the double-sided adhesive, leaving a polyvinyl alcohol-coated fluidic channel [2]. Inspect and clean any residual adhesive from the channel walls if necessary [3].

    - Talent showing the fully cured device

    - Talent peeling off the double-sided adhesive fully to leave a polyvinyl alcohol-coated fluidic channel.

    - Shot of the channel halfplate with cured PVA without adhesive

  - Align the top and bottom polycarbonate housings using dowel pins [1].

    - Talent aligning the top and bottom polycarbonate housings with dowel pins.

  - Next, fill a 5-milliliter syringe with a 30-gauge needle with the chosen solvent, such as dichloromethane or acetone. For this video, the syringe has already been filled. Position the syringe tip at the lip or ledge created in the device design [2]. Carefully dispense the solvent along the lip or ledge on both sides of the assembly, allowing the solvent to wick into the bonding interface and reach the polyvinyl alcohol barrier [3].

    - Talent positioning the syringe tip at the lip or ledge of the assembly; Talent dispensing solvent along both sides of the lip or ledge so it wicks into the bonding interface.

    - Talent dispensing solvent along both sides of the lip or ledge so it wicks into the bonding interface.

  - Then, place a moderate weight of approximately 2 to 3 pounds evenly over the assembled device and allow the device to cure fully at room temperature for 10 to 15 minutes [1].

    - Talent placing a moderate weight evenly over the assembled device for curing and setting the sample aside.

  - Ensure no large bubbles are present to reduce points of failure [1].

    - Talent viewing the sample for bubbles.

  - After curing, flow 80 degrees Celsius deionized water through the fluidic channel at low flows of 1 to 5 milliliters per minute for a minimum of 3 hours [1]. Inspect the device visually for residual polyvinyl alcohol in the channel [2], and repeat the water flow until no polyvinyl alcohol is visually aparent [3].

    - Talent flowing deionized water through the fluidic channel.

    - Talent visually inspecting the channel for residual polyvinyl alcohol.

    - Talent turning on the water flow again.

- Low-Density Polyethylene (LDPE) Bonding of Polycarbonate Microfluidic Device

Demonstrators: Jackie Kading, Charlotte Li

  - Prepare the low-density polyethylene bonding sheets. For a 100 to 200 micrometer low-density polyethylene sheet, set the referenced desktop laser cutter to 75% power and 100% speed [1]. Laser-cut the low-density polyethylene sheets to match the shape of the channel wall regions [2].

    - Show the referenced desktop laser cutter settings set to 75% power and 100% speed

    - Talent operating the laser-cutting to prepare the shape of the channel wall regions. Show laser cutter operating.

  - Inspect the laser-cut low-density polyethylene sheets for ash, residue, or particulates [1]. Rinse the sheets thoroughly with isopropyl alcohol and allow them to dry completely [2].

    - Talent inspecting the laser-cut low-density polyethylene sheets for ash, residue, or particulates; Talent rinsing the sheets with isopropyl alcohol and leaving them to dry completely

    - .

  - Now, align the low-density polyethylene cutouts onto the side walls of the fluid channel using dowel pins [1]. Place the top polycarbonate piece onto the bottom assembly, sandwiching the low-density polyethylene layers at the bonding interface [2].

    - Talent aligning the low-density polyethylene cutouts onto the side walls of the fluid channel with dowel pins; Talent placing the top polycarbonate piece onto the bottom assembly to create the fluidic channel with the low-density polyethylene layers at the bonding interface.

    - Talent placing the top polycarbonate piece onto the bottom assembly to sandwich the low-density polyethylene layers at the bonding interface.

  - To bond the assembly, place the assembled housing into a preheated programmable oven set to 145 degrees Celsius with a temperature sensor [1]. Then, position a calibrated weight such as a metal plate on top of the assembly to apply moderate pressure of 2 to 3 pounds per square inch [2].

    - Talent placing the assembled housing into a programmable oven set to 145 degrees Celsius with a temperature sensor.

    - Talent placing a calibrated weight such as a metal plate on top of the assembly.

  - After 2 hours, remove the assembly from the oven [1]. Ensure even low-density polyethylene transparency across the entire bonded surface area  and allow the device to cool to room temperature under ambient conditions [2] .

    - Talent removing the assembly from the oven after 2 hours.

    - Shot of the bonded surface area for even low-density polyethylene transparency.

- Mechanical Gasket Sealing of Polycarbonate-Silicon Microfluidic Device

  - Cut the gasket cord material to the required length to match the fluidic channel path [1]. Place the gasket cord into the gasket gland on the bottom polycarbonate housing, ensuring a snug and continuous fit along the entire path [2]. Align the top and bottom housing halfplates to form the bounds of the fluidic channel [3].

    - Talent cutting the gasket cord material to the required length.

    - Talent placing the gasket cord into the gasket gland on the bottom polycarbonate housing.

    - Talent aligning the top and bottom housing halfplates to form the bounds of the fluidic channel.

  - Insert bolts through the top layer and into the hex nut, including washers or a compression plate as needed to evenly distribute pressure [1] and align the top polycarbonate piece with the bottom piece using bolts for alignment [2].

    - Talent inserting bolts through the top layer into the hex nut with washers or a compression plate; Talent aligning the plates in the process of inserting bolts

    - Talent aligning the plates in the process of inserting bolts

  - To tighten the bolts, hand-tighten each screw to hold the bolts and device in place [1]. Use a torque wrench to tighten each screw, moving in a cross pattern [2]. Tighten to 0.5 pound-inch to avoid cracking the silicon chip or deforming the polycarbonate housing, and tighten further as necessary [3].

    - Talent hand-tightening each screw to secure the device.

    - Talent using a torque wrench to tighten screws in a cross pattern.

    - Talent setting and applying 0.5 pound-inch torque on the screws.

- Pressure and Flow Characterization of the Hybrid Microfluidic Device

  - Prepare to connect the hybrid microfluidic device to a fluidic circuit [1]. Using peristaltic pump tubing, connect a water reservoir, roller pump, pressure sensor, and the hybrid microfluidic device in series [2]. Position the pressure sensor as close to the device inlet as possible to ensure accurate pressure measurements [2].

    - Talent arranging the hybrid microfluidic device and tubing for circuit connection.

    - Talent connecting the water reservoir, roller pump, pressure sensor, and device in series using peristaltic pump tubing; Talent positioning the pressure sensor close to the device inlet.

    - Talent positioning the pressure sensor close to the device inlet.

  - Pump water through the circuit at an initial flow rate of 10 milliliters per minute [1]. Increase the flow rate in increments of 10 milliliters per minute every minute [2]. Measure and record the pressure at each flow rate [2]. Add a needle restriction downstream of the device to increase circuit resistance if additional pressure is required [4-TXT].

    - Talent initiating water flow through the circuit at 10 milliliters per minute.

    - Talent adjusting the flow rate in increments of 10 milliliters per minute; Talent measuring and recording pressure readings at each flow rate

    - .

    - Talent adding a needle restriction downstream of the device. TXT: Define the rated pressure as the pressure at which the first visible leak occurs

  - Finally, insert a second pressure sensor immediately downstream of the device using the same circuit [1]. Flow water at a selected flow rate of 10 milliliters per minute [2] and calculate the pressure drop across the device as the difference between the pre-device and post-device pressure measurements [3].

    - Talent inserting a second pressure sensor downstream of the device.

    - Talent flowing water through the circuit at 10 milliliters per minute.

    - Talent working at the computer table.

# Results

- Results

  - The gasket-based mechanical sealing strategy was successfully used to obtain a prototype [1].

    - LAB MEDIA: Figure 4. Video editor: Show all panels in figure 4.

  - The gasket-based configuration sustained pressures up to 239 millimeters of mercury before leakage [1] and showed a low-pressure drop of 3 millimeters of mercury [2].

    - LAB MEDIA: Table 1. Video editor: Highlight the cell in the row “Rated Pressure (mmHg)” under the column “Gasket” showing 239 +or- 22.

    - LAB MEDIA: Table 1. Video editor: Highlight the cell in the row “Pressure Drop (mmHg) at 10 mL/min” under the column “Gasket”.