# Submission ID #: 69685

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21253973

# Title: Automated HPLC Separation Using LC-mate: An Integrated Repetitive Autosampler and Fraction Collector for Microscale Purification

# Authors and Affiliations:

# Chengyu Shi1,2, Weimin Zhao1,2*

# 1Natural Product Research Center, Shanghai Institute of Materia Medica, Chinese Academy of Sciences

# 2University of Chinese Academy of Sciences

# Corresponding Authors:

Weimin Zhao		(wmzhao@simm.ac.cn)

# Email Addresses for All Authors:

Chengyu Shi  		(s20020216@163.com)

Weimin Zhao		(wmzhao@simm.ac.cn)

## Author Questionnaire

- Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  No

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Yes

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=21253973

3. Filming location: Will the filming need to take place in multiple locations?   No

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials.  No

Current Protocol Length

Number of Steps: 22

Number of Shots: 42

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context.

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

1.1 Weimin Zhao: Our research focuses on designing an instrument that plugs right into HPLC setups to automate all those repetitive manual steps.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Weimin Zhao: Manual sample injection and fraction collection are very tedious. LC-mate offers a cross-platform, automated solution for HPLC sample purification.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Chengyu Shi: This system facilitates automated repetitive purification and precise fraction collection for efficient microscale HPLC workflows.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the most important consideration or challenge when performing this protocol.

  - Weimin Zhao: Correct signal wiring, proper mode selection and parameter setting are critical for successful application of the LC-mate instrument.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Weimin Zhao: We're planning to boost channel capacity and integrate new detectors for better cross-platform compatibility.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

- Integrated HPLC–LC-mate System Setup, Operation, and Automated Fraction Collection

Demonstrator: Chengyu Shi

  - To begin, verify the solvent plumbing connections between the HPLC system, the LC-mate, and the chromatographic column to ensure proper flow paths [1]. Check the connection of the signal transmission line between the ultraviolet detector and the LC-mate [2].

    - WIDE: Talent visually inspecting and tracing the solvent tubing connections between the high-performance liquid chromatography system, the LC-mate, and the chromatographic column.

    - Talent examining the signal transmission cable connected between the ultraviolet detector and the LC-mate.

  - Connect the black wire to the red terminal [1] and the white wire to the black terminal to match the polarity of the analog signal output [2]. Observe the LC-mate display to confirm that the analog signal waveform appears [3].

    - Talent connecting the black wire to the red terminal.

    - Talent connecting the white wire to the black terminal on the LC-mate.

    - SCREEN: 2.2.3ú¿English)(updated).MOV: 00:11-00:19

  - Reverse the wire connections by attaching the black wire to the black terminal and the white wire to the red terminal if no waveform is visible [1].

    - Talent switching the wire connections on the LC-mate terminals.

  - Now, check the trigger connection between the ultraviolet detector and the LC-mate [1]. Plug the single end of the data cable into the output port of the ultraviolet detector [2].

    - Talent inspecting the trigger connection ports on the ultraviolet detector and LC-mate.

    - Talent inserting the data cable into the ultraviolet detector output port.

  - Connect the black and white wires from the other end of the cable to the blue and black trigger terminals on the LC-mate, respectively [1]. Tighten the terminals to ensure a firm trigger connection [2].

    - Talent connecting the black and white wires to the corresponding trigger terminals on the LC-mate.

    - Talent tightening the trigger terminals on the LC-mate.

  - Next, ensure the proper startup of the HPLC system [1] and LC-mate [2].

    - Talent powering on and confirming the operational status of the high performance liquid chromatography system.

    - Talent powering on the LC-mate device.

  - Click on the LC-mate icon on the interface to automatically execute the system self-test program and initialize key hardware modules [1]. Wait for the brief prompt tone and the disappearance of the Self-Testing message to confirm that initialization is complete [2].

    - SCREEN: 2.7.1+2.7.2ú¿English)(updated).MOV: 00:00-00:03

    - SCREEN: 2.7.1+2.7.2ú¿English)(updated).MOV 00:04-00:06

  - Configure the HPLC separation conditions and begin elution to equilibrate the column [1].

    - Talent setting separation parameters on the high performance liquid chromatography system and starting elution.

  - Click on the Test icon on the left side of the LC-mate screen to open the test interface [1].

    - SCREEN: Show the LC-mate main interface and select the Test icon. NOTE: Not provided by the authors

  - In the test interface, input the Injection Volume, Elution Time, and Repetition Times [1], then press the Start icon to initiate the autosampler, fraction collector, and HPLC operation [2].

    - SCREEN: 2.10.1ú¿English)(updated).MOV: 00:00-end

    - SCREEN: 2.10.2ú¿English)(updated).MOV: 00:01-00:04

  - Identify feasible separation conditions, including resolution, running time, retention time, injection volume, and ultraviolet wavelength [1].

    - Talent observing chromatographic output and evaluating separation performance on the display.

  - Then select the Threshold function module on the left side of the interface to open the parameter setting screen [1]. Enter the Injection Volume, Repetition Times, Threshold, and Elution Time, and press Start to begin automated operation [2].

    - SCREEN: 2.12.1ú¿English)(updated).PNG

    - SCREEN: 2.12.2ú¿English)(updated).MOV: 00:00-00:16.

  - Monitor the status indicators. Observe Injection in Progress to confirm sample introduction [1]. Check the Separation in Progress indicator to verify that the chromatographic separation is going on [2]. Note Syringe Emptying indicator to identify when the syringe is purging residual liquid [3].

    - SCREEN: 2.13.1+2.13.2ú¿English)(updated).MOV: 00:00-end  Video editor: Highlight tab “Injection in Progress “

    - SCREEN: 2.13.1+2.13.2ú¿English)(updated).MOV: 00:00-end  Video editor: Highlight tab “seperation in“

    - SCREEN: 2.13.3ú¿English)(updated).MOV: 00:00-end

  - To perform fraction collection based on time, select the Timing Collection module from the left side of the interface to open the parameter menu [1]. Input Injection Volume, Repetition Times, Elution Time, and the required Collection Times [2]. Press the Start button to initiate the repeated automated sequence [3].

    - SCREEN: 2.14.1ú¿English)(updated).PNG

    - SCREEN: 2.14.2ú¿English)(updated).MOV: 00:18-00:36

    - SCREEN: 2.14.3ú¿English)(updated): 00:01-00:05-00:13-00:20

  - Next, select the Collect All function module on the left side of the interface [1]. Set a collection threshold value to collect all chromatographic peaks separately [2].

    - SCREEN: 2.15.1ú¿English)(updated).PNG

    - SCREEN: 2.15.2ú¿English)(updated).MOV: 00:05-end

  - For pipeline cleansing, prepare a small bottle of clean solvent [1]. Insert the injection pipeline of the autosampler into the solvent [2].

    - Shot of a labeled bottled of cleaning solvent.

    - Talent placing the autosampler injection pipeline into a bottle containing clean solvent.

  - Select the Cleansing module from the left side of the interface to open the parameter menu [1]. Then set the Injection Volume between 0 and 1000 and the Repetition Times for injector pipeline cleansing [2].

    - SCREEN: 2.17.1ú¿English)(updated).PNG

    - SCREEN: 2.17.2ú¿English)(updated).mov: 00:00-end

  - Click on Cleansing below Cleansing Completed to start injector pipeline cleansing [1]. Monitor the Cycle count to track completed wash cycles [2].

    - SCREEN: 2.18.1+2.18.2ú¿English)(updated).MOV:  00:20-00:26

    - SCREEN: 2.18.1+2.18.2ú¿English)(updated).MOV: 01:17-01:22 Video editor: highlight “3” present in the left side

  - Now, set the Cleansing Time for the collection pipeline cleansing [1]. Click on Cleansing below Current Pipeline to start collection pipeline cleansing [2].

    - SCREEN: 2.19.1+2.19.2ú¿English)(updated).MOV: 00:00-00:02

    - SCREEN: 2.19.1+2.19.2ú¿English)(updated).MOV:  00:02-end

  - Monitor the Current Pipeline display to confirm sequential cleaning from pipeline 0 to pipeline 10 [1].

    - SCREEN: 2.20.1ú¿English)(updated).MOV: 2.20.1.mp4: 00:02-00:20

  - Enable automatic saving to allow the system to save chromatograms after each elution cycle [1]. Alternatively, click on the Save button after elution to manually save data [2].

    - SCREEN: 2.21.1ú¿English)(updated).MOV: 00:00-end

    - SCREEN: 2.21.2ú¿English)(updated).pdf

  - After entering the metadata in the dialog box, choose the storage location from local storage, an SD card, or an external universal serial bus drive [1].

    - SCREEN: 2.21.3.mp4

# Results

- Results

  - Semi-preparative HPLC separation with LC-mate achieved continuous and stable sample injection and fraction collection during purification of 2”-O-galloyl-hyperoside [1]. Minor impurities were effectively removed to obtain purified 2”-O-galloyl-hyperoside using the LC-mate system [2].

    - LAB MEDIA: Figure 3. Video editor: Show the chromatogram panel

    - LAB MEDIA: Figure 3. Video editor: Zoom in on the main peak labelled as 2”-O-galloyl-hyperoside

  - Semi-preparative HPLC separation with LC-mate achieved continuous and stable sample injection and fraction collection during purification of phloridzin and trilobatin [1]. The LC-mate system achieved continuous and stable sample injection and fraction collection [2].

    - LAB MEDIA: Figure 4. Video editor: Show the chromatogram panel

    - LAB MEDIA: Figure 4. Video editor: Zoom in on the main peaks labeled phloridzin and trilobatin.

  - The relative standard deviation of the full ring injection peak area was 0.813 percent, indicating excellent stability and reproducibility of LC-mate injections [1].

    - LAB MEDIA: Table 2. Video editor: Highlight the row showing peak area values and the cell showing the relative standard deviation value of 0.813 percent.

- Chromatographic
Pronunciation link: https://www.merriam-webster.com/dictionary/chromatographic
IPA: /ˌkroʊ.mə.təˈɡræf.ɪk/
Phonetic Spelling: kroh·muh·tuh·graf·ik

- Ultraviolet
Pronunciation link: https://www.merriam-webster.com/dictionary/ultraviolet
IPA: /ˌʌl.trəˈvaɪ.ə.lət/
Phonetic Spelling: uhl·truh·vai·uh·luht

- Transmission
Pronunciation link: https://www.merriam-webster.com/dictionary/transmission
IPA: /trænzˈmɪʃ.ən/
Phonetic Spelling: tranz·mish·uhn

- Analog
Pronunciation link: https://www.merriam-webster.com/dictionary/analog
IPA: /ˈæn.əˌlɔːɡ/
Phonetic Spelling: an·uh·lawg

- Waveform
Pronunciation link: https://www.merriam-webster.com/dictionary/waveform
IPA: /ˈweɪvˌfɔːrm/
Phonetic Spelling: wave·form

- Polarity
Pronunciation link: https://www.merriam-webster.com/dictionary/polarity
IPA: /poʊˈler.ə.t̬i/
Phonetic Spelling: poh·lair·uh·tee

- Elution
Pronunciation link: https://www.merriam-webster.com/dictionary/elution
IPA: /iˈluː.ʃən/
Phonetic Spelling: ih·loo·shuhn

- Equilibrate
Pronunciation link: https://www.merriam-webster.com/dictionary/equilibrate
IPA: /ˌiː.kwəˈlɪb.ri.eɪt/
Phonetic Spelling: ee·kwuh·lib·ree·ayt

- Autosampler
Pronunciation link: https://www.merriam-webster.com/dictionary/autosampler
IPA: /ˌɔː.toʊˈsæm.plɚ/
Phonetic Spelling: aw·toh·sam·plur

- Chromatogram
Pronunciation link: https://www.merriam-webster.com/dictionary/chromatogram
IPA: /kroʊˈmæt̬.əˌɡræm/
Phonetic Spelling: kroh·mat·uh·gram

- Resolution
Pronunciation link: https://www.merriam-webster.com/dictionary/resolution
IPA: /ˌrez.əˈluː.ʃən/
Phonetic Spelling: rez·uh·loo·shuhn

- Retention
Pronunciation link: https://www.merriam-webster.com/dictionary/retention
IPA: /rɪˈten.ʃən/
Phonetic Spelling: rih·ten·shuhn

- Threshold
Pronunciation link: https://www.merriam-webster.com/dictionary/threshold
IPA: /ˈθreʃ.hoʊld/
Phonetic Spelling: thresh·hohld

- Syringe
Pronunciation link: https://www.merriam-webster.com/dictionary/syringe
IPA: /səˈrɪndʒ/
Phonetic Spelling: suh·rinj

- Purging
Pronunciation link: https://www.merriam-webster.com/dictionary/purging
IPA: /ˈpɝː.dʒɪŋ/
Phonetic Spelling: pur·jing

- Pipeline
Pronunciation link: https://www.merriam-webster.com/dictionary/pipeline
IPA: /ˈpaɪpˌlaɪn/
Phonetic Spelling: pipe·line

- Cleansing
Pronunciation link: https://www.merriam-webster.com/dictionary/cleansing
IPA: /ˈklen.zɪŋ/
Phonetic Spelling: klen·zing

- Metadata
Pronunciation link: https://www.merriam-webster.com/dictionary/metadata
IPA: /ˈmet̬.əˌdeɪ.t̬ə/
Phonetic Spelling: meh·tuh·day·tuh

- Semi-preparative
Pronunciation link: https://www.merriam-webster.com/dictionary/semi-preparative
IPA: /ˌsem.i prɪˈper.ə.t̬ɪv/
Phonetic Spelling: sem·ee·prih·pair·uh·tiv

- Hyperoside
Pronunciation link: No confirmed link found
IPA: /haɪˈpɪr.oʊ.saɪd/
Phonetic Spelling: hy·peer·oh·syde

- Phloridzin
Pronunciation link: No confirmed link found
IPA: /ˈflɔːr.ɪd.zɪn/
Phonetic Spelling: flor·id·zin

- Trilobatin
Pronunciation link: No confirmed link found
IPA: /traɪˈloʊ.bə.tɪn/
Phonetic Spelling: try·loh·buh·tin

- Reproducibility
Pronunciation link: https://www.merriam-webster.com/dictionary/reproducibility
IPA: /ˌriː.prəˌduː.səˈbɪl.ə.t̬i/
Phonetic Spelling: ree·pruh·doo·suh·bil·uh·tee

- Deviation
Pronunciation link: https://www.merriam-webster.com/dictionary/deviation
IPA: /ˌdiː.viˈeɪ.ʃən/
Phonetic Spelling: dee·vee·ay·shuhn

- 2”-O-galloyl-hyperoside
Pronunciation link: No confirmed link found
IPA: /tuː oʊ ˈɡæl.ɔɪl haɪˈpɪr.oʊ.saɪd/
Phonetic Spelling: too·oh·gal·oyl·hy·peer·oh·syde