TITLE:

Laser Capture Microdissection of Paraformaldehyde-Fixed Mouse Liver Tissue for RNA Analysis

AUTHORS AND AFFILIATIONS:

Wei-Lan Zeng1,2,#,*, Ming-Ming Zhang1,3,#, Wei-Jie Zeng4, Yan Wang1,2,3,5,*

1 Biomedical Research Center, Southern Medical University, Guangzhou, Guangdong 510515, China

2 School of Traditional Chinese Medicine, Southern Medical University, Guangzhou, Guangdong 510515, China

3 School of Pharmaceutical Science, Southern Medical University, Guangzhou, Guangdong 510515, China

4 Leica Microsystems (Shanghai) Trading Co., Ltd., Shanghai, 201206, China

5 The Seventh Affiliated Hospital, Southern Medical University, Foshan, Guangdong 528244, China.

#  These authors contributed equally.

Email addresses of the co-authors:

Ming-Ming Zhang	mingming96@smu.edu.cn

Wei-Jie Zeng		weijie.zeng@leica-microsystems.com

*Co-corresponding authors

Yan Wang		Yanwang@smu.edu.cn

Wei-Lan Zeng		weilanlan21810500@smu.edu.cn

Keywords:

Laser capture microdissection, paraformaldehyde fixation, RNA isolation, gene expression, liver tissue

SUMMARY:

This protocol describes a method to isolate RNA from defined histological regions of paraformaldehyde-fixed, Optimal Cutting Temperature compound-embedded mouse liver tissue using laser capture microdissection, enabling targeted downstream gene expression analysis.

ABSTRACT:

Laser capture microdissection (LCM) provides spatial access to specific cell populations within complex tissues through in situ visualization and isolation. To enable transcriptomic analysis of histologically defined regions in fixed tissue, a detailed LCM protocol is presented for RNA extraction from paraformaldehyde (PFA)-fixed, Optimal Cutting Temperature (OCT) compound-embedded mouse liver sections. The protocol details the identification and collection of microscale samples (approximately 1,000 cells) through a workflow encompassing tissue fixation, sucrose dehydration, OCT embedding, cryosectioning, hematoxylin staining, and laser-capture of targeted histological areas. Using this method, a high-purity RNA (A260/A280: 1.9–2.1) was obtained. The RNA integrity number (RIN) was 6.7 ± 0.9, reflecting the expected fragmentation associated with PFA fixation. However, quantitative PCR for β-actin yielded Ct values of 17–19, and RNA sequencing performed using fragmentation-optimized library preparation generated high-quality reads, with >90% of bases meeting Q20 and Q30 thresholds, confirming that the RNA is suitable for sensitive downstream analyses. Therefore, this protocol enables spatially resolved, targeted gene expression analysis by providing RNA of defined purity and integrity from specific histological regions of PFA-fixed liver tissue.

INTRODUCTION:

Accurate acquisition of spatially resolved gene expression profiles is crucial for understanding tissue heterogeneity in health and disease. Single-cell RNA sequencing (scRNA-seq) offers high cellular resolution1,2, but it loses native spatial context and may alter transcriptional profiles during tissue dissociation3,4. For cell types lacking publicly identified specific markers, obtaining pure populations using immunofluorescence-based sorting or immunomagnetic isolation remains difficult5,6. Spatial transcriptomics technologies such as Visium and GeoMx allow in situ capture of high-throughput gene expression data, but their spatial resolution is limited at the single-cell or subcellular level7,8. Conversely, imaging-based spatial transcriptomics can achieve nanometer-level resolution but has limited gene detection throughput9. This highlights a persistent gap for hypothesis-driven analysis of specific morphological regions.

Laser capture microdissection (LCM) allows microscope-guided selection and isolation of specific cells from tissue sections or live cell cultures based on observable phenotype. This approach preserves cellular structure and spatial information10. This technique has been extensively applied in liver disease research for downstream analyses. For instance, LCM has been used to collect biliary epithelial cells from liver biopsies to elucidate their immunoregulatory role in primary biliary cholangitis11. In addition, it has been used to isolate γ-glutamyl transferase (GGT)-positive tissue from experimental models to confirm site-specific Ggt1 gene expression12. But these studies have largely focused on fresh frozen tissues, which often lack clinical annotations or long-term follow-up13. In contrast, chemically fixed tissues offer advantages for histopathological diagnosis14. Although LCM-based RNA sequencing (RNA-Seq) has been successfully applied to archived, large-scale formalin-fixed and paraffin-embedded (FFPE) liver specimens15. However, a robust LCM-based gene expression profiling protocol specifically for microscale samples from paraformaldehyde (PFA)-fixed and Optimal Cutting Temperature (OCT) compound-embedded liver tissues is lacking.

To address this unmet need, a comprehensive protocol was optimized using PFA-fixed, OCT-embedded cryosections. This protocol addresses the limitations of FFPE workflows (e.g., prolonged processing and harsh chemicals) for sensitive LCM-RNA applications16. In contrast, the fixation and embedding process for PFA-fixed, OCT-embedded samples is relatively simpler and more flexible, enabling rapid preparation of high-quality sections that preserve both RNA integrity and tissue morphology. Furthermore, this protocol is performed using an ultraviolet (UV) laser-based LCM system. Compared to infrared systems, UV-LCM provides finer cutting resolution and enables direct, non-contact sample collection, minimizing contamination and making it suited for the precise isolation of microscale regions from complex tissue architectures17,18. Therefore, the operational value of this protocol lies in its ability to perform morphology-preserving capture of microscale regions (~1,000 cells) from PFA-fixed liver sections. It yields RNA suitable for targeted gene expression analysis, allowing direct correlation between histopathological features and local transcriptomic profiles.

Here, we describe a protocol for laser capture microdissection of microscale tissue (~1,000 cells) from PFA-fixed, OCT-embedded mouse liver sections, followed by RNA extraction and downstream analysis by quantitative reverse transcription PCR and RNA sequencing. These procedures can be completed within five days.

To guide readers in assessing the applicability of this protocol, the key scope and limitations are

summarized below. This protocol is specifically optimized for hypothesis-driven, morphology-guided sampling of specific histological regions from PFA-fixed, OCT-embedded tissues, with a primary application in mouse liver. It is suited for researchers who need to correlate spatially resolved transcriptomic data with precise histopathological features preserved in fixed specimens. The workflow is particularly valuable when working with archived or clinically annotated fixed tissues where fresh-frozen samples are unavailable. However, this protocol is not ideal for studies requiring high-integrity, non-fragmented RNA (e.g., RNA integrity number (RIN) > 8), as the inherent fragmentation from chemical fixation limits its use in applications that depend on full-length transcripts.

PROTOCOL:

All procedures for animal handling and care were performed in accordance with ethical guidelines and approved by the Animal Ethics Committee at Southern Medical University. The described method used adult male C57BL/6J mice and was maintained on a high-fat, high-sugar, choline-deficient diet to model metabolic-associated steatohepatitis (MASH). Before harvesting liver tissue, the animals were euthanized by decapitation.

- Preparation of PFA-fixed specimen

  - Fix fresh liver tissue samples (individual lobes, 5 mm thickness or less) in 30 mL of cold 4% PFA (freshly prepared) at 4 °C for 24 h.

NOTE: It is recommended to perfuse the liver with phosphate-buffered saline (PBS, 1×, typically composed of 137 mM NaCl, 2.7 mM KCl, 10 mM Na₂HPO₄, and 1.8 mM KH₂PO₄, pH 7.4) solution at 1× concentration for 5 min, followed by perfusion fixation with 4% PFA for 5 min. Alternatively, the tissue may also be fixed in 10% formalin solution.

CAUTION: PFA fixative is toxic. All PFA waste must be collected in a designated, properly labeled chemical waste container in compliance with institutional and local regulations.

  - Wash the fixed tissue samples 3× times for 5 min with RNase-free 1× PBS solution.

  - Place the fixed liver tissue into a 50 mL tube filled with 50 mL of 30% sucrose (freshly prepared in sterile 1× PBS). Maintain the samples at 4 °C for 10–12 h.

  - Remove the tissue samples from the sucrose solution. Trim the samples to obtain easy-to-handle blocks (approximately 5 mm × 3 mm × 5 mm), then embed the samples into a cryomold (7 mm × 7 mm × 5 mm) containing OCT compound.

NOTE: Adjust the tissue orientation with forceps if necessary.

  - Add OCT compound on top to fully cover the tissue samples, then immediately place the OCT-embedded tissue block on liquid nitrogen or dry ice to rapidly freeze the OCT compound and form a solid block. This rapid freezing is essential for obtaining high-quality cryosections.

Caution: Wear cryogenic-resistant gloves to protect hands when handling cryogenic materials.

  - Place the OCT-embedded blocks at -80 °C until cryosectioning.

- Cryosectioning for laser capture microdissection

  - Expose the polyethylene naphthalate (PEN) membrane slides to UV light for 30 min, then coat the slides with 0.1 mg/mL of poly-L-lysine, followed by air-drying. Then, seal the edges of the PEN membrane slide with a neutral mounting medium.

NOTE: UV irradiation and poly-L-lysine treatment of PEN membrane slides can improve the adherence of tissue samples to the PEN membrane slides. Sealing the perimeter with a neutral mounting medium prevents water infiltration or bubble formation beneath the membrane, which is critical for maintaining LCM efficiency.

CAUTION: Protect the skin and eyes from the UV light. Wear UV-blocking safety goggles or a face shield (rated for UV protection), as well as a lab coat, nitrile gloves, and long sleeves to cover exposed skin.

  - Wipe all accessible surfaces of the cryostat and all tools with an RNase decontamination agent, and clean the brush with 75% ethanol.

CAUTION: Dispose of used ethanol wipes and solutions in accordance with local institutional environmental health and safety protocols.

  - Allow the instrument to reach and stabilize at −20 °C.

  - Retrieve the liver samples from -80 °C and place them into the cryostat for a minimum of 15 min.

  - Apply a thin layer (2–3 mm) of OCT compound onto the surface of the specimen stage. Place the specimen stage on a quick freeze shelf (or in the cryostat chamber) to pre-solidify the OCT layer.

  - Apply a small amount of OCT onto the pre-solidified OCT layer on the specimen stage. Retrieve the equilibrated tissue block (from step 2.4) and immediately place the block horizontally onto this OCT, hold it in place for 10–15 s to allow partial freezing and adhesion. Once secured, mount the specimen stage (with the attached tissue block) onto the cryostat specimen head.

  - Carefully insert the blade between the clamping plates and the back plates from one side of the cryostat, and then clamp the clamping lever.

  - Adjust the cryostat’s thickness setting to 10 µm.

NOTE: Adjust slice thickness as needed for the tissue, but maintain the thickness between 5–15 µm19.

  - Cut and mount the sections directly onto the pretreated PEN membrane slides inside the cryostat chamber maintained at -20 °C.

NOTE: Each PEN membrane slide can accommodate 6–8 tissue sections. It is recommended that the entire sectioning process be completed within 1 h. Limiting sectioning time minimizes tissue exposure to ambient conditions, thereby reducing the risks of RNA degradation, tissue dehydration, and morphological artifacts, which are critical for preserving RNA quality and histological integrity for LCM and downstream analysis.

- Hematoxylin staining for laser capture microdissection

  - Thaw the slides at room temperature (RT) for 20 min before use.

  - Wash the slides with RNase-free 1× PBS (3 times, 10 s each).

  - Immerse the tissue slides in filtered Mayer's hematoxylin solution (composed of 1.0 g/L hematoxylin, 0.2 g/L sodium iodate, 50 g/L aluminum ammonium sulfate dodecahydrate, 50 g/L chloral hydrate, and 1.0 g/L citric acid) for 5–10 s.

NOTE: Pass the hematoxylin working solution through a 0.22 µm syringe filter membrane before staining.

  - Wash the slides with RNase-free water (3 times, 10 s each) to remove residual hematoxylin solution.

  - Assess staining quality under a microscope.

NOTE: Optimally stained nuclei should appear crisp blue-purple against a clear background. Insufficient staining yields faint nuclei, while over-staining obscures cellular details. Only slides meeting these optimal staining criteria should proceed to the next step.

  - Blot the slides dry using absorbent paper and then air-dry completely (10–15 min).

NOTE: Ensure that the slides are fully dried before proceeding with LCM. Then place the slides into the 50 mL centrifuge tubes and store them temporarily on ice until LCM is performed.

- Laser capture microdissection

NOTE: The duration of the LCM experiment for each slide should be maintained between 30 min and 1 h. It is recommended that the entire LCM experiment be completed within 4 h. This limit is set to minimize RNA degradation at RT and to prevent the evaporation of the collection buffer in the tube cap, which is crucial for efficient sample retrieval. In practice, the target yield of approximately 1,000 cells from morphologically defined regions can be reliably harvested within 4 h.

  - Wipe the surfaces of the laboratory bench using the RNase decontamination agent.

  - Turn on the controller and allow the microscope to complete its initialization process.

  - Rotate the key clockwise from the Off position to the On position to initiate the laser system.

NOTE: The laser system requires about 10 min to warm up.

  - Turn on the computer and open the LCM software, then wait for the initialization process to complete.

  - Insert a sterile 0.2 mL PCR tube (collection tube) into the collection device.

  - Add 10 µL of a proteinase K digestion buffer (see the Table of Materials) to the cap of the PCR tube.

NOTE: The total volume of proteinase K digestion buffer will be brought to 150 µL in a subsequent step (see Step 5.4.1) as part of the complete lysis procedure.

  - Load the collection device into the microscope’s specimen holder.

  - Within the Change Collector Device window, first select Move to Reference Point button. Next, adjust the camera focus to ensure a clear view of the reference point. If necessary, use the arrows in the window to center the reference point in the field of view.

  - Load a tissue section into the specimen holder, ensuring the tissue section is oriented face down, and the label is to the right.

  - Select the 10× objective lens, identify and outline regions of interest exhibiting specific pathological features, such as microvesicular steatosis (characterized by hepatocytes with numerous small, clear cytoplasmic vacuoles and centrally located nuclei). Then, navigate to the Laser menu and click the Calibrate button. Confirm the on-screen message by selecting Yes and proceed with calibration.

NOTE: The 10× objective lens is not exclusive for microdissection; other magnification lenses (e.g., 20×) can be selected based on sample clarity and the desired resolution.

  - In the laser control panel, set the cutting parameters to: Laser Power = 50, Aperture (laser beam diameter) = 5, and Speed = 9.

NOTE: The above-listed parameter values serve as a reference for the UV-LCM system with the tissue samples used in this protocol. Users of other UV-LCM systems should consider these as a starting point and will likely need to adjust settings empirically. Optimal settings may vary with tissue type, section thickness, and other LCM systems.

  - Draw a circle around the target area using the graphic toolbar. Select the Single Shape, click the Start Cut button to perform LCM, and then position the collection device to collect the dissected tissue.

NOTE: Successful detachment is indicated by a separation of the selected area, often with a visible gap or fragment movement. For consistency, collect from a pre-defined number of regions of similar size per sample. If the target area does not detach automatically, click the Move and Cut button to cut the specimen manually using the mouse. For a more precise estimate of cell number, LCM software can automatically record the total cut area. This value can be divided by the average cross-sectional area of a target cell to approximate the total number of cells captured.

CAUTION: During the LCM process, do not stare directly at the laser beam to avoid laser radiation exposure.

  - Optionally, capture a Before and After image of the dissection site by selecting the Save Image As command in the File menu. In the dialog box, choose a location, filename, and save in a preferred format (e.g., .tif).

  - Upon collecting the desired number of cells, click the Unload button to remove the collector from the stage. Before sealing the PCR tube, verify that the proteinase K digestion buffer volume in the cap remains at 10 µL; replenish the buffer if evaporation occurred during LCM (see Step 4.6).

NOTE: Adding 10 µL of the buffer ensures that the cap surface is fully wetted, which facilitates the efficient transfer of the micro-dissected tissue pellet during centrifugation in Step 4.16.

  - Vortex the collection tubes thoroughly.

  - Spin the collection tubes briefly and store the sample tubes at -80 °C until RNA extraction.

NOTE: Repeat Steps 4.5 to 4.16 for the next sample. The protocol can be paused until RNA extraction.

- Total RNA extraction of micro-dissected samples

  - Wipe the surfaces of the laboratory bench using the RNase decontamination agent.

  - Transfer the LCM samples from the -80 °C freezer to an ice bath. If desired, pool different batches of the same captured samples.

  - Incubate the micro-dissected sample tubes upside down at 42 °C for 30 min.

NOTE: This inversion ensures that any tissue remaining adhered to the cap after brief centrifugation (Step 4.16) is released into the digestion buffer. This ensures maximal sample transfer into the buffer prior to the volume adjustment and high-speed centrifugation in Step 5.4.1.

  - Isolate total RNA from LCM samples using a standard RNA extraction kit optimized for fixed tissues (see the Table of Materials).

    - Adjust the volume in each sample tube to a final 150 µL with the proteinase K digestion buffer (see the Table of Materials). Spin the tubes at 11,000 × g (1 min), then transfer the samples to a sterile 1.5 mL tube.

    - Add 10 µL of proteinase K (20 mg/mL) per tube, then gently mix the samples through repeated pipetting. This results in a final volume of 160 µL and a final proteinase K concentration of approximately 1.25 mg/mL in the digestion reaction per sample.

    - Perform a two-step sample incubation: first at 56 °C for 15 min, followed by 15 min at 80 °C. During each incubation, briefly vortex the tubes every 5 min.

NOTE: For optimal workflow, use two pre-heated water baths. If using a single bath, pause after the 56 °C incubation: hold the samples at RT while the bath heats to 80 °C, then proceed.

    - Place the samples for 3 min on ice, followed by centrifugation at 20,000 × g (15 min).

    - Aspirate the clarified supernatant to a new 1.5 mL tube, taking care to avoid the pellet.

    - Add 16 µL of DNase incubation buffer and 10 µL of DNase I enzyme (2.73 units/µL, see the Table of Materials) to each sample. Then, mix the contents thoroughly by gently inverting the tubes. This results in a final volume of 186 µL and a final DNase I concentration of approximately 0.15 units/µL in the digestion reaction per sample.

    - Perform a quick spin to collect residual liquid from the tube walls.

    - Incubate the sample tubes at RT for 15 min.

    - Transfer 320 µL of high-salt wash buffer (see the Table of Materials) to each sample and mix the samples thoroughly. This results in a final volume of 506 µL per sample.

    - Transfer 720 µL of absolute ethanol to each tube, mix the samples by pipetting (no centrifugation), and proceed immediately. This results in a final volume of 1,226 µL per sample.

CAUTION: Absolute ethanol is highly flammable. Collect waste ethanol in a designated non-halogenated organic solvent waste container for proper disposal.

    - Transfer 700 µL of the sample mixture to a purification column (placed in a 2 mL collection tube, see the Table of Materials). Gently close the lid, and centrifuge at 8,000 × g (15 s).

    - Discard the flowthrough of the collection tube. Retain the purification column.

    - Repeat Steps 5.4.11 to 5.4.12.

    - Add 500 µL of ethanol-based wash buffer (see the Table of Materials) to each sample tube. Gently close the lids, and centrifuge at 8,000 × g (15 s). Then, discard the flow-through of the collection tube.

    - Repeat Step 5.4.14 in its entirety, with the sole modification of increasing the centrifugation time to 2 min at 8,000 × g.

    - Transfer each purification column to a new 2.0 mL tube. Centrifuge at 14,000 × g (5 min) with the cap open, then discard the tube and residual liquid.

    - Transfer each purification column to a new 1.5 mL tube. Pipette 16 µL of RNase-free water onto the membrane, close the cap, and centrifuge at 14,000 × g (5 min) to facilitate the elution of the total RNA.

NOTE: Keep all eluted RNA samples on ice for Steps 5.5 and 5.6 to prevent degradation.

  - Quantify the RNA samples and check the purity (A260/A280 ratio) by spectrophotometry.

NOTE: The measurement of concentration and purity (A260/A280 ratio) in this step, along with the integrity assessment in Step 5.6, serves as the definitive quantitative checkpoint for successful RNA recovery. A clear RNA eluate is expected.

  - Perform agarose gel electrophoresis with a 1% gel concentration and determine the RNA integrity number (RIN) using a Bioanalyzer system to evaluate RNA quality.

  - Store the RNA samples at -80 °C until performing qRT-PCR or RNA-Seq.

- RNA analysis

NOTE: qRT-PCR is described in Steps 6.1 and 6.2, and RNA-Seq is described in Step 6.3.

  - Prepare the reverse transcription reaction mixture and add 50 ng of total RNA per reaction on ice. Proceed with the reverse transcription reaction using a thermal cycler according to the instructions provided (see the Table of Materials).

  - On ice, prepare quantitative PCR master mix for β-actin according to the manufacturer’s instructions (see the Table of Materials for primer sequences). Perform real-time PCR amplification on a quantitative PCR system using the manufacturer’s recommended cycling protocol with 10 ng of cDNA per reaction.

  - Using a low-input stranded kit (see the Table of Materials), construct RNA-Seq libraries from 50 ng of total RNA. Determine the library size and concentration with a bioanalyzer system (see the Table of Materials) and then perform paired-end sequencing on a high-throughput platform (see the Table of Materials).

RESULTS:

The experimental workflow, which includes PFA fixation, 30% sucrose dehydration, OCT embedding and cryosectioning, hematoxylin staining, identification and precise collection of the target region (~1,000 cells) by LCM, cell lysis, RNA isolation, RNA quality control metrics and the two key analytical applications, is outlined in Figure 1. The hematoxylin-stained liver section on the PEN slide is shown under the microscope before (Figure 2A) and after (Figure 2B) laser capture of microscale liver tissue. Regions of interest exhibiting microvesicular steatosis were precisely collected into the cap of a PCR tube, preserving intact cellular morphology (Figure 2C).

Total RNA was extracted from approximately 1,000 and 2,000 LCM-captured cells per replicate, yielding 110.7 ± 11.8 ng and 213.3 ± 26.2 ng, respectively (Figure 3A). All samples showed high purity, with A260/A280 ratios between 1.9 and 2.1 (Figure 3B), characteristic of pure RNA20. Gel electrophoresis revealed the presence of distinct 28S and 18S ribosomal RNA bands (Figure 3C). Bioanalyzer profiles of RNA from as few as 1,000 cells gave the RIN of 6.7 ± 0.9 (Figure 3D,E), which aligns with expected values for PFA-fixed samples21, 22.

The utility of RNA isolated from approximately 1,000 LCM-captured cells was validated through downstream molecular assays (Figure 4). qRT-PCR for the housekeeping gene β-actin yielded cycle threshold (Ct) values ranging from 17 to 19 (Figure 4A), suggesting the functional integrity of cDNA derived from this limited fixed-tissue input. Next, sequencing libraries were constructed from 50 ng of total RNA. Bioanalyzer assessment of these libraries showed a characteristic peak between 300 and 500 bp (Figure 4B). Consistent with the expected size distribution for sequencing libraries constructed from fragmented RNA, RNA sequencing of these libraries yielded high-quality data, with both Q20 (Phred quality score > 20) and Q30 (Phred quality score > 30) scores exceeding 90%, and a low error rate across replicates (Table 1). This confirms that RNA obtained from this LCM workflow is competent for the construction of sequencing-ready libraries.

FIGURE AND TABLE LEGENDS:

Figure 1: Workflow for laser capture microdissection (LCM)-based RNA analysis. Paraformaldehyde-fixed mouse liver tissue was cryosectioned, followed by brief hematoxylin staining for morphological guidance. Target regions were isolated by LCM. RNA was then extracted from the captured cells for downstream gene expression analysis via quantitative PCR and RNA sequencing.

Figure 2: LCM-captured microscale samples in liver section. (A, B) Representative hematoxylin-stained liver sections from fixed mouse liver tissue before (A) and after (B) LCM. The regions targeted for microdissection (outlined in black) were selected based on the presence of microvesicular steatosis, characterized by numerous small, clear cytoplasmic vacuoles that give the hepatocytes a foamy appearance while the nucleus remains centrally located. (C) The captured LCM samples are visualized on the tube cap. Scale bar = 100 µm.

Figure 3: Quality assessment of RNA isolated from LCM-captured microscale samples. (A) Yield and (B) purity (A260/A280 ratio) of total RNA extracted from LCM samples containing approximately 2,000 and 1,000 cells. (C) Gel electrophoresis showing the integrity of RNA extracted from 2,000 (200 ng) and 1,000 cells (100 ng), and liver tissue without LCM (100 ng). 18S and 28S ribosomal RNA bands were indicated. (D, E) Quality assessment of RNA from 1,000 cells. (D) A representative electrophoretic profile (RNA integrity numbers (RIN) = 6.80). (E) RIN values across biological replicates. Data are presented as mean ± standard deviation. N = 3 samples per group.

Figure 4: Downstream validation of RNA from LCM-captured samples. (A) Amplification plot for the housekeeping gene β-actin using cDNA reverse-transcribed from total RNA isolated from approximately 1,000 LCM-captured cells. Ct values ranged from 17 to 19. (B) Size distribution of cDNA libraries constructed from 50 ng of total RNA derived from LCM-captured microscale samples (approximately 1,000 cells), as determined by Bioanalyzer. The predominant peak within the 300–500 bp range indicates successful library preparation suitable for sequencing. Abbreviations: Ct = Cycle threshold

Table 1. Table 1. RNA sequencing quality metrics for laser capture microdissection samples.
Summary of clean reads, error rate, Q20, Q30, and GC content for RNA sequencing libraries prepared from microscale samples of approximately 1,000 laser capture microdissection–isolated cells (N = 3).

DISCUSSION:

The strength of LCM lies in its precise, in situ capture of selected cells or tissue regions, facilitating the study of specific areas within a heterogeneous tissue. But its low sample yield, time-consuming process, high cost, and the need for specialized equipment limit the widespread application of this technique in laboratories23. In this study, the LCM-RNA protocol was established to overcome key historical limitations. This protocol implemented a rapid hematoxylin staining step (~10 s) and, by utilizing PFA/OCT embedding instead of FFPE, eliminated the time-consuming xylene deparaffinization steps required in many fixed-tissue protocols. The single-tube collection strategy avoids intermediate transfers, reducing sample loss and pipetting steps. Sealing the membrane perimeter with mounting medium prevents common hydration issues, increasing the first-attempt success rate and cutting efficiency22. Efficient recovery of cross-linked RNA was ensured through the use of a PFA-adapted RNA extraction kit. When coupled with an ultra-sensitive library preparation kit, this approach enabled successful RNA extraction from the picogram amounts of RNA yielded. Furthermore, tissue section thickness was set to a range of 5–15 µm. As supported by prior methodological studies10,19, sections thinner than 5 µm may not provide a complete cellular profile for reliable capture, while sections thicker than 15 µm can impede laser microdissection efficiency. These optimizations can reduce processing time and sample loss, making reliable RNA extraction and analysis from specific morphological cell types in PFA-fixed liver tissues feasible.

Obtaining high-quality RNA for downstream analysis critically depends on optimal sample preparation and a well-designed LCM workflow. Fresh frozen tissues allow for rapid sample processing, but they have a short storage time, and the tissues are prone to ice crystal formation during freezing, which can disrupt cell integrity. Conversely, formalin and PFA are widely used chemical fixatives in histology, not only for long-term tissue preservation but also for better retention of tissue morphology10,24. However, PFA-fixed samples exhibit the cross-link formation between nucleic acids and proteins. This process inherently yields partially fragmented RNA, typically with RIN values in the range of 5–7; this expected outcome should not be interpreted as an extraction failure16,21,22,25. Consequently, this RNA is suitable for applications such as quantitative PCR and RNA sequencing with fragmentation-compatible libraries, but it is not ideal for techniques that require long, intact transcripts. Previous studies have enhanced RNA quality by developing and optimizing sample collection and processing procedures for LCM workflows26,27. The combined use of proteinase K and heat effectively reverses formaldehyde cross-links, enabling the extraction of RNA from fixed samples28,29. Subsequent strategies like rRNA depletion or transcript amplification maximize the recovery of genetic information from the often limited and fragmented RNA, ensuring its suitability for downstream analyses30,31. Minimizing the post-collection interval before fixation is a critical, time-sensitive step to prevent tissue autolysis and nuclease-driven nucleic acid degradation. Fixation times exceeding 24 h may lead to more irreversible formaldehyde cross-linking24,31. Recent studies have successfully coupled immuno-LCM with RNA-Seq using a combination of perfusion and immersion fixation to achieve optimal RNA preservation for brain tissue research21. This protocol aims to address several practical challenges in liver histopathology. First, while the combination of perfusion and immersion fixation is ideal for some tissues (e.g., brain), it is less feasible for certain disease models or clinical samples. Therefore, the present approach utilizes immersion-fixed, OCT-embedded sections. This design offers a workflow that is compatible with standard pathological processing. Second, hematoxylin staining for cell selection avoids potential issues associated with antibodies, such as epitope masking, staining variability, and the need for prior knowledge of specific markers. This allows for the isolation of cell populations based on discernible histopathological features.

While this protocol provides a reliable approach, attention to common technical challenges and their solutions is critical for ensuring experimental success. For instance, if hematoxylin staining yields insufficient contrast, this is often due to suboptimal staining time or aged dye; optimal results can be achieved by titrating the staining duration between 5–10 s and using freshly filtered hematoxylin. A frequent concern during LCM is that the dissected tissue does not fall into the cap. This issue can typically be resolved by verifying laser calibration, ensuring the PEN membrane is fully dry and sealed, and confirming that the collection buffer has not evaporated. Low RNA yield or degradation often stems from RNase contamination or over-fixation; rigorous use of RNase-free techniques and adherence to the specified fixation and digestion times are critical.

This experimental protocol provides a reliable method to obtain high-quality RNA from morphologically defined, microscale regions of PFA-fixed liver tissue. The RNA produced is competent for qPCR and RNA-Seq library construction, thereby establishing a crucial methodological foundation for future investigations into cell type-specific transcriptional alterations in health and disease. Additionally, this protocol serves as a valuable reference for RNA extraction and analysis from fixed-tissue samples in other organs.

ACKNOWLEDGMENTS:

This work was supported in part by grants from the National Natural Science Foundation of China (82070589), the Guangdong Natural Science Foundation (2022A1515010218), and the Guangzhou Science and Technology Plan Project (206077078001) to Prof. Yan Wang. We are grateful to the Biomedical Research Center at Southern Medical University for technical support and access to the laser capture microdissection system.

DISCLOSURES:

The authors have no conflicts of interest to declare.

REFERENCES:

1.	Shapiro, E., Biezuner, T., Linnarsson, S. Single-cell sequencing-based technologies will revolutionize whole-organism science. Nat Rev Genet. 14 (9), 618-630 (2013).

2.	Halpern, K. B. et al. Single-cell spatial reconstruction reveals global division of labour in the mammalian liver. Nature. 543 (7647), 352-356 (2017).

3.	Aizarani, N. et al. A human liver cell atlas reveals heterogeneity and epithelial progenitors. Nature. 572 (7768), 199-204 (2019).

4.	Van Den Brink, S. C. et al. Single-cell sequencing reveals dissociation-induced gene expression in tissue subpopulations. Nat Methods. 14 (10), 935-936 (2017).

5.	Baron, C. S. et al. Cell Type Purification by Single-Cell Transcriptome-Trained Sorting. Cell. 179 (2), 527-542.e519 (2019).

6.	Gross, A. et al. Technologies for Single-Cell Isolation. Int J Mol Sci. 16 (8), 16897-16919 (2015).

7.	Dries, R. et al. Advances in spatial transcriptomic data analysis. Genome Res. 31 (10), 1706-1718 (2021).

8.	Chen, A. et al. Spatiotemporal transcriptomic atlas of mouse organogenesis using DNA nanoball-patterned arrays. Cell. 185 (10), 1777-1792.e1721 (2022).

9.	Ståhl, P. L. et al. Visualization and analysis of gene expression in tissue sections by spatial transcriptomics. Science. 353 (6294), 78-82 (2016).

10.	Espina, V. et al. Laser-capture microdissection. Nat Protoc. 1 (2), 586-603 (2006).

11.	Baba, N. et al. Gene expression profiling in biliary epithelial cells of primary biliary cirrhosis using laser capture microdissection and cDNA microarray. Transl Res. 148 (3), 103-113 (2006).

12.	Mena, J. E. T. et al. Laser capture microdissection after γ-glutamyl transferase histochemistry: An optimization for gene expression analysis. Anal Biochem. 447 126-132 (2014).

13.	Liu, Y. et al. Next-generation RNA Sequencing of Archival Formalin-fixed Paraffin-embedded Urothelial Bladder Cancer. Eur Urol. 66 (6), 982-986 (2014).

14.	Ludyga, N. et al. Nucleic acids from long-term preserved FFPE tissues are suitable for downstream analyses. Virchows Arch. 460 (2), 131-140 (2012).

15.	Mägel, L., Bartels, S., Lehmann, U. Next-Generation Sequencing Analysis of Laser-Microdissected Formalin-Fixed and Paraffin-Embedded (FFPE) Tissue Specimens. Methods Mol Biol. 1723 111-118 (2018).

16.	Von Ahlfen, S., Missel, A., Bendrat, K., Schlumpberger, M. Determinants of RNA Quality from FFPE Samples. PLoS One. 2 (12), e1261 (2007).

17.	Böhm, M., Wieland, I., Schütze, K., Rübben, H. Microbeam MOMeNT: non-contact laser microdissection of membrane-mounted native tissue. Am J Pathol 151 (1), 63-67 (1997).

18.	Schütze, K. et al. Cut out or poke in--the key to the world of single genes: laser micromanipulation as a valuable tool on the look-out for the origin of disease. Genet Anal. 14 (1), 1-8 (1997).

19.	Datta, S. et al. Laser capture microdissection: Big data from small samples. Histol Histopathol. 30 (11), 1255-1269 (2015).

20.	Desjardins, P. Conklin, D. NanoDrop microvolume quantitation of nucleic acids. J Vis Exp. (45), 2565 (2010).

21.	Ge, S. J., Gan, W. Q., Karlinsey, K., Zhou, B. Y., Pachter, J. S. Immuno-laser capture microdissection of perfusion-fixed mouse brain tissue coupled to RNA-seq. J Neurosci Methods. 423 110548 (2025).

22.	Takahashi, K., Beltran, W. A., Sudharsan, R. An optimized workflow for transcriptomic analysis from archival paraformaldehyde-fixed retinal tissues collected by laser capture microdissection. Exp Eye Res. 246 109989 (2024).

23.	Mahalingam, M. Laser Capture Microdissection: Insights into Methods and Applications. Methods Mol Biol. 1723 1-17 (2018).

24.	Kim, S. O., Kim, J., Okajima, T., Cho, N. J. Mechanical properties of paraformaldehyde-treated individual cells investigated by atomic force microscopy and scanning ion conductance microscopy. Nano Converg. 4 (1), 5 (2017).

25.	Erickson, H. S. et al. Quantitative RT-PCR gene expression analysis of laser microdissected tissue samples. Nat Protoc. 4 (6), 902-922 (2009).

26.	Pan, D. Y. et al. Laser Capture Microdissection-Based RNA Microsequencing Reveals Optic Nerve Crush-Related Early mRNA Alterations in Retinal Ganglion Cell Layer. Transl Vis Sci Technol. 9 (11), 30 (2020).

27.	Civita, P. et al. Laser Capture Microdissection and RNA-Seq Analysis: High Sensitivity Approaches to Explain Histopathological Heterogeneity in Human Glioblastoma FFPE Archived Tissues. Front Oncol. 9 482 (2019).

28.	Phan, H. V. et al. High-throughput RNA sequencing of paraformaldehyde-fixed single cells. Nat Commun. 12 (1), 5636 (2021).

29.	Thomsen, E. R. et al. Fixed single-cell transcriptomic characterization of human radial glial diversity. Nat Methods. 13 (1), 87-93 (2016).

30.	Lin, X. J. et al. A comparative analysis of RNA sequencing methods with ribosome RNA depletion for degraded and low-input total RNA from formalin-fixed and paraffin-embedded samples. Bmc Genomics. 20 (1), 831 (2019).

31.	Luzzi, V., Mahadevappa, M., Raja, R., Warrington, J. A., Watson, M. A. Accurate and reproducible gene expression profiles from laser capture microdissection, transcript amplification, and high density oligonucleotide microarray analysis. J Mol Diagn. 5 (1), 9-14 (2003).