# Submission ID #: 69913

# Scriptwriter Name: Sulakshana Karkala

# Project Page Link: https://review.jove.com/account/file-uploader?src=21323513

# Title: Automated Hospital Room Disinfection Utilizing a Novel Aerosolized Hydrogen Peroxide Microdroplet Disbursing Technology

# Authors and Affiliations:

# Catherine Lazarus1, Christopher Joseph2, Meghan Johnson3, Timothy Short3, Savannah Stevenson3, Ranya Sikka4, Marcela Mercado Montoya5, Tatiana Gomez Bustamante5, Kristin Grimes6, Erik Kulstad2, Barbara Hawkins7, Kathy Warye8

# 1Northwestern University

# 2University of Texas Southwestern Medical Center

# 3Rehabilitation Hospital of Southern New Mexico/Advanced Care Hospital of Southern New Mexico

# 4Granite Bay High School

# 5IN SILICO STEM S.A.S.

# 6Mount Graham Regional Medical Center

# 7Little Colorado Medical Center

# 8Infection Prevention Partners

☐   All author names and affiliations are correct (city/state/country information not included in video title page).

# Corresponding Authors:

Catherine Lazarus		(catelazarus2025@u.northwestern.edu)

# Email Addresses for All Authors:

Catherine Lazarus		(catelazarus2025@u.northwestern.edu)

Christopher Joseph 		(christopher.joseph@utsouthwestern.edu)

Meghan Johnson 		(MeghanJJohnson@ernesthealth.com)

Timothy Short 		(TimothyShort@ernesthealth.com)

Savannah Stevenson 		(savannahvillescas@gmail.com)

Ranya Sikka			(ranya.sikka@icloud.com)

Marcela Mercado Montoya	(marcela.mercadom@insilicostem.com)

Tatiana Gomez Bustamante 	(tatiana.gomez@insilicostem.com)

Kristin Grimes			(kristincollin04@hotmail.com)

Erik Kulstad 			(erik.kulstad@utsouthwestern.edu)

Barbara Hawkins 		(bhawkins@lcmcwmh.com)

Kathy Warye 			(kwarye@infectionpreventionpartners.com)

## Author Questionnaire

1. Microscopy: Does your protocol require the use of a dissecting or stereomicroscope for performing a complex dissection, microinjection technique, or something similar?  Enter Yes or No.

If Yes, can you record movies/images using your own microscope camera?

Enter Yes or No.

If your protocol involves microscopy but you are not able to record movies/images with your microscope camera, JoVE will need to use our scope kit.

If your microscope does not have a camera port, the scope kit will be attached to one of the eyepieces and you will have to perform the procedure using one eye.

Enter make and model of microscope.

If a dissection or stereo microscope is required for your protocol, please list all shots from the script that will be visualized using the microscope (shots are indicated with the 3-digit numbers, like 2.1.1, 2.1.2, etc.).

Click here to list microscope shots, using the shot numbers from the protocol section of the video script.

2. Software: Does the part of your protocol being filmed include step-by-step descriptions of software usage?  Enter Yes or No.

If Yes, we will need you to record using screen recording software.

We recommend using the screen capture program OBS. JoVE’s tutorial for using OBS Studio is provided at this link: https://review.jove.com/v/5848/screen-capture-instructions-for-authors?status=a7854k

As these files are necessary for finalizing your script, please upload all screen-captured video files to your project page as soon as possible.

3. Filming location: Will the filming need to take place in multiple locations?   Enter Yes or No.

If Yes, how far apart are the locations? Click to enter distance between locations.

4. Testimonials (optional): Would you be open to filming two short testimonial statements live during your JoVE shoot? These will not appear in your JoVE video but may be used in JoVE’s promotional materials. Enter Yes or No.

If Yes, please provide the full name and position (e.g., Director of [Institute Name], Senior Researcher [University Name], etc.) of the author willing to participate.

Enter full author name, Enter author position

To ensure that your script can be filmed in one day, the protocol sections are cumulatively restricted to 55 shots (shots are the 3-digit numbers like 2.1.1, 2.1.2…etc)

Current Protocol Length

Number of Steps: 19

Number of Shots: 38

# Introduction

Videographer: Obtain headshots for all authors available at the filming location.

Answers to these questions will become interview statements that you will deliver on camera.

- Introduction: Answer the 1st REQUIRED prompt and 1 additional question; Conclusion: Answer any 3 questions. No more than 5 interview statements will be included in the video.

- Enter the full name of the author who will deliver the statement.

- Speak naturally and avoid reading the lines.

- Limit the length of each statement to 20 words or fewer.

- Answers will be edited for length, clarity, and consistency with journal style guidelines.

Authors: Please note that the questions will not appear on screen. Please answer in stand-alone full sentence with sufficient context.

INTRODUCTION:

REQUIRED: Summarize your research focus and the main questions it aims to answer.

  - Enter full author name.: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

    - INTERVIEW: Named Talent says the statement above in an interview-style shot, looking slightly off-camera.

Describe the research gap or limitation in existing methods and how this protocol addresses them.

  - Enter full author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the research contexts or systems in which this protocol can be applied.

  - Enter full author name: Click here to answer question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

CONCLUSION:

What does this protocol allow researchers to study or measure?

  - Enter full author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Describe the most important consideration or challenge when performing this protocol.

  - Enter full author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Are there any additional methods or analyses that can be performed following this procedure?

  - Enter full author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

How can future studies build on this work? Describe the questions or research directions that can be investigated next.

  - Enter full author name: Click here if you choose this question. Please write in a style that you will be comfortable memorizing and speaking aloud. Limit length to 20 or fewer words.

Videographer: Obtain headshots for all authors available at the filming location.

# Protocol

Please review this section to make sure that it accurately describes your protocol. Use Track Changes when making edits or revisions.

- The two-digit steps (e.g., 2.1., 2.2.) with purple font are the narration.  JoVE is responsible for the narration of the protocol and results.

- Red italics are pronunciation guides indicating how the word will be spoken.

- Filming should take no more than 10 minutes per step. If a step takes more than 10 minutes, prepare the product for that step in advance.

- The three-digit shots (e.g., 2.1.1., 2.2.2.) are the actions that the videographer will capture.

- Title

Demonstrator: Click here to enter name of demonstrator(s)

If the same person is the demonstrator throughout, mention them once here and remove the "Demonstrator" field from the other sections; if the demonstrator changes, retain the field in the respective sections.

  - To begin, install the smart controller in the controller holder located on the handle of the cart that supports the fogging unit [1]. Ensure the smart controller is securely seated in the holder [2].

    - WIDE: Talent placing the smart controller into the holder mounted on the handle of the mobile cart supporting the fogging unit.

    - Talent adjusting the smart controller to ensure it is firmly secured in the holder.

  - Plug the controller cable into the controller charge port [1]. Insert the other end of the cable into the USB receptacle on the mobile cart [2]. Using the supplied cable clips or ties, secure the cable along the cart handle to hold it in place [3].

    - Talent inserting the controller cable into the charge port of the smart controller.

    - Talent connecting the other end of the cable into the universal serial bus receptacle on the mobile cart.

    - Talent fastening the cable neatly along the cart handle using the supplied cable clips or ties.

  - Now plug the power cable included in the mobile cart into the receptacle at the bottom of the aerosolized hydrogen peroxide disinfecting fogger [1]. Install the fogger into the designated depression in the mobile cart [2].

    - Talent inserting the power cable into the receptacle at the bottom of the aerosolized hydrogen peroxide disinfecting fogger.

    - Talent placing the fogger securely into the molded depression of the mobile cart.

  - Fill the tank of the aerosolized hydrogen peroxide disinfecting fogger with 1 gallon of 5 percent hydrogen peroxide solution [1]. Screw the cap back onto the tank securely [2-TXT].

    - Talent pouring 1 gallon of 5 percent hydrogen peroxide solution into the tank.

    - Talent tightening the cap back onto the tank. TXT: Refill tank as needed

  - Next, set up a cloud account using the email invitation sent to the user [1]. If no invitation has been received, request one [2-TXT].

Authors: Please create screen capture videos of the shots labeled as SCREEN, create a screenshot summary, and upload the files to your project page as soon as possible: https://review.jove.com/account/file-uploader?src=21323513

    - SCREEN: Display the email invitation and show the account setup process on the cloud platform registration page.

    - SCREEN: Show composing an email request to info@breezymed.com to obtain an invitation. TXT: Contact address: info@breezymed.com

  - Log in to the cloud dashboard in a web browser [1]. Select Devices in the left navigation menu on the cloud webpage dashboard to view the aerosolized hydrogen peroxide disinfecting fogger devices [2].

    - SCREEN: Show a web browser address bar with https://cloud.breezymed.com/ entered and the login page displayed.

    - SCREEN: Show the dashboard with the left navigation menu and click on Devices to display the list of foggers.

  - Press the VIEW ALL DATA button to display an overview of disinfection activity on the home screen and the details of the disinfection events [1].

    - SCREEN: Show the cursor clicking on the VIEW ALL DATA button on the dashboard home screen.

  - Next plug the aerosolized hydrogen peroxide disinfecting fogger power cord into a wall outlet [1].

    - Talent inserting the fogger power cord into a wall outlet.

  - Verify that the waistband light on the fogger is illuminated [1]. If the waistband light is yellow, press the button on top of the device to toggle the waistband light to blue [2].

    - Shot of the illuminated waistband light on the fogger.

    - Talent pressing the top button and the waistband light changing from yellow to blue.

  - Power on the smart controller [1]. Verify that the battery indicator shows the controller is charging [2].

    - Talent pressing the power button on the smart controller.

    - Shot of the smart controller screen displaying the charging battery icon.

  - Now sign in to the smart controller using the cloud account username and password if prompted [1]. Confirm that the screen displays the Scanning interface when signed in [2].

    - SCREEN: Show the login screen on the smart controller and the user entering credentials.

    - SCREEN: Display the Scanning screen on the smart controller after successful login.

  - Verify that the fogger on the cart appears in the list of devices on the smart controller screen [1]. Confirm that the serial number displayed matches the serial number located on the bottom of the fogger [2].

    - SCREEN: Show the list of available devices on the smart controller screen.

    - Shot of the serial number label on the bottom of the fogger and the matching name on the smart controller screen.

  - Press the + Add Favorite Device (Add-favorite-devices) button on the smart controller screen [1]. Choose the fogger from the list and press Continue [2].

    - SCREEN: Show the cursor tapping the + Add Favorite Device button.

    - SCREEN: Show selecting the fogger from the device list and pressing Continue.

  - From the Scanning screen, press the button corresponding to the fogger to be controlled [1]. Confirm that the device is now shown as a Favorite device on the screen [2]. Press Pair in the pop-up window if prompted [3].

    - SCREEN: Show tapping the fogger button on the Scanning screen.

    - SCREEN: Display the fogger listed under Favorite devices.

    - SCREEN: Show the pop-up window and the cursor pressing the Pair button.

  - To disinfect, position the aerosolized hydrogen peroxide disinfecting fogger in the corner of the room to be treated [1]. Orient the fogger to face the center of the room [2-TXT].

    - Talent moving the fogger into the corner of the room.

    - Talent adjusting the fogger to face the center of the room. TXT: Ensure that there are no objects within 5 ft of jets

  - Then plug the mobile cart cord into a wall outlet in the room to be disinfected [1].

    - Talent inserting the mobile cart cord into a wall outlet inside the treatment room.

  - Press the button on the smart controller screen corresponding to the fogger on the mobile cart [1]. Press the button for the appropriate fog time [2]. Then leave the room within the 7 second countdown displayed on the screen [3].

    - SCREEN: Show tapping the favorite fogger device on the smart controller screen.

    - SCREEN: Show selecting the desired fog time on the controller interface.

    - Talent exiting the room while a 7 second countdown is visible on the smart controller screen.

  - After fogging completes, keep the room closed and unoccupied for 10 minutes to allow disinfectant contact time [1]. Re-enter the room after 10 minutes or up to 15 minutes depending on the dosage and disinfectant used [2].

    - Shot of the closed room door with a visible timer indicating 10 minutes of contact time.

    - Talent opening the door and re-entering the room after the contact period.

  - Use commercial software COMSOL (Com-sol) Multiphysics version 6.3 with computational fluid dynamics and particle tracing modules to perform code verification [1]. Access the verification and validation models on the COMSOL website to review the available verification models [2].

    - SCREEN: Show the COMSOL Multiphysics version 6.3 interface with computational fluid dynamics and particle tracing modules enabled.

    - SCREEN: Display the COMSOL verification and validation models webpage with available verification models listed.

# Results

Please review this section to make sure that it accurately reflects your findings.

- This section will not be recorded by the videographer. It only includes the figures/tables from your manuscript (called LAB MEDIA).

- Use Track Changes when making edits or revisions. Ensure the voiceover length is below 200 words. Current word count: 126

- Please note that the video cannot include voiceover without an accompanying visual.

- Results

  - The standardized room geometry was defined as 11 feet by 12 feet with an 8 foot ceiling in the mathematical model [1].

    - LAB MEDIA: Figure 4.

  - The model showed that aerosolized hydrogen peroxide particle velocity ranged from 0.2 meters per second to over 25 meters per second [1].

    - LAB MEDIA: Figure 5.

  - Room coverage from surface contact with hydrogen peroxide was completed at 1 minute into the treatment course [1].

    - LAB MEDIA: Figure 6. Video editor: Show the room filled with green dots indicating full surface coverage.

  - Over the 24 month timeframe from January 2023 to July 2025, 24 patients known to be colonized or infected with carbapenem-resistant organisms were admitted [1].

    - LAB MEDIA: Figure 7.

  - Only 2 cases were detected after admission during this period [1]. Compared to the expected 9.6 in-facility transmissions based on a per-admission reproduction number of 0.40, the observed transmissions represented a significant reduction [2].

    - LAB MEDIA: Figure 7. Video editor: Highlight the bar labeled “With Intervention”

    - LAB MEDIA: Figure 7. Video editor: Highlight the “Without Intervention” bar