Replace Jove transcript 2.1 to 2.18.2 with below

2. Smart Controller and Aerosolized Hydrogen Peroxide Disinfecting Fogger Setup, Configuration, and Operation 

2.1. Insert controller on cart bracket and install controller cover with two screws
2.2. Plug USB cable into controller and cart charge port
2.3. Install USB cable on cart handle using clips
2.4. Install Breezy blue on cart by plugging device into cart power cord and setting device in cart pocket
2.5. Fill breezy blue tank
2.6. Plug Breezy Buggy power cord into wall outlet in room to be disinfected
2.7. Power on controller (volume up  + ON/OFF buttons)
Option 1 – Site already exists at location
2.8. PC SCREEN Site administrator invites new user to the site from cloud.breezymed.com Site Settings & Users
2.9. PC SCREEN If the site administrator is unknown, email Breezy Med at info@breezymed.com for more information 
2.10. PC SCREEN Using the Breezy Cloud invitation email link, setup account on cloud.breezymed.com
2.11. PC SCREEN Use the verification code sent to the email address to complete account setup
2.12. Controller SCREEN Use new account to login to controller
Option 2 – no site exists at location
2.13. Controller SCREEN If no site exists at the location, create a user account on the controller
2.14. Controller SCREEN Use the verification code sent to the email address to verify the account on the controller
2.15. Controller SCREEN Create site on the controller

2.16. Controller SCREEN Verify the Breezy Blue on the cart shows up in scanning list
2.17. Controller SCREEN Press +Add Favorite Device button, select the Breezy Blue on the cart from the list and press Submit
2.18. Controller SCREEN Select the Breezy Blue device shown as the favorite on the scanning screen
2.19. Controller SCREEN Press the Pair button 
2.20. Disinfect the room by pressing the appropriate room size button on the fogging screen
2.21. After fogging and settling has completed, re-enter the room, unplug breezy buggy from the wall outlet, wrap the cord around the handle and roll device out of room
2.22. PC SCREEN Login into cloud.breezymed.com 
2.23. PC SCREEN view the site fogging logs (Home -> VIEW LOGS)
2.24. PC SCREEN view the site devices (Devices ->Breezy Blue , Devices -> Smart Controller)



