TITLE:

Metagenomic Next-Generation Sequencing of Cerebrospinal Fluid for the Detection of Central Nervous System Pathogens

AUTHORS AND AFFILIATIONS:

Sumanth P. Reddy1, Jarod Morgenroth-Rebin1, Shiyin Wang1, Michael R. Wilson1

1University of California, San Francisco Weill Institute for Neurosciences, Department of Neurology, University of California, San Francisco.

Email addresses of co-authors:

Jarod Morgenroth-Rebin		(jarod.morgenrothrebin@ucsf.edu)

Shiyin Wang				(shiyin.wang2@ucsf.edu)

Michael R. Wilson			(michael.wilson@ucsf.edu)

Corresponding author:

Sumanth P. Reddy			(Sumanth.Reddy@ucsf.edu)

SUMMARY:

This research protocol outlines an optimized workflow for metagenomic next-generation sequencing (mNGS) of cerebrospinal fluid. By addressing challenges unique to low-biomass samples, the method enables robust pathogen detection across diverse resource settings, providing a versatile foundation for neuroinfectious disease research applications.

KEYWORDS:

Metagenomic next-generation sequencing, mNGS, CSF, cerebrospinal fluid, CNS infection

ABSTRACT:

Metagenomic next-generation sequencing (mNGS) has emerged as a powerful tool for unbiased pathogen detection and host transcriptional profiling in clinical and research settings. While its utility in diagnosing central nervous system (CNS) infections is increasingly recognized, cerebrospinal fluid (CSF) samples pose unique challenges due to low nucleic acid abundance and susceptibility to degradation. This study presents an optimized research-based protocol for Illumina sequencing platforms tailored to CSF mNGS, spanning sample handling, nucleic acid extraction, library preparation, sequencing, and bioinformatic analysis. Quality control approaches adaptable to both high and low-resource settings are also provided, including alternatives to capillary electrophoresis. This study demonstrates the protocol’s robustness through two representative cohorts: a high-depth, NovaSeq-based workflow and a cost-conscious NextSeq-based workflow. Across these cohorts, pathogens were detected in over 50% of cases, underscoring the method’s diagnostic potential even with resource-constrained adaptations. This protocol facilitates reproducible CSF mNGS, providing a foundation for diverse applications in neuroinfectious disease research and diagnostics.

INTRODUCTION:

Metagenomic next-generation sequencing (mNGS) is an increasingly utilized molecular technique that allows for the detection of a broad spectrum of nucleic acids within a given environmental or host sample1. By amplifying and sequencing the vast majority of the genetic material within a sample, mNGS does not require a preconceived hypothesis of the specific microorganisms that may be present. mNGS has shown utility as a diagnostic tool in a variety of clinical settings, such as for the detection of uncommon pathogens or atypical presentations of infectious diseases1–3. Furthermore, the host transcriptional profile is also captured by the RNA-sequencing data generated by the mNGS assay, allowing researchers to better understand the host immune response to various disease states4. The applicability of mNGS to a broad range of clinical and research contexts, along with the dramatic decrease in cost of generic sequencing in recent decades, has led to a tremendous growth in the use of this technology5.

In particular, mNGS has shown promise in the detection of central nervous system (CNS) infections, which remain challenging to diagnose2,6,7. Prior to the advent of clinical mNGS testing, the California Encephalitis Project examined causes of encephalitis across the state and found that over half of cases remained undiagnosed, despite utilizing an expanded panel of molecular tests and detailed clinical evaluation8. Now, over two decades later, a clinically validated cerebrospinal fluid (CSF) mNGS assay detected the presence of pathogenic organisms in the CSF of nearly 15% of patients with suspected CNS infections over a seven-year period7. Furthermore, among patients with confirmed infectious diagnoses, one in five were only diagnosed by this clinical CSF mNGS assay, emphasizing the utility of this testing approach7. In addition, studies have also demonstrated the potential of predicting the likelihood of various CNS infections based on the host transcriptome generated by the mNGS assay9–11.

However, the accessibility of this promising technique is limited by challenges unique to CSF sample handling and processing. In particular, the quantity of CSF DNA and RNA is often lower than the minimum quantity specified by library preparation kits12,13. Nucleic acid quality can also be suboptimal, given that CSF samples collected for clinical purposes are often stored at temperatures at which DNAses and RNAses can remain active14. This paper describes a research-based CSF mNGS protocol for Illumina sequencing platforms that has been optimized to maximize the likelihood of generating usable data for pathogen detection (Figure 1).

PROTOCOL:

This protocol was developed and tested using samples collected as part of studies approved by the UCSF Institutional Review Board.

[Place Figure 1 here]

- CSF sample handling considerations:

NOTE: mNGS results are optimal if CSF library preparation occurs at the time of collection, or if CSF samples are preserved via snap-freezing in liquid nitrogen, followed by long-term storage at -80 °C15. Alternatively, the use of preservative solutions such as DNA/RNA shield at the time of sample collection can allow for storage at room temperature for longer periods of time16,17. This may be preferable in situations when there are concerns about the reliability of the cold chain. In addition, avoidance of unnecessary freeze-thaw cycles can also help preserve nucleic acid integrity, and it may be advisable to pre-aliquot samples at the time of collection if multiple analytic workflows are anticipated for the CSF samples18,19.

- Extraction of CSF nucleic acid

CAUTION: Perform in an appropriate biosafety level cabinet with gloves and a laboratory coat to reduce the risk of transmission of infectious organisms. Reagents are harmful if swallowed, inhaled, or come into contact with skin. Dispose of used materials as hazardous chemical waste in compliance with local regulations.

2.1.	The Quick-DNA/RNA Kit generally works well for samples with a low abundance of nucleic acids.

2.2. Thaw the CSF samples on ice, then transfer 100 µL–1 mL of each sample to a 1.5 mL tube and centrifuge at 16,000 x g at 4 °C for 10 min.

2.3. Carefully pipette off the supernatant, leaving behind the pellet (which is often invisible to the naked eye) and 100 µL of supernatant.

2.4. Add 100 µL of DNA/RNA Shield to the remaining 100 µL of supernatant and pellet, and pipette mix thoroughly.

2.5. Then, add 600 µL of DNA/RNA lysis buffer and pipette mix until the solution is homogeneous.

2.6. Follow the remainder of the DNA/RNA extraction protocol as described by the manufacturer in Supplementary Appendices A, with the following important modifications specific to CSF:

2.6.1. Perform all centrifugation steps for 1 min at max speed (~21,000 x g) except for the final DNA/RNA wash step, which should run for 3 min at max speed.

2.6.2. Elute the sample by adding 22 µL (rather than 25 µL) of nuclease-free water directly onto the column matrix, then incubate for 3 min at room temperature. The higher volume generally results in a greater total quantity of nucleic acid eluted from the column.

2.6.3. After elution, pipette the flow-through back onto the column matrix and centrifuge again for one minute at max speed. Transfer the sample to a new 1.5 mL storage tube, and either proceed immediately to library preparation or store at -80 °C.

- CSF RNA library preparation

CAUTION: Perform in an appropriate biosafety cabinet, wearing gloves and a laboratory coat. Reagents are harmful if swallowed, inhaled, or come into contact with skin. Dispose of used materials as hazardous chemical waste in compliance with local regulations.

NOTE: Important considerations prior to beginning: rRNA depletion is preferred over polyadenylation enrichment for isolation of mRNA, given that the rRNA depletion does not require additional bead clean steps (which can reduce the yield of usable mRNA), and some neuroinvasive RNA virus transcripts lack polyadenylation. ERCC (External RNA controls consortium) RNA Spike-In can be added as a quality control measure, though it is not strictly necessary. If the RNA was stored at -80 °C following extraction, thaw on ice prior to beginning.

3.1. RNA fragmentation and priming: Pipette mix the following in a polymerase chain reaction (PCR) tube: 3.5 µL RNA sample, 0.5 µL 1:2500 diluted ERCC spike-in, 4 µL first-strand reaction buffer (5x), 1 µL random primers, and 1 µL 1:100 diluted rRNA (Table of Materials). In a thermocycler, set the heated lid to 105 °C and incubate the samples at 75 °C for 2 min, 70 °C for 2 min, 65 °C for 2 min, 60 °C for 2 min, 55 °C for 2 min, 37 °C for 5 min, and 25 °C for 5 min.

3.2. First-strand cDNA synthesis: Pipette mix the fragmented and primed RNA (10 µL) with 8 µL nuclease-free water and 2 µL first-strand synthesis enzyme mix. In a thermocycler, set the heated lid to 105 °C and incubate the samples at 25 °C for 10 min, 42 °C for 15 min, and 70 °C for 15 min.

3.3. Second-strand cDNA synthesis: Pipette mix the first-strand synthesized DNA (20 µL) with 8 µL second-strand synthesis reaction buffer and 4 µL second-strand synthesis enzyme mix. In a thermocycler, incubate the samples for 1 h at 16 °C, with the heated lid turned off.

3.4. Perform a magnetic bead purification (Supplementary Appendices B) using a 1.8x ratio of SPRI magnetic beads (144 µL). Elute into 53 µL nuclease-free water and transfer 50 µL of the final supernatant (containing purified double-strand cDNA) to a clean nuclease-free PCR tube. At this point, the samples can be safely frozen at -20 °C overnight if needed.

3.5. cDNA library end preparation: If the sample was stored at -20 °C overnight, thaw on ice prior to restarting. Pipette mix the purified ds-cDNA (50 µL) with 7 µL of end-prep reaction buffer and 3 µL of end-prep enzyme mix.

3.6. In a thermocycler, incubate the samples at 20 °C for 30 min, and 65 °C for 30 min, with the heated lid turned off.

3.7. Adapter ligation (perform this step on ice): Create a 1:100 dilution of adapter in nuclease-free water. Pipette the end-preparation reaction mixture (60 µL) into a tube containing 30 µL ligation master mix, 1 µL ligation enhancer, and 2.5 µL 1:100 diluted adapter. Ensure the adapter is added separately (e.g., not premixed with the ligation master mix or ligation enhancer) to avoid adapter dimer formation.

3.8. In a thermocycler, incubate the samples at 20 °C for 15 min, with the heated lid turned off.

3.9. Proceed immediately to magnetic bead purification (Supplementary Appendix 2) using a 0.9x ratio of SPRI magnetic beads (87 µL).

3.10. Elute into 17 µL nuclease-free water, and transfer 15 µL of the final supernatant to a clean nuclease-free PCR tube

3.11. Barcoding PCR: Pipette the purified, adapter-ligated cDNA (15 µL) into a tube with 3 µL USER enzyme, 25 µL Q5 Master Mix, and 10 µL unique barcoded primers. In a thermocycler, set the heated lid to 105 °C, incubate the samples at 37 °C for 15 min, then 98 °C for 30 s, and then perform 19 cycles of 98 °C for 10 s and 65 °C for 75 s. Finish the PCR by incubating at 65 °C for 5 min.

3.12. Perform a final magnetic bead purification using a 0.8x ratio of magnetic beads (43 µL). Elute into 23 µL nuclease-free water, and following the final bead separation step, transfer 20 µL to a clean nuclease-free PCR tube. At this point, the samples can be safely frozen at -20 °C overnight if needed.

- CSF DNA library preparation

CAUTION: Perform in an appropriate biosafety level cabinet with gloves and a laboratory coat. Reagents are harmful if swallowed, inhaled, or come into contact with skin. Dispose of used materials as hazardous chemical waste in compliance with local regulations.

4.1. If the extracted DNA was stored at -20 °C overnight, thaw on ice prior to restarting.

4.2. DNA fragmentation and end-preparation: Mix the first-strand reaction buffer thoroughly by vortexing and pipette mixing the solution to resuspend any precipitate. Pipette mix 3.5 µL of the extracted DNA, 22.5 µL nuclease-free water, 7 µL first-strand reaction buffer, and 2 µL first-strand enzyme mix.

4.3. In a thermocycler, set the heated lid to 105 °C and incubate the samples at 37 °C for 5 min, then 65 °C for 30 min.

4.4. The adapter ligation and barcoding PCR process is identical for the DNA and RNA libraries. Repeat sections 7–10 of the “CSF RNA Library Preparation” with the fragmented and end-prepped DNA to complete the DNA library preparation.

- Library quality control and pooling

5.1. Quantify the concentration of the sample libraries via the DNA quantification kit and corresponding fluorometer. Water controls should have substantially less (or unquantifiable) concentrations compared to the rest of the samples.

5.2. Assess the sample library sizes by running the samples on an automated capillary electrophoresis machine.

5.3 Alternatively, if an automated capillary electrophoresis machine is not available, further amplify 1-2 µL of the sample by performing a PCR reaction with Illumina universal primers, and then perform electrophoresis of 10 µL of the final product on a 2% agarose gel. Dispose of used materials as hazardous chemical waste in compliance with local regulations.

NOTE: The optimal library length is approximately 400–500 base pairs, so future fragmentation incubations should be adjusted accordingly if the library sizes are larger or shorter than desired.  In addition, it is critical to assess for the presence of adapter dimers, which are approximately 150 base pairs in length and occur when two adapter molecules inadvertently ligate together without an insert sequence. Even if present in low quantities, adapter dimers tend to cluster and sequence efficiently on the flow cell, reducing the proportion of usable reads during a sequencing run. As such, if adapter dimers are present, perform an additional magnetic bead purification with a 0.8x ratio of SPRI magnetic beads. (Of note, if >10% of the sample is adapter dimer, it may take 2-3 rounds of magnetic bead purification to remove most of the adapter dimers).

5.3. To ensure a similar depth of sequencing across samples within the same sequencing run, it is necessary to pool equimolar quantities of each sample. If the library sizes are comparable across all samples, simply pool equivalent masses of each RNA library and repeat separately for each DNA library.

- Sequencing considerations

6.1. The pooled samples are now ready for sequencing on Illumina sequencing devices. If sequencing in-house (rather than via a specialized sequencing core facility), carefully denature and dilute the pooled libraries according to the sequencing device instructions prior to loading the pool onto the reagent cartridge (typical concentration is 4 nM for most Illumina sequencing platforms).

6.2. Calculate the estimated sequencing depth by dividing the expected data yield for the Illumina flow cell by the number of multiplexed samples in the run.

NOTE: Determining the desired sequencing depth depends on several factors, including cost, the sequencing device used, and the objectives of the experiment. Two examples of different sequencing depths are presented below in the Representative Results section.

- Data analysis for pathogen detection

7.1. Utilize the open-source web-based platform Chan Zuckerberg ID (CZID, https://www.czid.org, Illumina mNGS Pipeline v8.3) to perform the metagenomic analysis for pathogen detection. CZID also includes options to upload samples via command line interface and direct transfer from an Illumina BaseSpace account to CZID. A detailed overview of CZID mNGS data analysis is available at https://chanzuckerberg.zendesk.com/hc/en-us/articles/13770737266196-Guide-to-mNGS-Data-Analysis

7.2. Log into CZID, click “upload,” select “metagenomics” for analysis type, and select the input fastq files from the sequencing run. Input the necessary information for each sample, including sample name, sample type (CSF), and nucleotide (RNA or DNA).

NOTE: The uploaded files undergo automated processing via the CZID pipeline, which incorporates several computational tools such as STAR, Bowtie2, Trimmomatic, PriceSeq, and GSNAP, among others, to remove low-quality reads and human read sequences. The pipeline then aligns and assembles the filtered sequences utilizing the National Center for Biotechnology Information (NCBI) nucleotide (NT) and protein (NR) databases via the Minimap2 and DIAMOND software programs, respectively. The final CZID output includes NT/NR taxon counts and contig counts (contiguous, overlapping segments generated during assembly) for each sample.

7.3. Once processed, click the project folder containing the samples and navigate to the summary dashboard.

7.4. Scroll down to examine the number of reads per sample (and compare to the expected number of reads based on the Illumina flow cell used and the number of samples multiplexed in the sequencing run), the percentage of reads passing quality control filtering, and the duplicate compression ratio to assess for biased PCR overamplification.

7.5. The presence of contamination by environmental nucleic acid fragments - either acquired during the CSF sample collection process, or during extraction/library preparation - is a common occurrence20. To determine whether bacterial, fungal, and parasitic reads are true positives, potential methods utilizing the normalized read counts (reads per million, or rPM) can be considered:

7.5.1. Create a background model (e.g., using water samples that have been extracted, prepared, and sequenced) directly within CZID to filter out any taxons with rPMs that are similar to or lower in abundance relative to the background model (and therefore, likely to represent contamination). Within the project folder, check each of the water control samples and then click “background model” to create. While analyzing each sample, select the background model and click “threshold filters” to apply an “NTZ Score” > 1 (or higher) to filter out likely contaminants.

7.5.2. Utilize pre-specified rPM thresholds, such as rpMsample > 10*(average rPMcontrols), or log10(rPMsample) at least 1 log greater than log 10(average rPMentire cohort). First, download the taxon-specific rPM for each sample by checking each sample on the project page, clicking download, and selecting “Combined Sample Taxon Results.” The resulting .csv file can be analyzed using statistical software such as R or Stata as desired.

NOTE: Given the lower relative abundance of viral pathogens, any virus with known neuroinvasive potential and at least one read aligning to the viral genome should be considered a positive result (and ideally, confirmed via repeat mNGS or pathogen-specific clinical testing). In addition, the analytic pipeline requires each read in the fastq file to be computationally aligned to the most likely reference genome in the NCBI database. As a result, some reads may be relatively non-specific to a particular pathogen and could theoretically have aligned to a broad range of organisms rather than the specific taxon assigned in CZID. Given this risk, manual confirmation of positive calls should be performed by confirming the alignment via the web-based NCBI BLAST tool (the specific reads for a taxon can be directly sent from CZID to NCBI BLAST by clicking the blast icon next to the relevant taxon).

REPRESENTATIVE RESULTS:

Library quality control

Following DNA/RNA extraction, the measured concentration of the extracted DNA is generally 1–10 ng/µL. However, the concentration of the extracted RNA is often on the order of pg/µL, and thus it is frequently undetectable by spectrophotometers. Following library preparation, the purified and amplified cDNA and DNA libraries are generally at a concentration of 0.5–10 ng/µL and 10–30 ng/µL, respectively. The capillary electrophoresis and electropherogram data in Figure 2A demonstrate a cDNA sample library that is over-fragmented and contains adapter dimers. Prior to pooling and sequencing, the adapter dimers should be removed via additional rounds of bead cleanup, using a 0.8x magnetic bead-to-sample ratio. In contrast, the data in Figure 2B show a higher quality cDNA library without significant adapter dimer contamination that is ready for pooling and sequencing.

[Place Figure 2 here]

Representative cohort A

The mNGS results of two representative cohorts with suspected CNS infections are shown below. Cohort A contains 68 samples from a TB-endemic population with a high prevalence of HIV. The samples were submitted to the UCSF Center for Advanced Technology for sequencing via the Illumina NovaSeq Platform as part of a larger pooled sequencing run, and the average depth was 116,150,823 DNA reads (standard deviation, SD: 37,770,481 reads), and 95,966,675 RNA reads (standard deviation, SD: 44,601,339 reads). After undergoing automated CZID computational filtering, there were an average of 2,168,623 DNA reads (SD: 1,064,346 reads) and 8,663,265 RNA reads (SD: 6,749,945 reads) that passed filters.

Representative cohort B

Cohort B, in contrast, contains 43 samples from a different resource-limited setting and was sequenced in-house at a shallower read depth on the Illumina NextSeq 500/550 system with High-Output 150-cycle reagent kits. To minimize costs and maximize the number of samples that could be analyzed, ERCC was omitted, and the protocol was miniaturized by performing each reaction at 0.5x volume, thereby considerably reducing consumable costs. Capillary electrophoresis was also unavailable, so quality control was performed by over-amplifying the samples via additional rounds of PCR with universal Illumina primers to estimate library size and adapter dimer concentration prior to pooling. The RNA libraries were sequenced in a single run with an average depth of 13,033,943 reads per sample (SD 9,110,614 reads), whereas the DNA libraries were split into two separate runs with an average depth of 30,203,814 reads per sample (SD: 18,634,933 reads). After computational processing, an average of 268,664 DNA reads (SD: 588,478 reads) and 3,696,052 RNA reads (SD: 3,756,262 reads) passed filters.

Pathogen detection

Figure 3 demonstrates the diversity of the pathogens detected in these two cohorts. At least one pathogen was detected in 73.5% of Cohort A (50/58 samples) and 51.2% of Cohort B (22/43 samples).

[Place Figure 3 here]

FIGURES AND TABLE LEGENDS:

Figure 1: Cerebrospinal fluid (CSF) metagenomic next-generation sequencing (mNGS) sample processing and analytic workflow for pathogen detection. This diagram demonstrates the primary steps in the CSF mNGS wet-lab and bioinformatic pathways.

Figure 2: Example of capillary electrophoresis and electropherogram results. (A–B) represent the capillary electrophoresis trace and electropherogram from a sample that contains adapter dimers (blue arrow). In addition, the average library size is slightly lower than desired, either due to sample degradation or over-fragmentation. (C–D) demonstrate a high-quality library with few adapter dimers and an optimal size (~400–500 base pairs) for Illumina sequencing. Abbreviations; bp = base pairs.

Figure 3: Pathogens detected via CSF mNGS in two distinct clinical cohorts. Counts reflect total pathogens detected; some samples contained multiple pathogens. Abbreviations; CSF = cerebrospinal fluid, mNGS = metagenomic next-generation sequencing

Supplementary Appendices: (A) CSF nucleic acid extraction. This contains additional steps for the CSF nucleic acid extraction protocol. (B) Magnetic bead purification. This contains the general workflow for magnetic bead purification, which is used several times during the library preparation

DISCUSSION:

There are important considerations that may require alterations to this protocol, depending on the desired sequencing depth, the sequencing platform used, and the availability of equipment for quality control. Nonetheless, the two representative cohorts presented here demonstrate that robust pathogen detection is possible across a broad range of experimental conditions. Specifically, Cohort B demonstrated that, despite several cost-saving measures – shallow-depth sequencing, the exclusion of ERCC controls, miniaturization of the protocol, and the use of traditional gel electrophoresis for quality control (rather than a bioanalyzer or TapeStation) – a diverse range of neuroinvasive pathogens was still detected in the samples.

There are several critical steps and considerations to maximize the chances of success during the protocol. It is important to work efficiently during extraction (particularly prior to the addition of DNA/RNA Shield) to avoid unwanted RNA degradation. During the various steps of library preparation, the mixtures should be thoroughly mixed to ensure homogenization, while avoiding excess bubbles, which can reduce library yield. If bubbles are apparent after mixing, gently flick the PCR tube strip and pulse-spin. Adapter dimers are another critical consideration during the protocol. To reduce the likelihood of adapter dimer excess, add the diluted adapter separately from the master mix while the PCR tubes are on ice, and mix thoroughly (ideally with a multichannel pipette if preparing multiple samples). When excess adapter dimers are present during quality control assessment, it is often necessary to perform 2–3 additional magnetic bead cleans (at a 0.8x bead-to-library volume) to adequately remove them. If this problem persists despite optimal pipetting technique, consider further diluting the adapter before addition.

An important limitation specific to CSF is that low RNA yield is common, particularly in samples with low pathogen burdens and low pleocytosis. The concentration of RNA is often unquantifiable via common accessible quantification methods such as fluorometry until the final PCR step is performed, which can make troubleshooting of low cDNA library yield challenging. A common pitfall that can result in low RNA yield is over-drying of the beads during the cleanup steps – to avoid this, elute in water once the beads appear matte to the naked eye, but have not developed any visible cracks. In addition, ensuring adequate mixing of the reagents with the sample is critical, especially before cDNA synthesis. This is particularly important if performing the miniaturized version of the protocol, where the sample and reagent volumes are minute and more prone to inconsistencies. If low RNA yield persists despite no identifiable issues in the library preparation protocol, the initial sample may be overly degraded (due to sample handling, etc.), rendering the initial input RNA concentration too low for the enzymes to efficiently convert the RNA to cDNA. If this is the case, the researcher can consider eluting the final RNA product from the extraction into a smaller volume to boost the concentration. The addition of ERCC can also be helpful in these instances, as it can act as a “carrier” for the sample RNA to boost the concentration and prevent the sample RNA from sticking to the tube walls.

Additional modifications can also be made to this protocol to accommodate other research goals. For instance, a semi-automated workflow (e.g., utilizing Agilent Echo and Agilent Bravo automated liquid handling platforms)21 for higher-throughput mNGS applications has been successfully implemented in the author’s laboratory. In addition, apart from distinct modifications in the initial steps of the nucleic acid extraction, a similar library preparation protocol can be applied to serum and brain tissue specimens22,23. The protocol can also be adapted to allow for the incorporation of unique molecular identifiers (UMIs) to allow for more accurate read quantification in bulk RNA-seq transcriptomic applications24. Furthermore, although authors utilized CZID for the computational processing and to confirm detected pathogen reads, it is possible to use other workflows or to create custom pipelines that may be better suited to the researcher’s needs25,26.

It is important to note that this is a research-based protocol rather than a validated clinical assay; thus, the results obtained from this protocol should not be viewed as a substitute for standard diagnostic methods or relied on in clinical care. Nonetheless, this protocol offers an accessible entry point for researchers in a variety of settings to apply this technique to better understand the epidemiology of CNS infections. Unfortunately, the specific sensitivity and specificity of this protocol are not available for the two reference cohorts presented in the results above, due to the very limited clinical testing available in both resource-limited settings. However, a meta-analysis of high-quality studies using clinically validated or research-based mNGS assays (including multiple studies from author lab, utilizing similar variations of this protocol) found a pooled sensitivity of 0.70 (95% confidence interval [CI] 0.67–0.72) and a pooled specificity of 0.93 (95% CI 0.92–0.94) for the diagnosis of CNS infection, relative to the final diagnosis obtained by clinical adjudication27. In comparison, the pooled sensitivity and specificity for conventional microbiologic testing (including culture, PCR, and antibody-based methods) were 0.37 [95% CI 0.34–0.41] and 0.95 [0.94–0.96], respectively27.

As such, clinically validated mNGS assays represent a useful adjunctive diagnostic tool, likely best used in conjunction with the high-sensitivity tests that exist for several treatable CNS infections (cryptococcus, human herpesvirus-1, varicella zoster virus). They are particularly valuable in immunocompromised hosts, who may harbor opportunistic and atypical CNS infections28. Although the cost and accessibility of such assays limit their practical use in many settings, mNGS has the potential to reduce the diagnostic gap in CNS infections, improve antibiotic stewardship, and reduce overall healthcare utilization as the technology continues to scale. As discussed above, interpretation of the sequencing results requires nuance and manual confirmation because of challenges related to the background contamination and sample handling considerations, and multidisciplinary interpretation with infectious disease specialists, neurologists, and laboratory medicine is strongly recommended29,30.

Despite these challenges, CSF mNGS remains an important clinical and research-based tool for a variety of applications. The protocol presented here is reproducible and specifically optimized for CSF samples, which are the most paucicellular and challenging to sequence among commonly analyzed human bodily fluids. In addition, this workflow highlights CSF sample handling considerations to preserve nucleic acid integrity, modifications to standard kit-based workflows to optimize performance, and quality control steps critical to ensuring that the final sample libraries are suitable for high-throughput sequencing.

ACKNOWLEDGMENTS:

This study was supported by funding from the Fogarty International Center of the National Institutes of Health under Award Number D43TW009343. We would like to thank the many researchers at the Wilson Lab, UCSF, and CZI who have contributed to the development and optimization of mNGS protocols for diverse research applications over the past several years.

DISCLOSURES:

M.R.W. receives unrelated research grant funding from Roche/Genentech, Novartis, and Kyverna Therapeutics and is a founder and board member of Delve Bio Inc. He has provided consulting services to Pfizer, Vertex Pharmaceuticals, Ouro Medicines, and Indapta Therapeutics.

REFERENCES:

1.	Chiu, C. Y., Miller, S. A. Clinical metagenomics. Nat Rev Genet. 20 (6), 341–355 (2019).

2.	Wilson, M. R. et al. Clinical metagenomic sequencing for diagnosis of meningitis and encephalitis. N Engl J Med. 380 (24), 2327–2340 (2019).

3.	Forbes, J. D., Knox, N. C., Ronholm, J., Pagotto, F., Reimer, A. Metagenomics: the next culture-independent game changer. Front. Microbiol. 8, doi: 10.3389/fmicb.2017.01069 (2017).

4.	Wang, Z., Gerstein, M., Snyder, M. RNA-Seq: a revolutionary tool for transcriptomics. Nat Rev Genet. 10 (1), 57–63 (2009).

5.	DNA Sequencing Costs: Data at <https://www.genome.gov/about-genomics/fact-sheets/DNA-Sequencing-Costs-Data>.

6.	Wilson, M. R. et al. Chronic meningitis investigated via metagenomic next-generation sequencing. JAMA Neurology. 75 (8), 947–955 (2018).

7.	Benoit, P. et al. Seven-year performance of a clinical metagenomic next-generation sequencing test for diagnosis of central nervous system infections. Nat Med. 30 (12), 3522–3533  (2024).

8.	Glaser, C. A. et al. In search of encephalitis etiologies: diagnostic challenges in the California encephalitis project, 1998—2000. Clinical Infectious Diseases. 36 (6), 731–742 (2003).

9.	Ramachandran, P.S. et al. Integrating central nervous system metagenomics and host response for diagnosis of tuberculosis meningitis and its mimics. Nat Commun. 13, 1675 (2022).

10.	Xing, Z. et al. Integrating DNA/RNA microbe detection and host response for accurate diagnosis, treatment, and prognosis of childhood infectious meningitis and encephalitis. J Transl Med. 22, 583 (2024).

11.	Omura, C. et al. 298. Host response profiling from clinical metagenomic sequencing data for diagnosis of central nervous system infections. Open Forum Infect Dis. 12 (1), ofae631.088 (2025).

12.	Andriuk, L. V., Skochiĭ, P. G. Nucleic acids of the cerebrospinal fluid in acute disorders of cerebral circulation. Zh Nevropatol Psikhiatr Im S S Korsakova. 88 (7), 32–34 (1988).

13.	Olausson, J. et al. Optimization of cerebrospinal fluid microbial DNA metagenomic sequencing diagnostics. Sci Rep. 12 (1), 3378 (2022).

14.	Becskei, A., Rahaman, S. The life and death of RNA across temperatures. Comput Struct Biotechnol J. 20, 4325–4336 (2022).

15.	Wang, C., Liu, H. Factors influencing degradation kinetics of mRNAs and half-lives of microRNAs, circRNAs, lncRNAs in blood in vitro using quantitative PCR. Sci Rep. 12 (1), 7259 (2022).

16.	Cb, T. et al. DNA/RNA preservation in glacial snow and ice samples. Frontiers in microbiology. 13, doi: 10.3389/fmicb.2022.894893 (2022).

17.	J, K., E, M., A, K., A, K., A, M. Optimisation of sample storage and DNA extraction for human gut microbiota studies​. BMC microbiology. 21 (1), doi: 10.1186/s12866-021-02233-y (2021).

18.	X, J. et al. The impact of repeated freeze-thaw cycles on the quality of biomolecules in four different tissues. Biopreservation and biobanking. 15 (5), doi: 10.1089/bio.2017.0064 (2017).

19.	Bp, K. et al. Multiple freeze-thaw cycles lead to a loss of consistency in poly(A)-enriched RNA sequencing. BMC genomics. 22 (1), doi: 10.1186/s12864-021-07381-z (2021).

20.	Zinter, M. S., Mayday, M. Y., Ryckman, K. K., Jelliffe-Pawlowski, L. L., DeRisi, J. L. Towards precision quantification of contamination in metagenomic sequencing experiments. Microbiome. 7 (1), 62 (2019).

21.	Mayday, M. Y., Khan, L. M., Chow, E. D., Zinter, M. S., DeRisi, J. L. Miniaturization and optimization of 384-well compatible RNA sequencing library preparation. PLoS One. 14 (1), e0206194 (2019).

22.	Sakiyama, Y. et al. Brain biopsy and metagenomic sequencing enhance aetiological diagnosis of encephalitis. Brain Commun. 7 (3), fcaf1650 (2025).

23.	Sakiyama, Y., Yuan, J., Yoshimura, A., Takei, J., Ando, M., Takashima, H. Etiological insights of encephalitis and encephalitis-like central nervous system diseases: brain tissue metagenomic next-generation sequencing. Neurology. 102 (7–1), 3782 (2024).

24.	Sun, J. et al. Correcting PCR amplification errors in unique molecular identifiers to generate accurate numbers of sequencing molecules. Nat Methods. 21 (3), 401–405 (2024).

25.	Tan, J. K. et al. Laboratory validation of a clinical metagenomic next-generation sequencing assay for respiratory virus detection and discovery. Nat Commun. 15(1), 9016 (2024).

26.	Akshatha, Prasanna. Clin-mNGS: Automated pipeline for pathogen detection from clinical metagenomic data. at <https://github.com/AkshathaPrasanna/Clin-mNGS> (2025).

27.	Marra, P. S. et al. Metagenomic next-generation sequencing in patients with infectious meningoencephalitis: a comprehensive systematic literature review and meta-analysis. Open Forum Infect Dis. 12 (5), ofaf274 (2025).

28.	Ramachandran, P. S., Wilson, M. R. Metagenomics for neurological infections - expanding our imagination. Nat Rev Neurol. 16 (10), 547–556 (2020).

29.	Salter, S. J. et al. Reagent and laboratory contamination can critically impact sequence-based microbiome analyses. BMC Biol. 12, 87 (2014).

30.	Simner, P. J., Miller, S., Carroll, K. C. Understanding the promises and hurdles of metagenomic next-generation sequencing as a diagnostic tool for infectious diseases. Clin Infect Dis. 66 (5), 778–788 (2018).