TITLE:

Measurement of Nitric Oxide Production in Mouse Tracheal Epithelial Cell Cultures Differentiated at the Air Liquid Interface

AUTHORS AND AFFILIATIONS:

Muddassar Iqbal1, David Mahan1, Adam Clark2, Akueba Bruce1, Anthony W. DeMartino1, Paola Corti2,*

1Department of Medicine, University of Maryland School of Medicine, Baltimore MD 21201, USA

2Department of Biochemistry and Molecular Biology, University of Maryland School of Medicine, Baltimore MD 21201, USA

Email addresses of the co-authors:

Muddassar Iqbal: miqbal@som.umaryland.edu

David Mahan: dmahan@fpi.umaryland.edu

Adam Clark: Adam.Clark@som.umaryland.edu

Akueba Bruce: Akueba.Bruce@som.umaryland.edu

Anthony W. DeMartino: ADeMartino@som.umaryland.edu

Corresponding author: Paola Corti (PCorti@som.umaryland.edu)

SUMMARY

This protocol describes a method to quantify nitric oxide (NO) production in airway epithelial cultures differentiated at the air–liquid interface (ALI) by measuring nitrite in apical secretions using a triiodide–based chemiluminescence assay.

ABSTRACT

Airway epithelial cultures differentiated at the air–liquid interface (ALI) provide a physiologically relevant model to study nitric oxide (NO) signaling in ciliated cells. Here, we describe a protocol to quantify NO production by measuring nitrite (NO₂⁻), a stable oxidation product of NO, in samples collected from ALI-differentiated mouse tracheal epithelial cells. Apical washes, basal media and cell lysates are collected and analyzed using a triiodide-based chemiluminescence assay coupled to a nitric oxide analyzer (NOA). Upon injection into the reaction vessel, nitrite is chemically reduced to NO, which is detected by ozone-based chemiluminescence. Signal intensity is quantified and converted to picomoles of nitrite using a standard calibration curve, and values are normalized to total cellular protein content. Assay performance is validated using pharmacological modulation of NO levels. Upon treatment with the NO donor Diethylenetriamine NONOate (DETA–NONOate) and the NO synthase inhibitor N(ω)–Nitro–L–arginine methyl ester (L–NAME) we observe significant differences in nitrite levels across conditions. The most robust and reproducible signal is observed in apical washes, corresponding to the amount of NO released from the airway epithelium. This method reliably detects NO in samples from ALI–differentiated airway epithelia, providing an accurate in vitro platform to quantify NO production and model diseases associated with abnormal NO metabolism.

INTRODUCTION

A hallmark of several airway diseases, including asthma, chronic obstructive pulmonary disease (COPD), cystic fibrosis (CF) and primary ciliary dyskinesia (PCD), is altered nitric oxide (NO) homeostasis measurable as changes in exhaled NO levels, which can largely vary in these diseases, being elevated in COPD and asthma1,2, while drastically reduced in CF and PCD3,4. These diseases commonly exhibit abnormalities in airway epithelial motile cilia, the key effectors of mucociliary clearance5. Because NO and its downstream signaling pathways are critical regulators of ciliary motility, understanding how NO production and signaling are altered in these contexts is essential for elucidating mechanisms of mucociliary dysfunction. Accordingly, a method that enables precise detection of NO production and release in airway epithelia provides a valuable tool for dissecting the regulation and physiological role of this signaling network. Such approaches are particularly useful in experimental settings where direct measurement of NO is challenging due to its short half-life and rapid reactivity.

NO regulates airway cilia biogenesis and function by influencing ciliated cell polarity and the coordinated beating of motile cilia6–8. However, the mechanisms that maintain NO homeostasis in the airways remain incompletely understood. Airway epithelial cultures differentiated at air–liquid interface (ALI) provide a valuable in vitro system for investigating these mechanisms. At present, this model is widely used to study the impact of environmental insults on the airway epithelium9–11 and to elucidate the pathophysiology of airway diseases12. Compared with conventional submerged cultures13,14, ALI–differentiated epithelia more faithfully reproduce the architecture and function of the native airway, including epithelial polarization, mucociliary differentiation, and realistic gas–liquid exchange, making them especially suitable to study apical NO release, signaling, and related ciliary function. This system is therefore well-suited for experiments aimed at quantifying NO-derived metabolites released into the airway lumen, particularly under conditions of epithelial differentiation or pharmacological manipulation.

Accurate measurement of NO concentration is essential, as both excessive and deficient NO levels can impair ciliary function. Currently, fluorescent probe–based methods (e.g., diaminofluoresceins, DAFs) available for detecting NO are only semi–quantitative and are limited by pH sensitivity, oxygen dependence and lack of specificity, as they can react with cellular oxidants and antioxidants, resulting in semi–quantitative measurements15. In addition, electrochemical sensors (e.g., NO–sensitive amperometric electrodes) enable real–time detection of NO16 but can be difficult to implement in complex biological systems due to NO short half–life, interference from redox–active species biofouling of the electrode surface, and challenges in maintaining stable calibration17. Alternative chemiluminescence–based approaches including direct gas–phase NO detection and reductive assays targeting different nitrogen oxide species, have also been described, but their application to airway epithelial cultures remains unexplored. These limitations highlight the need for sensitive, quantitative methods that can reliably detect low levels of NO-derived species in small-volume biological samples. To overcome these limitations, we established a protocol to quantify NO in apical secretions collected from ALI–differentiated airway epithelia using a Nitric Oxide Analyzer (NOA). The NOA is a highly sensitive chemiluminescence–based detector traditionally used to measure NO metabolites in chemical samples and more recently adapted for the analysis of biological specimens, including blood, tissue homogenates, and cell culture supernatants7,18,19. NO is a highly reactive signaling molecule rapidly oxidized to nitrite in tissues under conditions of normal oxygen levels, with an exceptionally short half–life. Thus, nitrite as end product of NO oxidation serves as an indicator of total NO production in biological systems20–22. This sensitive and reproducible chemiluminescence method allows detection of nitrite/NO at picomolar levels23,24, making it ideal for measuring small changes in low–volume biological samples such as those derived from MTEC ALI cultures. However, this approach is best suited for conditions where nitrite accumulation reflects recent NO production and may be influenced by factors such as sample handling, background nitrite contamination, and epithelial differentiation state. Therefore, appropriate controls and careful experimental design are required to ensure accurate interpretation of results.

PROTOCOL:

All animal procedures were reviewed and approved by the University of Maryland Institutional Animal Care and Use Committee (Protocol # AUP–00004632).

NOTE: The protocol requires approximately 2 days for isolation of  mice tracheal epithelial cells (MTEC), followed by 7–10 days for expansion, 5–6 days for proliferation in transwell inserts, 21 days for differentiation (time needed to acquire sufficient multi–ciliated cells), 24 h of drug treatment and one additional day for sample collection, NOA calibration, injections and data acquisition.

- Preparation of Media and Solutions

NOTE: Refer to Table 1 for details on media ingredients.

  - Ham’s F–12 media

Prepare the solution inside a sterile laminar flow hood. Add 5 mL of 100X Penicillin/Streptomycin antibiotics and 0.5 mL Amphoterecin B (Fungizone) to 500 mL Ham’s F12 medium in a sterile bottle. Mix by gently inverting the bottle 5-10 times to ensure complete mixing. Filter sterilize the solution using a 0.22 μm filter under sterile conditions, and store at 4 °C.

CAUTION: Penicillin/streptomycin and amphotericin B are antimicrobial agents that may cause irritation or allergic reactions. Handle using appropriate personal protective equipment, including gloves and lab coat.

  - Pronase solution

Add 15 mg pronase to 10 mL sterile Ham’s F12 medium in a sterile tube to obtain 0.15% pronase solution. Mix by gently inverting or pipetting up and down until the pronase is fully dissolved. Prepare fresh immediately before use.

  - DNase I solution

Add 5 mg crude pancreatic DNase I to 9 mL Ham’s F12 medium and 1 mL 10 mg/mL Bovine Serum Albumin (BSA) solution in a sterile tube to obtain a 0.5 mg/mL DNase I solution. Gently mix by pipetting up and down until the DNase I is fully dissolved, avoiding bubble formation. Once fully dissolved, aliquot the solution into 1 mL sterile microcentrifuge tubes and store at –20 °C.

  - Ham’s F12 media with fetal bovine serum (FBS)

Prepare the solution at room temperature inside a sterile laminar flow hood. Add 10 mL fetal bovine serum (FBS) to 40 mL Ham’s F12 media with antibiotics in a sterile conical tube to obtain 20% FBS in F12 medium. Mix gently inverting the tube 5-10 times until the solution is homogeneous. Filter sterilize the solution using a 0.22 μm filter under sterile conditions, and store at 4 °C.

  - MTEC Expansion media

Prepare the MTEC expansion media at room temperature inside a sterile laminar flow hood by combining the components listed in Table 1 in a sterile bottle. Mix thoroughly by gently swirling or inverting the container until all components are fully combined. Filter sterilize the solution using a 0.22 μm filter under sterile conditions. Add ROCK inhibitor Y–27632 (5 mM stock solution) to a final concentration of 10 μM and Notch γ–secretase inhibitor DAPT (10 mM stock solution) to a final concentration of 5 μM immediately before use, and mix gently by inversion.

  - MTEC Proliferation media

Prepare the MTEC proliferation media at room temperature inside a sterile laminar flow hood by combining the components listed in Table 1. Mix thoroughly by gentle inversion or swirling until all components are fully combined. Filter sterilize the medium using a 0.22 μm filter under sterile conditions. Immediately before use, add ROCK inhibitor Y–27632 (5 mM stock solution) to a final concentration of 10 μM and retinoic acid (5 μM stock solution) to a final concentration of 0.05 μM. Mix gently by inversion before use.

  - MTEC Differentiation media

Prepare the MTEC differentiation medium as at room temperature inside a sterile laminar flow hood by combining the components listed in Table 1 in a sterile container. Mix thoroughly by gentle inversion or swirling until all components are fully combined. Filter sterilize the medium using a 0.22 μm filter under sterile conditions. Immediately before use add retinoic acid to a final concentration of 0.05 μM and mix gently by inversion.

- Isolation of Mouse Tracheal Epithelial Cells (MTEC)

  - Euthanize three adult mice (8–20 weeks old) by carbon dioxide (CO2) inhalation in an approved euthanasia chamber according to institutional IACUC guidelines. Confirm death before tissue collection. Place the animal dorsally (chest facing up) on an aseptic surface. Using clean forceps, gently lift the skin and cut it open with surgical scissors from the lower lip to the abdomen to expose the trachea.

  - Move the salivary glands a side andr remove the connective tissue and trachealis muscle to expose the cartilage rings. Dissect the full length of trachea and place it in a Petri dish containing sterile Ham’s F12 medium. Combine three tracheas in one dish.

  - Transfer the dish into a laminar flow hood and transfer the tracheas to another dish with fresh Ham’s F12 medium. Remove any extratracheal tissue and cut the trachea longitudinally through the lumen.

  - Transfer all the tracheas into a 50 mL screw–cap conical tube containing sterile 10 mL 0.15% pronase solution. Incubate at 4 °C for 16-18 hours.

- Cell Isolation and Initial Expansion

  - The next day invert the tube containing the tissues 10–15 times to dislodge epithelial cells and incubate at 4 °C for 45 min. Stop the pronase reaction by adding 10 mL Ham’s media containing 20% FBS and mix by inversion 10–15 times.

  - Add 10 mL Ham’s F12 medium with FBS to three 15 mL tubes. With a Pasteur pipette transfer the tracheal tissue sequentially from the 50 mL tube to the first-second and third 15 mL tube and invert 10–15 times in each tube to dislodge the tracheal epithelial cells. Discard the tissue after the final transfer.

  - Combine the whole media from the three 15 mL conical tubes with the medium in the 50 mL tube. Centrifuge at 360 x g for 10 min at 4 °C. Carefully remove the supernatant and add 0.5 mL of 0.5 mg/mL DNase I solution. Incubate on ice for 5 min.

  - Centrifuge at 360 x g for 5 min at 4 °C. Carefully discard the supernatant and resuspend the cell pellet in 7–10 mL MTEC expansion medium.

- Expansion and Differentiation of MTEC

  - Prepare a collagen coated T75 cell culture flask using 10 mL collagen I (rat tail) at 50 μg/mL and incubate at 37 °C for 16-18 hours. Aspirate the collagen solution from the T75 flask and wash the flask three times with PBS (1X). Allow the flask to dry completely.

  - Seed the cell suspension consisting of 7-10 mL MTEC expansion medium containing the cells into the collagen coated flask. Place the flask in an incubator at 37 °C and 5% CO2. Replace the expansion medium every 2 or 3 days and monitor cell number until 70% confluency.

At 70% confluency, add 1 mL Accutase solution and incubate for 10–15 min at 37 °C to detach the cell monolayer. Resuspend the cells in 10 mL of MTEC proliferation medium. Centrifuge at 1057 × g for 5 min at room temperature.

  - Discard the supernatant, resuspend the cells in proliferation medium and count the cells.

  - Seed 2 x 105 cells in 0.5 mL volume per transwell insert (12 mm membrane diameter, 0.4 μm membrane pore size) precoated with collagen. Add 1.5 mL medium to the basal chamber of each insert and place the transwell in the incubator at 37 °C.

  - Change media in both the apical and basal chamber on day 3 and allow the cells to reach 100% confluency, usually by day 5 or 6. Include one insert without cells as a control

  - Once confluent, wash the cells apically with 0.5 mL MTEC differentiation media. Initiate air-liquid interface (ALI) by adding 0.75 mL differentiation medium to the basal chamber only.  leaving the cells in apical chamber exposed to air.

  - Every 2–3 days, wash cells in the apical surface with differentiation medium and replace basal medium. Allow cells to differentiate at ALI for 21 days.

- Immunolabeling and imaging of ciliated epithelial cells

  - At ALI day 7, 14 and 21, fix the cells with 4% paraformaldehyde (PFA).

  - Perform immunolabelling of cells with anti–FoxJ1 antibody to detect FoxJ1, a transcription factor essential for ciliogenesis 25,26, and anti–acetylated α–tubulin antibody to detect acetylated α–tubulin, a structural marker of motile cilia according to published protocols 27–29.

CAUTION: PFA is a hazardous chemical, so wear appropriate PPE (nitrile gloves and lab coat at a minimum) when fixing the cells.

  - Acquire images of the immunolabelled cells by confocal microscopy using a 40x or 60x objective. Collect Z–stacks spanning a depth of 10 µm with each step size of 0.5 µm. Set laser power, detector gain and exposure settings to avoid signal saturation and maintain comparable acquisition parameters across samples.

  - Quantify the FoxJ1–positive or acetylated tubulin–positive cells as a percentage of total cells (DAPI–positive) per region of interest (ROI) using the cell counter Plugin in FIJI imaging software.

  - In the GraphPad Prism software (v. 10.0) generate graphs and perform statistical analysis using a two–tailed Student’s t–test after assessing the data distribution for normality using Shapiro–Wilk test.

- Drug Treatment of differentiated airway epithelia

  - At ALI day 21 prepare prepare differentiation medium (3 mL total volume) in a sterile tube at room temperature.

  - Add 7.5 μL Diethylene triamine–NONOate or DETA/NO from a 100 mM stock solution (prepared in 10 mM NaOH) to the differentiation medium to achieve the final concentration of 250 μM in 3 mL. Mix gently by inversion and use immediately.

  - Add 750 μL of the final solution in three basal chambers of differentiated cells and incubate for 24 h at 37 °C.

  - Prepare 1 mM Nω–Nitro–L–Arginine Methyl Ester (L–NAME) by adding 30 μL from the 100 mM L–NAME stock solution (prepared in DMSO) to MTEC differentiation media for a total volume of 3 mL.

  - Add 750 μL of the final solution in three basal chambers and incubate for 24 h at 37 °C.

  - Add 30 µL dimethyl sulfoxide (DMSO) only (untreated group) to 3 mL basal media (1% v/v).

  - Add 750 μL of the final solution in the basal chamber of three inserts (untreated).

- Sample collection

NOTE: For NOA analysis it is necessary to pool samples  media, apical wash, and cell lysate samples from three inserts.

  - At ALI day 22, collect samples from the differentiating airway epithelia beginning with                                                                                              750 μL of basal media from each of three inserts per experimental group.

  - Collect basal media incubated in inserts without cells to serve as blank control accounting for potential external nitrite contamination.

  - Pool the basal media from three inserts and store immediately at –20 °C until injection into NOA.

  - To collect the nitrite/NO present in the apical mucus of the epithelium, add 50 μL pre–warmed PBS (1X) buffer (sterile, aliquoted and frozen) on top of cells in each insert and incubate for 30 min at 37 °C.

  - NOTE: During the 30–min incubation of cells with PBS to collect apical wash, PBS can also be added in the basal chamber to maintain hydration on both sides of the cells.

  - Collect the PBS containing the mucus overlying the epithelium.

  - Pool the PBS washes from three inserts to obtain around 150 µL apical mucus/PBS from each experimental condition which will be enough for two NOA injections.

  - Store the apical wash samples at –20 °C until further use.

  - Store also the PBS wash buffer for use as a blank control during NOA injections.

  - Wash the apical and basal sides of the cells in three inserts with 500 μL prewarmed PBS twice.

  - Add 500 μL Accutase solution apically and incubate for 15–20 min until the cells completely detach from the membranes.

  - Once fully detached, add serum–free RPMI–1640 media to the cells (at 3:1 media to Accutase ratio).

  - Pool the detached cells from the three inserts in a single tube.

  - Spin the cells at 360 x g for 5 min at 4 °C.

  - Resuspend the cell pellet in 50 μL lysis solution (RIPA buffer + Halt 1X protease inhibitor cocktail) and incubate on ice for 45 min.

  - Spin the lysate at 14,800 x g for 20 min at 4 °C and collect supernatant.

  - Use 3 μL assay to determine the protein concentration through the BCA assay.

  - Dilute the remaining lysate solution 1:1 in RIPA buffer to make enough volume for two 50 μL NOA injections.

  - Store the lysates at –80 °C until injecting into the NOA and the RIPA buffer at –20 °C.

  - Store the RIPA buffer used to resuspend the cell pellet at –20 °C to be used as blank and contamination control for lysate samples.

- NOA Calibration and Sample Injections

  - Prepare triiodide reagent by dissolving 400 mg potassium iodide and 250 mg iodine in 28 mL glacial acetic acid and 8 mL ultrapure water.

  - Perform reagent preparation in a chemical fume hood using appropriate personal protective equipment (lab coat, gloves, and eye protection). CAUTION: Iodine and glacial acetic acid are hazardous and can cause irritation and corrosion to skin and eyes.

  - Store at room temperature and use within two weeks of preparation.

  - Prepare 1 mM stock solution of sodium nitrite in nitrite–free water.

  - Store the stock solution at −20 °C until use.

  - Prepare fresh working standards (0.125–4.0 µM) in PBS immediately before injections.

NOTE: For NOA calibration, use sterile, tightly capped PBS to minimize nitrite contamination.

  - Turn on the nitrogen (N2) carrier gas and initiate gas flow through the system.

  - Assemble the glass purge vessel (Part # 015–043, ECO PHYSICS) and add 4 mL of triiodide reagent into the vessel with 20 μL anti–foam diluted 1:10.

  - Securely cap it with a septum held in place by the screw cap to ensure an airtight seal.

  - Pour 10 mL sodium hydroxide (NaOH) into the NaOH trap.

  - On the display screen of the NOA instrument (nCLD 88) select “Measurement”, “Select Gas and Range”.

  - Set the instrument to “MR2:50.00”, corresponding to a measurement range between 0 and 50 parts per billion of NO, and select “Fast” for filter setting.

  - Adjust the nitrogen gas pressure until the reactor pressure stabilizes between 27–29 psi (186–200 kPa).

  - Open the eDAQ chart software on the computer connected to the NOA.

  - Set the acquisition range to 10/s and the voltage to 10 V.

  - Navigate to Setup>Channel Settings, select Channel 1.

  - Set the display scale to 2:1 on the NOA software.

  - Start the data acquisition on the NOA instrument and in the software.

  - Monitor the baseline until it stabilizes.

  - Inject 10 μL of PBS through the septum into the purge vessel using a Hamilton syringe to confirm baseline stability.

  - Record the chemiluminescent signal for each injection.

  - Allow the signal trace to return fully to baseline before the next injection.

  - Inject 10 μL of the 0.125 μM nitrite standard into the purge vessel.

  - Repeat twice for each nitrite standard, to obtain duplicate measurements that will later be averaged.

  - Annotate each injection in the “Comment” tab (e.g., “10 μL 0.5 μM nitrite”).

  - Wipe the syringe on a lint–free tissue and rinse 3–4 times with sterile nitrite–free water to remove excess triiodide.

  - Repeat this cleaning step after each injection.

  - Allow the signal to return to baseline between injections (20–30 s).

  - Open the Flow Analysis window.

  - Review peak height and time values for each injection in the analysis table.

  - Highlight each peak and adjust the selection to calculate the area under the curve (AUC).

  - Toggle the display from “S” to “C” to enable calibration mode.

  - Enter the known nitrite amounts for each standard in the “Amount” column.

  - Click “Calibration curve” button to generate the standard curve.

  - Verify that the calibration curve meets acceptance criteria (R²>= 0.99).

  - Thaw biological samples previously collected from ALI cultures on ice.

  - Inject 50 μL of blank cell–free medium into purge vessel.

  - Inject 50 μL of each basal media sample into the purge vessel.

  - Record the chemiluminescence signal for each injection.

  - Perform duplicate injections (50 μL each) for each sample.

NOTE: The 50 μL injections result in frequent bubble formation. To limit this issue, replace the triiodide reagent with anti–foam in the purge vessel regularly.

  - Annotate each sample in the “Comment” tab prior to injection.

  - Inject 50 μL of cell–free PBS (apical blank) into purge vessel.

  - Inject apical wash samples (two 50 µL injections per sample) and record the corresponding signal peaks.

  - Inject RIPA buffer twice, followed by lysate samples (two 50 µL injections per sample) and record the corresponding signal peaks. NOTE: Always allow the chemiluminescent trace to reach back to the baseline and stabilize. Wait 20-30 seconds, then inject the next sample.

  - At the end of sample injections, open the Flow Analysis window.

  - Highlight each sample peak to calculate the area under the curve (AUC).

  - Determine nitrite concentrations for each sample using the calibration curve generated from standard solutions.

- Data collection, normalization and analysis

  - Export the data, including peak height (V), area under the curve (V.s) and mean area under the curve from the “Flow Analysis” function in the eDAQ software to an Excel file. The exported file will tabulate the height (V) versus time (s) of signal traces for the entire NOA experiment.

  - Determine the picomoles (pmol) of nitrite injected into the NOA by multiplying the nitrite concentration (μM) by the injection volume (μL).

  - Average the picomoles (pmol) of nitrite measured in controls (Cell–free media, PBS wash buffer PBS or RIPA buffer).

  - Subtract the corresponding control values from the pmol measured in each 50 μL sample injection.

  - Determine the total picomoles (pmol) of nitrite in each sample by multiplying the pmol measured in the injected portion of the sample by the exact total sample volume collected. Use 750 μL for the media, –150 μL for the apical wash and around 90–100 μL for the lysate.

  - Additionally, for lysates only, multiply the total nitrite pmol by 2 to account for the dilution factor.

  - Normalize the pmol of nitrite obtained for each sample by dividing these values by the total protein amount (µg) present in the lysate of the corresponding sample, obtained previously through the BCA assay.

  - Average the nitrite levels (pmol/µg) from duplicate injections for each sample.

  - Plot normalized nitrite levels (pmol/µg of total protein) for each sample using GraphPad Prism software.

  - Assess data distribution for normality using the Shapiro–Wilk test.

  - Perform statistical analysis using a two–tailed Students’ t-test.

RESULTS:

The overall goal of this procedure is to quantify nitric oxide (NO) levels in mouse tracheal epithelial cell (MTEC) cultures differentiated at air–liquid interface (ALI). Primary cells are isolated and differentiated according to established protocols27,28. Over the course of 21 days at ALI, the cells are differentiated into a pseudostratified airway epithelium containing ciliated cells27,28. Levels of ciliation were assessed by immunolabeling for Foxj1 or acetylated α–tubulin. As differentiation progressed, a gradual increase in FoxJ1 expression within the nuclei of differentiating cells was detected (Figure 1A), accompanied by a corresponding increase in the number of cells carrying multiple cilia labelled by acetylated α–tubulin (Figure 1B). By ALI day 21, approximately 10% of the airway epithelial cells display motile cilia, consistent with previous reports27,28.

We previously described the use of ozone–based chemiluminescence for detection of NO in biological samples7,18,21,22. In the present study, we apply this method for the quantification of nitrite, the primary stable oxidation product of NO, as a proxy for NO production at the time of sample collection. In this assay, samples from airway epithelial cell cultures are injected into a purge vessel (Figure 2A) containing triiodide reagent, where the nitrite (NO2–) present in the sample is protonated to nitrous acid and subsequently reduced by iodide to yield stoichiometric amounts of NO (Equation 1). The released NO is then carried by an inert gas into a chemiluminescent detector, where it reacts with ozone to produce excited nitrogen dioxide (Equation 2). As this molecule returns to ground state, it emits a photon that is detected by a photomultiplier tube, amplified, and converted into voltage signal23–24 (Figure 2A).

Equation 1: NO2– + 2l– + 2H+           NO + I2 + H2O

Equation 2: NO + O3            NO2* + O2

To generate a calibration curve, standard sodium nitrite solutions prepared in PBS are analyzed under identical conditions (Figure 2B, left). Upon injection into triiodide, the reaction proceeds immediately, producing a voltage response proportional to the total NO derived from nitrite in the standard solution. Chemiluminescent signal traces increase linearly with rising nitrite concentrations. A valid calibration curve, expressed as mean area under the curve (AUC) versus picomoles (pmol) of nitrite, should yield a correlation coefficient (r²) greater than 0.99, as shown in Figure 2B (right).

This procedure was used to analyze samples from airway epithelial cells. As proof of concept that we are indeed measuring NO (via nitrite), we treated cells with N(ω)–Nitro–L–arginine methyl ester (L–NAME), the specific inhibitor of NO synthases (NOSs)29, and with the NO donor Diethylenetriamine NONOate (DETA/NO)19. Of note, NOS enzymes are the primary source of NO, but NO can also be generated through NOS–independent pathways, such as the reduction of nitrite under certain conditions. Airway epithelial cells were harvested at ALI day 22 for NO measurements. We first collected samples from basal media and apical washes containing apical mucus secreted by the cells. After removing a small aliquot of the lysate for protein quantification, the remaining lysate was also analyzed. Following calibration, samples were injected into the NOA; each sample was injected twice. The pmol nitrite/NO were normalized by protein content in the corresponding cell lysate to obtain nitrite levels (pmol/µg) from at least three independent experiments and values were plotted in GraphPad Prism software (Figure 3).

Our results show that in the basal medium, the amount of nitrite/NO is non–significantly altered upon L–NAME treatment compared to untreated controls. In contrast, the positive control DETA/NO treatment shows an expected increase due to NO release from the compound over the 24 h incubation period (Figure 3A). The observed negative values in some of the samples arose due to background subtraction during data processing, as the nitrite signal measured in the corresponding blank (cell–free media) slightly exceeded that of the sample, resulting in negative values after subtraction. This reflects variability in background nitrite levels rather than true negative nitrite concentrations. In the cell lysate samples, nitrite/NO levels were similar across untreated cells (0.0692 ± 0.01282 pmol/µg), the L–NAME–treated cells (0.071 ± 0.03 pmol/µg) and DETA/NO–treated cells (0.0699 ± 0.027 pmol/µg DETA/NO) (Figure 3B).

Importantly, we observed a significant decrease in nitrite/NO levels in the apical washes from L–NAME treated cells as compared to the untreated cells (0.221 ± 0.011 pmol/µg in untreated versus 0.143 ± 0.0154 pmol/µg in L–NAME treated) (Figure 3C). Additionally, DETA/NO treatment resulted in more than 6.5–fold increase in detected nitrite/NO (1.457 ± 0.62 pmol/µg). This polarized pattern of NO distribution likely reflects the physiology of ALI–differentiated epithelia, where NO production and diffusion are directed toward the apical compartment, where mucus efficiently traps nitrite as a stable end product. We observed a minimal change in basal medium or lysate fractions, may be due to intracellular scavenging, dilution, or limited NO retention in these compartments.

Overall our data confirms that the mucus secreted by airway epithelial cells and collected in the apical wash is a reliable reservoir reflecting NO release from ciliated cells, validating the sensitivity and accuracy of the current chemiluminescent method. By enabling direct quantification of NO in an in vitro system that accurately recapitulates airway physiology, this approach offers unique potential to dissect NO–related mechanisms and therapeutic modulation in airway disease.

FIGURE AND TABLE LEGENDS:

Figure 1. Ciliated cell population increases progressively with ALI differentiation.

Mouse tracheal epithelial cells (MTECs) differentiated at ALI at various time points and immunostained with anti–FoxJ1 antibody (A) or anti–acetylated Tubulin (B). Nuclei were counterstained with DAPI. Images are 3D confocal projections. Representative images from 10–15 different regions of interest (ROI) are shown. Scale bar = 50 µm. On the right, quantification of the percentage of cells expressing Foxj1 or acetylated tubulin at different states of ALI. Each bar is Mean ± SD (n = 10–15 technical repeats from a single culture). Statistical analysis was performed using Student’s t-test (two–tailed); ns, not significant; ** p < 0.01 and ***; p < 0.001.

Figure 2. NOA chemiluminescent signals increase linearly with nitrite concentration.

(A) Schematic depicting triiodide based chemiluminescent detection of nitrite/NO using the nitric oxide analyzer (NOA) in samples (basal media, apical wash and cellular lysates) collected from ALI–differentiated airway epithelia. Image created with BioRender.com. B) Trace peaks of nitrite standards with each injection at 10 µL volume into the Purge vessel (left). Calibration curve (right) depicts the mean area under the curve for each nitrite standard solution correlating with the picomoles (pmol) of nitrite injected into the NOA.

Figure 3. Triiodide based chemiluminescent assay detects nitrite/NO in samples from ALI differentiated airway epithelia.

(A) Trace peaks (left), absolute values (middle) and quantification (right) of normalized nitrite/NO detected in the basal medium of ALI–differentiated airway epithelium. (B) Trace peaks (left), absolute values (middle) and quantification (right) of normalized nitrite/NO in the cellular lysate. (C) Trace peaks (left), absolute values (middle) and quantification (right) of normalized nitrite/NO secreted and trapped in the overlying mucus collected as apical washes. Each sample was injected twice. The detected total picomoles (pmol) of nitrite/NO were normalized to the total μg of protein and averaged for duplicate injections. Each bar represents Mean ± SD nitrite/NO levels normalized to pmol/μg protein (n = 3, independent cultures). Statistical analysis was performed using Student’s -test (paired, two–tailed); ns, not significant; ** p < 0.01.

Table 1. Media compositions for mouse tracheal epithelial cell (MTEC) culture and ALI differentiation

The media ingredients for MTEC expansion in T75 cell culture flask, proliferation in 12 mm (0.4 µm pore size) transwell inserts and differentiation at air liquid interface (ALI) are listed along with the vendor–provided or self–prepared stock and final concentrations.

DISCUSSION

Multi–ciliated cells within the airway epithelium are critical for NO synthesis and signaling, which in turn is essential for regulating ciliary beating, mucociliary clearance, and airway homeostasis6–8. NO is a highly reactive molecule and even small changes in NO levels are associated with airway diseases1–4. Therefore, maintaining finely tuned NO levels is essential for normal airway function. Here, we describe a method to detect nitrite as a proxy for NO production in ciliated airway epithelial cells cultured at the air–liquid interface (ALI). Mouse tracheal epithelial cells (MTEC) were isolated and differentiated at ALI, with slight modifications from previously published methods27–28, to obtain a sufficient population of multi–ciliated cells needed to produce measurable nitrite/NO levels (Figure 1A, B).

Nitrite in biological samples originates primarily from the oxidation of NO by oxygen or reactive oxygen/nitrogen species such as superoxide and peroxynitrite21,22. Minor contributions to nitrite may arise from alternative enzymatic or redox pathways, however, under physiological conditions, these are limited, and most cellular nitrite derives from NO oxidation30,20,31. Accordingly, nitrite quantification serves as a reliable readout of NO production in this assay. Critical steps in this protocol include maintaining low background nitrite levels and ensuring consistent NOA operating conditions. Because NO levels in these samples are expected to be very low, relatively large injection volumes are required to achieve detectable signal intensity. The low yield of secreted NO/nitrite from each 12 mm transwell insert necessitates pooling the apical washes from three inserts per condition to reach detection levels in the NOA. However, sample volume must be carefully controlled, as excessive injection of biological material, particularly from lysates with high protein content, can cause protein–induced bubbling in the purge vessel, leading to overflow into the NOA gas lines. This represents a key technical limitation and requires careful optimization of injection volume and sample composition. To prevent this, the injection volume is maintained at 50 µL into 4 mL of triiodide solution, which provides a balance between signal sensitivity and instrument safety. ALI differentiation for a longer period (28–35 days, for example) can possibly increase the secreted NO in each insert due to a higher number of multi–ciliated cells. Because multiciliated cells are major contributors to NO synthesis, differences in epithelial differentiation state are expected to influence NO output. This highlights the potential of the assay to detect physiologically relevant changes in NO production across different epithelial differentiation stages. We note that increased differentiation is also associated with greater mucus accumulation, which may introduce variability due to differences in residual components within apical samples. Despite this, such differences remain biologically relevant when interpreting NO measurements. Additionally, bigger commercially available inserts (24 mm membrane diameter) that can seed a greater number of airway epithelial cells can be used to address the issue of low nitrite/NO yield per insert. It is worth noting from a method standpoint that all the apical wash samples collected from pooling three inserts were injected into the NOA in duplicate. Importantly, we analyzed biological replicates, and the results shown here represent data from at least three independent cultures.

A major troubleshooting consideration when using this highly sensitive assay is nitrite contamination from external sources. Because nitrite is ubiquitous in the environment, contamination can be difficult to avoid. Therefore, it is essential to run appropriate controls alongside experimental samples. To control for nitrite contamination in the media, we incubated cell–free media in the basal chambers of empty (cell–free) inserts and used these as blanks to subtract background signal in the NOA readings. Similarly, PBS and RIPA buffer samples prepared at the time of apical washes and cell lysis were collected and analyzed in parallel. Particularly in the basal medium, we detected an elevated nitrite background likely originating from serum components and protein additives, which can either contain trace nitrite/nitrate or undergo slow spontaneous oxidation of amino acids and peptides that generate nitrite over time. This explains why nitrite levels in cell–free basal medium can be as high as untreated samples. Consequently, this buffer is not ideal for NOA measurements, as the high background can mask small biological differences.

In addition to controlling for exogenous nitrite, it is important to include both positive and negative experimental controls. We inhibited NO synthesis employing the NOS inhibitor L–NAME to reduce NO signals29 and applied the NO donor DETA/NO to release NO19. When all reagents are freshly prepared and samples immediately frozen prior to NOA analysis, these controls effectively minimize the likelihood of false readings due to contamination. Under these conditions, the apical wash compartment provided the most robust and reproducible readout of nitrite levels. This finding reflects the polarized nature of airway epithelia, where NO generated by the cells predominantly diffuses toward the apical compartment and becomes oxidized to nitrite, which accumulates in the mucus over time. In the basal medium no clear differences were detected, also due to substantial endogenous nitrite contamination. In the cell lysate, the measured nitrite/NO in the DETA/NO treated cells were unexpectedly low, suggesting that most of the released NO is rapidly directed apically and trapped as nitrite rather than retained intracellularly. This highlights a limitation of the method in intracellular or basal compartments and supports the use of apical washes as the preferred sample type for this assay. Thus, the apical wash represents the most reliable compartment for assessing NO release and validating the sensitivity of this detection method. Overall, this method provides a sensitive and reproducible approach to quantify NO–derived nitrite in ALI cultures and can be applied to investigate NO signaling under physiological and disease–relevant conditions.

ACKNOWLEDGMENTS

We thank Dr. Mark T. Gladwin (Dean and Professor, University of Maryland School of Medicine) for guidance on study design. We also thank the University of Maryland School of Medicine’s Confocal Microscopy Core (Baltimore, MD) for access to confocal microscopy resources. This work was supported by National Institutes of Health (NIH) grant 5R01HL168775 to Dr. Paola Corti.

DISCLOSURES

The authors declare no conflicts of interest.

REFERENCES:

- Alcázar–Navarrete, B., Ruiz Rodríguez, O., Conde Baena, P, Romero Palacios, P. J., Agusti, A. Persistently elevated exhaled nitric oxide fraction is associated with increased risk of exacerbation in COPD. Eur Respir J. 2018;51(1): 1701457.

- Krantz, C., Janson, C., Alving, K., Malinovschi, A. Nasal nitric oxide in relation to asthma characteristics in a longitudinal asthma cohort study. Nitric Oxide 2021;106:1–8.

- Thomas, S. R., Kharitonov, S. A., Scott, S. F., Hodson, M. E., Barnes, P. J. Nasal and exhaled nitric oxide is reduced in adult patients with cystic fibrosis and does not correlate with cystic fibrosis genotype. Chest 2000;117(4):1085–9.

- Shapiro, A. J. et al. Nasal Nitric Oxide Measurement in Primary Ciliary Dyskinesia. A Technical Paper on Standardized Testing Protocols. Ann Am Thorac Soc. 2020;17(2): e1–e12 .

- Legendre, M., Zaragosi, L. E., Mitchison, H. M. Motile cilia and airway disease. Semin Cell Dev Biol. 2021;110:19–33.

- Li, D., Shirakami, G., Zhan, X., Johns, R. A. Regulation of ciliary beat frequency by the nitric oxide–cyclic guanosine monophosphate signaling pathway in rat airway epithelial cells. Am J Respir Cell Mol Biol. 2000;23(2): 175–81.

- Rochon, E. R. et al. Cytoglobin regulates NO–dependent cilia motility and organ laterality during development. Nat Commun. 2023;14(1): 8333.

- Mikhailik, A. et al. nNOS regulates ciliated cell polarity, ciliary beat frequency, and directional flow in mouse trachea. Life Sci Alliance 2021;4(5): e202000981.

- Schamberger, A. C., Staab–Weijnitz, C. A., Mise–Racek, N., Eickelberg, O. Cigarette smoke alters primary human bronchial epithelial cell differentiation at the air–liquid interface. Sci Rep. 2015;5: 8163.

- Gindele, J. A. et al. Intermittent exposure to whole cigarette smoke alters the differentiation of primary small airway epithelial cells in the air–liquid interface culture. Sci Rep. 2020;10(1): 6257.

- Ghosh, B. et al. Effect of sub–chronic exposure to cigarette smoke, electronic cigarette and waterpipe on human lung epithelial barrier function. BMC Pulm Med. 2020;20(1): 216.

- Baldassi, D., Gabold, B., Merkel, O. Air–liquid interface cultures of the healthy and diseased human respiratory tract: promises, challenges and future directions. Adv Nanobiomed Res. 2021;1(6): 2000111.

- Bessa, M. J. et al. In Vitro Toxicity of Industrially Relevant Engineered Nanoparticles in Human Alveolar Epithelial Cells: Air–Liquid Interface versus Submerged Cultures. Nanomaterials (Basel) 2021;11(12): 3225.

- Allouche, Y. et al. Comparison of submerged and air liquid interface exposure: limitations and differences in the toxicological effects evaluated in bronchial epithelial cells. Environ Res. 2025;279(Pt 2): 121856.

- Li, J., LoBue, A., Heuser, S. K., Leo, F., Cortese–Krott, M. M. Using diaminofluoresceins (DAFs) in nitric oxide research. Nitric Oxide 2021;115:44–54.

- Jiang S. et al. Real–time electrical detection of nitric oxide in biological systems with sub–nanomolar sensitivity. Nat Commun. 2013;4:2225.

- Brown M. D, Schoenfisch M. H. Electrochemical Nitric Oxide Sensors: Principles of Design and Characterization. Chem Rev. 2019;119(22):11551–75.

- Corti, P. et al. Globin X is a six–coordinate globin that reduces nitrite to nitric oxide in fish red blood cells. Proc Natl Acad Sci U S A 2016;113(30):8538–43.

- Zhao, X. J. et al. Mechanisms for cellular NO oxidation and nitrite formation in lung epithelial cells. Free Radic Biol Med. 2013;61:428–37.

- Lundberg, J. O. Weitzberg E, Gladwin MT. The nitrate–nitrite–nitric oxide pathway in physiology and therapeutics. Nat Rev Drug Discov. 2008;7(2):156–67.

- Rochon, E. R., Corti, P. Globins and nitric oxide homeostasis in fish embryonic development. Mar Genomics 2020;49:100721.

- Giordano, D., Verde, C., Corti, P. Nitric Oxide Production and Regulation in the Teleost Cardiovascular System. Antioxidants (Basel) 2022’11(5):957.

- MacArthur, P. H., Shiva, S., Gladwin, M. T. Measurement of circulating nitrite and S–nitrosothiols by reductive chemiluminescence. J Chromatogr B Analyt Technol Biomed Life Sci. 2007;851(1–2):93–105.

- Basu, S., Ricart, K., Gladwin, M. T., Patel, R. P, Kim–Shapiro, D. B. Tri–iodide and vanadium chloride based chemiluminescent methods for quantification of nitrogen oxides. Nitric Oxide 2022;121: 11–19.

- Brody, S. L., Yan, X. H., Wuerffel, M. K., Song, S. K., Shapiro, S. D. Ciliogenesis and left–right axis defects in forkhead factor HFH–4–null mice. Am J Respir Cell Mol Biol. 2000;23(1): 45–51.

- Jain, R. et al. Temporal relationship between primary and motile ciliogenesis in airway epithelial cells. Am J Respir Cell Mol Biol. 2010;43(6): 731–9.

- Eenjes, E. et al. A novel method for expansion and differentiation of mouse tracheal epithelial cells in culture. Sci Rep. 2018;8(1): 7349.

- You, Y., Richer, E. J., Huang, T., Brody, S. L. Growth and differentiation of mouse tracheal epithelial cells: selection of a proliferative population. Am J Physiol Lung Cell Mol Physiol. 2002;283(6): L1315–21.

- Pfeiffer, S., Leopold, E., Schmidt, K., Brunner, F., Mayer, B. Inhibition of nitric oxide synthesis by NG–nitro–L–arginine methyl ester (L–NAME): requirement for bioactivation to the free acid, NG–nitro–L–arginine. Br J Pharmacol 1996;118(6):1433–40.

- Rochon, E. R., Corti, P., Gladwin, M. T. Hemoglobin α in Pulmonary Endothelium: Ironing Out Nitric Oxide Signaling. Am J Respir Cell Mol Biol. 2017;57(6):639–41.

- van Faassen, E. E. et al. Nitrite as regulator of hypoxic signaling in mammalian physiology. Med Res Rev. 2009;29(5):683–741.